Project title: Upgrading national research infrastructures – RIUM

Compliance declaration with minimum technical REquirements of the offered Equipment

Name of the tenderer:

Address of the tenderer:

Subject of the public contract: Purchase of micro/nano lithographic and machining SEM/FIB system

Designation of the public contract: 302/18-RIUM7- FERI03.3/2020

INSTRUCTIONS:

- This Attachment No. 2 must be completed and signed. The tenderer must indicate the actual technical specifications of the equipment offered in line with all requirements.

- The tenderer shall upload this Attachment No. 2 to the e-JN information system in the section “Other Attachments”.

We declare that the offered equipment fully meets the minimum technical requirements of the client for the public contract titled "Purchase of a micro/nano lithographic and machining SEM/FIB system", and that we are ready to realize the order according to the following specification:

|  | EQUIPMENT |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  |  |  |  | DATASHEET |
| 1 | Ultra-High-Resolution FE-SEM Column: |  |  |  |  |  |  |
|  | - In-column SE and axial BSE detectors with energy filtering capability | SE in BSE detectors included: |  | - YES |  | - NO |  |
|  |  | Detectors with energy filtering capability: |  | - YES |  | - NO |  |
|  |  | In-column SE and axial BSE detectors position: |  | - YES |  | - NO |  |
|  | - SE detector for low vacuum conditions | - YES |  | - NO |  |  |  |
|  | - Retractable BSE detectors and SE detector | - YES |  | - NO |  |  |  |
|  | - Electron Beam Resolution: equal or better than 1.2 nm at 1 keV, and equal or better than 0.7 nm at 15 keV | Resolution at 1 keV: _________ nm |  |  |  |  |  |
|  |  | Resolution at 15 keV: _________ nm |  |  |  |  |  |
|  | - Electron beam landing energy range: minimal range from 200 eV–30 keV | Minimum electron beam landing energy: _________ eV |  |  |  |  |  |
|  |  | Maximum electron beam landing energy: _________ keV |  |  |  |  |  |
|  | - Electrostatic beam blanker | - YES |  | - NO |  |  |  |
|  | - Spot optimization without mechanical aperture positioning | - YES |  | - NO |  |  |  |
|  | - Probe current: minimal range from 5 pA–100 nA, continuously adjustable, the stability of SEM beam current < +/- 1 % for both short term (1h) and long term (24h) | Minimum probe current: ___________ pA |  |  |  |  |  |
|  |  | Maximum probe current: ___________ nA |  |  |  |  |  |
|  |  | Short term (1h) and long term (24h) probe current stability range: ±________ % |  |  |  |  |  |
|  | Focused Ion Beam Column: |  |  |  |  |  |  |
|  | - Gallium liquid metal ion source, lifetime equal or better than 3000 µAh (guaranteed) | Gallium liquid metal ion source lifetime: __________ µAh |  |  |  |  |  |
|  | - Ion Beam Resolution: equal or better than 3 nm at 30 keV | Ion Beam Resolution at 30 keV: ________ nm |  |  |  |  |  |
|  | - Piezo driven or motorized apertures | - YES |  | - NO |  |  |  |
|  | - Electrostatic beam blanker with integrated Faraday cup | - YES |  | - NO |  |  |  |
|  | - Ion beam minimum energy range: minimal range from 500 eV–30 keV | Ion beam minimum energy: ______________ eV |  |  |  |  |  |
|  |  | Ion beam maximum energy: ______________ keV |  |  |  |  |  |
|  | - Probe current: minimal range from 1 pA–90 nA | Minimum probe current: ___________ pA |  |  |  |  |  |
|  |  | Maximum probe current: ___________ nA |  |  |  |  |  |
|  | - Field of view: equal or larger than 0.5 mm | Field of view: _______ mm |  |  |  |  |  |
|  | Sample stage: |  |  |  |  |  |  |
|  | - Motorized, at least 5-Axis Stage (automatic positioning of the sample within focus must be provided during sample movement and rotation) | Motorized stage system: |  | - YES |  | - NO |  |
|  |  | Number of axis: _________ |  |  |  |  |  |
|  |  | Automatic positioning of the sample within focus: |  | - YES |  | - NO |  |
|  | - X, Y axis travel range: minimal 120 mm | X-axis travel range: ________ mm |  |  |  |  |  |
|  |  | Y-axis travel range: ________ mm |  |  |  |  |  |
|  | - Z-axis travel range: minimal 50 mm | Z-axis travel range: ________ mm |  |  |  |  |  |
|  | - Rotation: 360° | Rotation: _________ ° |  |  |  |  |  |
|  | Vacuum Chamber: |  |  |  |  |  |  |
|  | - High vacuum: 9x10-3 Pa to Low vacuum: 60 Pa or more | High vacuum: ______________ Pa |  |  |  |  |  |
|  |  | Low vacuum: ______________ Pa |  |  |  |  |  |
|  | - Pumping time: less than four (4) minutes | Pumping time: _________ min |  |  |  |  |  |
|  | - Width or Diameter: minimal 330 mm | Width or Diameter: __________ mm |  |  |  |  |  |
|  | - Height/Depth: minimal 270 mm | Height/Depth: ___________ mm |  |  |  |  |  |
|  | - Ports: minimal 18 | Number of ports: ___________ |  |  |  |  |  |
|  | - Feedthrough for 2 (two) standard telecom single mode optical fibers | - YES |  | - NO |  |  |  |
|  | - Suspension type: active | - YES |  | - NO |  |  |  |
|  | - Chamber view: minimum 2 (two) cameras | - YES |  | - NO |  |  |  |
|  | - Integrated plasma cleaner (Decontaminator) | - YES |  | - NO |  |  |  |
|  | - Sample exchange under vacuum conditions through airlock | - YES |  | - NO |  |  |  |
|  | - Possibility to make charge neutralization by electron flood gun, electron gun, or equivalent system, which actively prevents a charge build-up | - YES |  | - NO |  |  |  |
|  | - Additional systems for observation of non-conductive samples in the high vacuum, such as beam de-acceleration/gentle-beam sample potential control or equivalent (the system must provide good imaging capability for unconductive samples in high vacuum conditions) | Additional systems for observation of non-conductive samples in the high vacuum: |  | - YES |  | - NO |  |
|  |  | System provides good imaging capability for unconductive samples in high vacuum conditions: |  | - YES |  | - NO |  |
|  | Gas Injection Systems (for depostion of thin local films): |  |  |  |  |  |  |
|  | - GIS with 5 (five) independent precursors (including suitable reservoirs and capillaries) | - YES |  | - NO |  |  |  |
|  | - 3-axis motorized | - YES |  | - NO |  |  |  |
|  | - Choice of Precursors for the deposition of the following materials: Pt, W, C, SiOx, Hf2O, XeF2, Ta2O5, TiO2, Zr2O3 | - YES |  | - NO |  |  |  |
|  | Scanning Systems: |  |  |  |  |  |  |
|  | - 2 (two) independent scanning systems for SEM and FIB | - YES |  | - NO |  |  |  |
|  | - Line, frame accumulation and drift-corrected frame accumulation | - YES |  | - NO |  |  |  |
|  | - Dwell time: minimal range from 50 ns–10 ms, adjustable in steps or continuously | Minimum dwell time: ______________ ns |  |  |  |  |  |
|  |  | Maximum dwell time: ______________ ms |  |  |  |  |  |
|  |  | Dwell time adjustable in steps or continuously: |  | - YES |  | - NO |  |
|  | Image Acquisition: |  |  |  |  |  |  |
|  | - Equal or higher frame size than 16K x 16K, panorama size not limited | Frame size: _______ K x _______ K |  |  |  |  |  |
|  |  | Panorama size not limited: |  | - YES |  | - NO |  |
|  | - Minimal 5 (five) channels or more can be acquired simultaneously | - YES | - NO |  |  |  |  |
|  | Microscope Control PC: |  |  |  |  |  |  |
|  | - High Performance PC: equal or better than Intel Core i7/Xeon or equivalent, equal or larger than 16 GB RAM | - YES |  |  | - NO |  |  |
|  | - Equal or larger than 2 TB SSD NVME, graphic card: equal or better than Nvidia RTX 2060 or equivalent, Windows 10 Pro 64-bit or equivalent or newer | - YES |  |  | - NO |  |  |
|  | - Equal or better than 32” QHD Monitor | - YES |  |  | - NO |  |  |
|  | Electron and ion lithography modules: |  |  |  |  |  |  |
|  | - The system must contain all the necessary machine components for performing ion and electronic lithography | - YES | - NO |  |  |  |  |
|  | - Lithography software fully integrated into SEM-FIB software or separate third-party lithography software or combination of both fully compatible (integrated and third-party) | - YES | - NO |  |  |  |  |
|  | - Ability to draw and expose shapes like Point, Line, Rectangle, Circle, Annulus, Polygon, etc., plugin for easy editing of specific shapes and properties | - YES | - NO |  |  |  |  |
|  | - GDSII data format with EBL technology metadata (implemented using standard GDSII user properties) | - YES | - NO |  |  |  |  |
|  | - Possibility to directly control the spot size and probe current | - YES | - NO |  |  |  |  |
|  | - Digital nanopatterning engine, 16-bit DAC | - YES | - NO |  |  |  |  |
|  | - Available processing modes: Electron exposition, Electron etching, Electron Deposition, Ion etching, Ion deposition | - YES | - NO |  |  |  |  |
|  | - Dwell time equal or less than 80 ns/pixel for all shapes | Dwell time: ___________ns/pixel |  |  |  |  |  |
|  | - Multiple write fields (stitching) | - YES | - NO |  |  |  |  |
|  | - Stage navigation | - YES | - NO |  |  |  |  |
|  | - Proximity effect correction | - YES | - NO |  |  |  |  |
|  | - FIB end‐point detection | - YES | - NO |  |  |  |  |
|  | - Automatic pattern alignment | - YES | - NO |  |  |  |  |
|  | - Manual alignment mark detection | - YES | - NO |  |  |  |  |

| Place and date: | Seal: | Authorized signature: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.