Naziv projekta: Nadgradnja nacionalnih raziskovalnih infrastruktur – RIUM

Project title: Upgrading national research infrastructures – RIUM

Tehnične specifikacije za mikro/nano obdelovalni in litografski sistem s SEM/FIB

Technical specifications for a micro/nano lithographic and machining SEM/FIB system

- PREDMET JAVNEGA NAROČILA 						1. SUBJECT-MATTER OF THE PUBLIC CONTRACT

| Predmet javnega naročila je nabava mikro/nano obdelovalnega in litografskega sistema s SEM/FIB, vključno z dostavo, vnosom, namestitvijo, montažo, konfiguracijo, inštalacijo in zagonom naprave, skladno s tehničnimi specifikacijami. Lokacija dostave in vnos opreme: Univerza v Mariboru, Fakulteta za elektrotehniko, računalništvo in informatiko, Koroška cesta 46, 2000 Maribor, Slovenija, Laboratorij za elektro-optične in senzorske sisteme objekt G, čisti prostor G-105, skladno s prilogo 1 teh specifikacij, kjer je podan tloris prostora, kamor bo izbrani ponudnik moral opremo vnesti in namestiti. Naročnik bo ponudnikom omogočil ogled prostora na lokaciji vnosa in namestitve opreme, skladno s točko 11 Navodil ponudnikom. Rok dobave: najpozneje 100 (sto) dni od dneva pričetka veljavnosti pogodbe. Garancijski rok: najmanj 3 (tri) leta od dneva podpisa prevzemnega zapisnika. Celotna 3-letna garancija vključuje letni servis (brez potrošnega materiala). Za vso opremo velja, da mora zanjo v času garancijskega obdobja proizvajalec opreme zagotavljati dostop do vseh posodobitev programske opreme brez dodatnih stroškov. Vsi popravki, servisni paketi ter nadgradnje in podpora morajo biti od proizvajalca strojne oziroma programske opreme. Ponudnik mora za opremo zagotoviti: - da so vsi ponujeni moduli/komponente združljivi in medsebojno povezljivi, - da je mogoče vse zahtevane module/komponente povezati v enoten in brezhibno delujoč sistem v skladu s tehničnimi specifikacijami, - odpraviti vse napake, ki bi se pojavile zaradi morebitne medsebojne nezdružljivosti ali nepovezljivosti oz. zaradi neustrezne povezave v enoten, brezhibno delujoč sistem. V ceno je vključeno tudi izobraževanje ob prevzemu in inštalaciji opreme na lokacije dobavljene opreme. V ponudbi mora biti zajeti ves droben material, ki je potreben za priključitev naprav, tudi če ni izrecno naveden. | Subject-matter of the public contract is the purchase of an R&D platform for direct micromachining (FIB milling and FIB-assisted deposition), electron beam lithography, ion-beam lithography and sample observation with high spatial resolution, including delivery, placement, assembly, configuration, installation and starting of the device, in line with technical specifications. Equipment delivery and placement: University of Maribor, Faculty of Electrical Engineering and Computer Science, Laboratory for Electro-Optics and Sensor Systems, Koroška cesta 46, 2000 Maribor, Slovenia, Building G, clean room G-105, in line with Attachment 1 to these specifications providing the floor plan of the clean room where the equipment shall be delivered and placed by the successful tenderer. The contracting authority will enable the tenderers to inspect the clean room on site in line with Point 11 of the Instructions to tenderers. Delivery time: at the latest 100 (one hundred) days from the day of entry into force of the contract. The guarantee period: At least 3 (three) years from the day of signature of the take-over protocol, including annual service (without consumables). The equipment manufacturer is obliged to provide access to all software updates at no additional cost during the guarantee period. All patches, service packs, upgrades and support must be provided by the hardware or software manufacturer. The tenderer of the equipment must guarantee: - that all modules/components offered are compatible and interconnected, - that all required modules/components can be integrated into a single and seamlessly functioning system in line with technical specifications, - to eliminate all errors resulting from possible incompatibility or mismatch, or from improper connection to a single, seamlessly functioning system. The price includes training upon acquisition and installation of the device on site. The offer must include all the fine material, necessary for connecting the devices, even if it is not explicitly stated. |
| --- | --- |

TEHNIČNE ZAHTEVE

- TECHNICAL SPECIFICATIONS

| Mikro/nano obdelovalni in litografski sistem s SEM/FIB Univerza v Mariboru naroča vsestransko raziskovalno-razvojno platformo za neposredno mikro obdelavo (FIB rezkanje in možnost depozicije s FIB), elektronsko litografijo, litografijo z ionskim curkom in opazovanje vzorcev z visoko prostorsko ločljivostjo. Sistem mora vključevati celostni vakuumski sistem z veliko vakuumsko komoro z več priključki, vir elektronskih žarkov, vir ionskega curka, detektorje, vsaj 5-osno računalniško krmiljeno mizo za vzorce, večkanalni sistem za dovod procesnih plinov in ustrezno programsko okolje s podporo za zajemanje slik vzorcev (elektronsko ali z ionskim curkom), litografijo, rezkanje s FIB in depozicijo tankih plasti s pomočjo FIB. Sistem mora biti zasnovan in zgrajen tako, da podpira: - izdelavo litografskih aplikacij z elektronskim curkom (osvetljevanje različnih fotorezistov na polprevodniških rezinah z elektronskim osvetljevanjem), - izdelavo litografskih aplikacij z ionskim curkom (osvetljevanje različnih fotorezistov na polprevodniških rezinah z ionskim curkom), - izdelavo mikro in nano struktur z ionskim curkom (mikro in nano-rezkanje), - zajemanje slik vzorcev z uporabo SEM sistema v nanometrski ločljivosti, - zajemanje slik vzorcev s skanirnim sistemom z uporabo ionskega curka, - delovanje v področju visokega in nizkega vakuuma (načrtovano je frekventno opazovanje neprevodnih vzorcev, kot je na primer steklo), - opazovanje/pregled struktur, grajenih iz lahkih elementov, kar zahteva sisteme/tehnologije za upočasnjevanje snopa (»beam de-acceleration«), s katerimi je mogoče doseči visoko ločljivost ob nizkih pospeševalnih napetostih, - depozicijo ob podpori FIB v kontrolirani atmosferi (s sistemom za dovod procesnih plinov preko vsaj 5 (pet) neodvisnih prekurzorjev), - možnost hkratne uporabe FIB in SEM. V nadaljevanju so podane minimalne tehnične zahteve. Ponudba mora izpolnjevati ali presegati vse minimalne tehnične zahteve. | Micro/nano lithographic and machining SEM/FIB system The University of Maribor is purchasing a versatile R&D platform for direct micromachining (FIB milling and FIB-assisted deposition), electron beam lithography, ion-beam lithography, and sample observation with high spatial resolution. The system shall include a complete vacuum system with a large vacuum chamber and multiple ports, electron beam source, ion beam source, detectors, 5-axis fully computer controlled sample stage, multiple gas injection system, and computer control with appropriate software for electron and ion beam imaging, lithography, FIB-milling and FIB assisted deposition. The system shall be designed and built to support: - electron beam lithographic applications (e-beam illumination of wafers containing different photoresists), - ion beam lithographic applications (ion-beam illumination of wafers containing different photoresists), - ion beam micro and nano-milling applications, - imaging of samples using scanning electron beam (SEM function) with nanometer resolution, - imaging of samples using scanning ion beam, - high and low vacuum operation (unconductive samples, as glass will be frequently used), - examination of structures made from light elements (e.g. it is expected to include systems like beam de-acceleration technology or similar to achieve high imaging resolutions at low acceleration voltages), - FIB-assisted deposition using gas injection system that can support at least 5 (five) independent precursors. - simultaneous operation of FIB and SEM. Minimum technical requirements are given below. The offer must match or exceed all minimum technical requirements. |
| --- | --- |

| Številka postavke Item No. | Število kosov No. of pieces | ZAHTEVANO | REQUIRED |
| --- | --- | --- | --- |
| 1 | 1 | Minimalne tehnične zahteve: Kolona FE-SEM z ultra visoko ločljivostjo: - SE detektor in soosno postavljen BSE detektor z možnostjo filtriranja energije. - SE detektor združljiv z delovanjem v nizkem vakuumu - Izvlečna BSE in SE detektorja - Ločljivost elektronskega curka: enaka ali boljša od 1.2 nm pri 1 keV in enaka ali boljša od 0.7 nm pri 15 keV - Energija elektronskega curka: vsaj v območju med 200 eV–30 keV - Elektrostatični sistem za prekrivanje žarka - Možnost optimizacije pike brez mehanskega pomika zaslonke - Tok sonde: vsaj v območju med 5 pA–100 nA, kratkoročna (1h) in dolgoročna (24h) stabilnost SEM toka sonde v območju < +/- 1 % Kolona z virom za fokusiran ionski curek: - Galijev ionski vir, z življenjsko dobo ≥ 3000 µAh - Ločljivost ionskega curka: enaka ali boljša od 3 nm pri 30 keV - Motorizirane ali piezo krmiljene zaslonke - Elektrostatični sistem za prekrivanje žarka z vgrajeno Faradejevo šalčko (kletko) - Energija ionskega curka: vsaj v območju med 500 eV–30 keV - Tok sonde: vsaj v območju med 1 pA–90 nA - Vidno področje: ≥ od 0.5 mm Miza za vzorce - Motoriziran, ≥ 5-osni sistem z možnostjo avtomatskega pozicioniranja in fokusiranja vzorca med premikanjem - Pomik po X in Y-osi : ≥ 120 mm - Pomik po Z-osi: ≥ 50 mm - Rotacija: 360° Vakuumska komora - Črpalni sistem, sposoben doseganja vakuuma od 9x10-3 Pa (visokega vakuuma) do 60 Pa (nizkega vakuuma) ali več - Čas do doseganja delovnega vakuuma: v manj kot štirih (4) minutah - Širina ali premer: ≥ 330 mm - Višina/Globina: ≥ 270 mm - Priključki na vakuumski komori: ≥ 18 - Priključek za 2 (dve) standardni telekomunikacijski optični vlakni (»feedthrough«) - Vzmetenje: aktivno - Vizualni sistem: ≥ 2 (dve) kameri - Integriran plazemski čistilnik (Dekontaminator) - Možnost izmenjave vzorcev, medtem ko je sistem pod delovnim vakuumom - Možnost nevtralizacije naboja s poplavno elektronsko pištolo (»electron flood gun,«), elektronsko pištolo (»electron gun«) ali enakovrednim sistemom, ki aktivno preprečuje kopičenje naboja na vzorcu - Dodatni sistemi za opazovanje neprevodnih vzorcev v visokem vakuumu, kot je na primer sistem za upočasnjevanje snopa (»beam de-acceleration/gentle-beam sample potential control system«) ali enakovreden (sistem mora zagotavljati kvaliteten prikaz neprevodnih vzorcev pod visokim vakuumom). Sistem za dovod procesnih plinov (za nalaganje lokalnih tankih plasti): - GIS s 5 (petimi) neodvisnimi prekurzorji (s pripadajočimi rezervoarji in kapilarami) - Motoriziran 3-osni - Možnost izbire prekurzorjev za depozicijo naslednjih materialov: Pt, W, C, SiOx, Hf2O, XeF2, Ta2O5, TiO2, Zr2O3 Skanirni sistemi - Neodvisna skanirna sistema za SEM in FIB - Akumulacija črt in okvirjev ter akumulacija kompenziranih okvirjev (»Line, frame accumulation and drift-corrected frame accumulation«) - Zadrževalni čas nad vzorcem (»Dwell time«): vsaj v območju med 50 ns–10 ms, nastavljiv zvezno ali po korakih Sistem za zajem slike - Ločljivost sistema: ≥ 16K x 16K, brez omejitve panoramskega pogleda/velikosti - Hkraten zajem iz ≥ 5 kanalov Računalnik za nadzor mikroskopa - Delovna postaja: ≥ Intel Core i7/Xeon ali enakovreden, ≥ 16 GB RAM, - ≥ 2 TB diska SSD NVME, grafična kartica: ≥ Nvidia RTX 2060 ali enakovredna, operacijski sistem: Windows 10 Pro 64-bit oz. enakovreden ali novejši. - Monitor: ≥ 32” QHD Moduli za elektronsko in ionsko litografijo - Sistem mora vsebovati vse potrebne strojne komponente za izvajanje ionske in elektronske litografije - Programska oprema za litografijo, integrirana kot del osnovnega SEM-FIB sistema ali kot samostojen programski paket ali kot kombinacija (osnovnega in samostojnega), popolnoma združljiv(a) z osnovnim SEM-FIB sistemom - Programska oprema mora imeti podporo za risanje in razvijanje več različnih geometrijskih oblik, kot so pike, črte, pravokotniki, krogi (obročaste oblike), poligoni (ponavljajoči se vzorci), ipd. ter možnost (lahko v obliki vtičnika) za enostavno naknadno urejanje slednjih - Podpora za GDSII format s tehnologijo metapodatkov EBL (implementirano preko standardnih klicev GDSII lastnosti) - Možnost direktne nastavitve velikosti pike in toka na sondi - Digitalni sistem za izdelavo nano vzorcev, 16-bitni digitalno analogni pretvornik - Možni procesni načini: Elektronska ekspozicija, Elektronsko jedkanje, Elektronska depozicija, Ionsko jedkanje, Ionska depozicija - Zadrževalni čas nad vzorcem (»Dwell time«): enak ali krajši od 80 ns/pixel ne glede na obliko vzorca - Možnost združitve (»stitching«) več litografskih vzorcev - Sistem za vodenje mize za vzorce ob uporabi FIB litografije (vpisovanja vzorca) - Sistem za korekcijo vzorca zaradi približevanja ionskega curka (»Proximity effect correction«) - Detekcija končnega položaja FIB sonde - Avtomatska poravnava litografskih vzorcev - Sistem za detekcijo oznak za poravnavo vzorcev | Minimum technical requirements: Ultra-High-Resolution FE-SEM Column: - In-column SE and axial BSE detectors with energy filtering capability - SE detector for low vacuum conditions - Retractable BSE detectors and SE detector - Electron Beam Resolution: equal or better than 1.2 nm at 1 keV, and equal or better than 0.7 nm at 15 keV - Electron beam landing energy range: minimal range from 200 eV–30 keV - Electrostatic beam blanker - Spot optimization without mechanical aperture positioning - Probe current: minimal range from 5 pA–100 nA, continuously adjustable, the stability of SEM beam current < +/- 1 % for both short term (1h) and long term (24h) Focused Ion Beam Column: - Gallium liquid metal ion source, lifetime equal or better than 3000 µAh (guaranteed) - Ion Beam Resolution: equal or better than 3 nm at 30 keV - Piezo driven or motorized apertures - Electrostatic beam blanker with integrated Faraday cup - Ion beam minimum energy range: minimal range from 500 eV–30 keV - Probe current: minimal range from 1 pA–90 nA - Field of view: equal or larger than 0.5 mm Sample Stage - Motorized, at least 5-axis stage (automatic positioning of the sample within focus must be provided during sample movement and rotation) - X, Y axis travel range: minimal 120 mm - Z axis travel range: minimal 50 mm - Rotation: 360° Vacuum Chamber - High vacuum: 9x10-3 Pa to Low vacuum: 60 Pa or more - Pumping time: less than four (4) minutes - Width or diameter: minimal 330 mm - Height/Depth: minimal 270 mm - Ports: minimal 18 - Feedthrough for 2 (two) standard telecom single mode optical fibres - Suspension type: active - Chamber view: minimum 2 (two) cameras - Integrated plasma cleaner (Decontaminator) - Sample exchange under vacuum conditions through airlock - Possibility to make charge neutralization by electron flood gun, electron gun, or equivalent system, which actively prevents a charge build-up - Additional systems for observation of non-conductive samples in the high vacuum, such as beam de-acceleration/gentle-beam sample potential control or equivalent (the system must provide good imaging capability for unconductive samples in high vacuum conditions). Gas Injection Systems (for depostion of thin local films): - GIS with 5 (five) independent precursors (including suitable reservoirs and capillaries) - 3-axis motorized - Choice of precursors for the deposition of the following materials: Pt, W, C, SiOx, Hf2O, XeF2, Ta2O5, TiO2, Zr2O3 Scanning Systems - 2 (two) independent scanning systems for SEM and FIB - Line, frame accumulation and drift-corrected frame accumulation - Dwell time: minimal range from 50 ns–10 ms, adjustable in steps or continuously Image Acquisition - Equal or higher frame size than 16K x 16K, panorama size not limited - Minimal 5 (five) or more channels can be acquired simultaneously Microscope Control PC - High performance PC: equal or better than Intel Core i7/Xeon or equivalent; equal or larger than 16 GB of RAM - Equal or larger than 2 TB SSD NVME, graphic card: equal or better than Nvidia RTX 2060 or equivalent, Windows 10 Pro 64-bit or equivalent or newer - Equal or better than 32” QHD Monitor Electron and ion lithography modules - The system must contain all the necessary machine components for performing ion and electronic lithography - Lithography software fully integrated into SEM-FIB software or separate third-party lithography software or combination of both fully compatible (integrated and third-party) - Possibility to directly control the spot size and probe current - Ability to draw and expose shapes like point, line, rectangle, circle, annulus, polygon, etc., plugin for easy editing of specific shapes and properties - GDSII data format with EBL technology metadata (implemented using standard GDSII user properties) - Digital nanopatterning engine, 16-bit DAC - Available processing modes: Electron exposition, Electron etching, Electron Deposition, Ion etching, Ion deposition - Dwell time equal or less than 80 ns/pixel for all shapes - Multiple write fields (stitching) - Stage navigation - Proximity effect correction - FIB end‐point detection - Automatic pattern alignment - Manual alignment mark detection |

Priloga 1: Tloris prostora G-105
Attachment 1: Floor plan of room G-105

Priloga 1: Tloris prostora G-105

Attachment 1: Floor plan of room G-105

Operacijo sofinancirata Republika Slovenija, Ministrstvo za izobraževanje, znanost in šport in Evropska unija iz Evropskega sklada za regionalni razvoj. Operacija se izvaja v okviru Operativnega programa za izvajanje evropske kohezijske politike 2014-2020 v okviru prednostne osi 1: »Mednarodna konkurenčnost raziskav, inovacij in tehnološkega razvoja v skladu s pametno specializacijo za večjo konkurenčnost in ozelenitev gospodarstva«, prednostne naložbe 1.1 »Krepitev infrastrukture za raziskave in inovacije ter zmogljivosti za razvoj odličnosti na tem področju, pa tudi spodbujanje pristojnih centrov, zlasti takšnih, ki so evropskega pomena«, specifičnega cilja 1.1.1 »Učinkovita uporaba raziskovalne infrastrukture ter razvoj znanja/kompetenc za boljše nacionalno in mednarodno sodelovanje v trikotniku znanja«.

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.