TECHNICAL SPECIFICATIONS

| Institute: | Research Centre of the Slovenian Academy of Sciences and Arts |  |  |
| --- | --- | --- | --- |
| Equipment: | CLIMATE CHAMBER and compatible CONTROLER | Quantity: | 2+2 |

| Final location of equipment | BIJH ZRC SAZU, Zagorica 20, 1292 Ig |
| --- | --- |

| Contact details | Surname and first name | Phone | Email address |
| --- | --- | --- | --- |
| Contact person: | Kralj-Fišer Simona | 01 4706333 | simonakf@zrc-sazu.si |

| Technical specifications | CLIMATE CHAMBER 1 |
| --- | --- |
|  | - At least 4 shelves |
|  | - T = -10 – 44°C, RH = 45% – 80% |
|  | - Possible inflow of fresh air |
|  | - Height between shelves 25 - 27 cm |
|  | - Total shelf surface area > 2.2 m2 |
|  | - Even light source, LED; intensity 400μmol/m-2/s-1 |
|  | - Glass door, allowing easy access to shelves |
|  | - Capability of misting/watering |
|  | - USB or Etheret |
|  | - Backup battery in case of power outage |
|  | - Stainless steel interior |
|  | - Continuous condensation unit, for tighter temperature tolerance (up to ± 0,3 ° C) and longer lifespan |
|  | - Air filter before coil |

| Additional requirements: | - delivery to the location BIJH ZRC SAZU (Zagorica 20, 1292 Ig, Slovenia) | da/yes ne/no |
| --- | --- | --- |
|  | - installation of equipment | da/yes ne/no |
|  | - performing basic diagnostics of equipment operation | da/yes ne/no |
|  | - equipment warranty: 3 years | da/yes ne/no |
|  | - the equipment provider must provide an authorized official service in the Republic of Slovenia | da/yes ne/no |

| Technical specifications: | CLIMATE CHAMBER 2 |
| --- | --- |
|  | - At least 4 shelves |
|  | - T = 4 – 44°C, RH = up to 70% |
|  | - Possible inflow of fresh air |
|  | - Height between shelves 23 - 27 cm |
|  | - Total shelf surface area > 2.2 m2 |
|  | - Even light source, LED; intensity up to 1000μmol/m-2/s-1 |
|  | - Glass door, allowing easy access to shelves |
|  | - Capability of misting/watering |
|  | - USB or Etheret |
|  | - Backup battery in case of power outage |
|  | - Stainless steel interior |
|  | - Air filter before coil |

| Additional requirements: | - delivery to the location BIJH ZRC SAZU (Zagorica 20, 1292 Ig, Slovenia) | da/yes ne/no |
| --- | --- | --- |
|  | - installation of equipment | da/yes ne/no |
|  | - performing basic diagnostics of equipment operation | da/yes ne/no |
|  | - equipment warranty: 3 years | da/yes ne/no |
|  | - the equipment provider must provide an authorized official service in the Republic of Slovenia | da/yes ne/no |

| Technical specifications | KONTROLER 1 |
| --- | --- |
|  | - Temperature regulation in the range of -10-44°C +/- 0.5°C |
|  | - Humidity regulation in area between 40% to 80% |
|  | - Light intensity regulation |
|  | - Regulation of day – night temperature, humidity (and lighting) |
|  | - Long distance regulation (e-mail or similar) |
|  | - 24 hour settings program, for several days in advance |
|  | - Alarm |

| Additional requirements: | - delivery to the location BIJH ZRC SAZU (Zagorica 20, 1292 Ig, Slovenia) | da/yes ne/no |
| --- | --- | --- |
|  | - installation of equipment | da/yes ne/no |
|  | - performing basic diagnostics of equipment operation | da/yes ne/no |
|  | - equipment warranty: 3 years | da/yes ne/no |
|  | - the equipment provider must provide an authorized official service in the Republic of Slovenia | da/yes ne/no |

| Technical specifications | KONTROLER 2 |
| --- | --- |
|  | - Temperature regulation in the range of 4-44°C +/- 0.5°C |
|  | - Humidity regulation in area up to 70% |
|  | - Light intensity regulation |
|  | - Regulation of day – night temperature, humidity (and lighting) |
|  | - Long distance regulation (e-mail or similar) |
|  | - 24 hour settings program, for several days in advance |
|  | - Alarm |

| Additional requirements: | - delivery to the location BIJH ZRC SAZU (Zagorica 20, 1292 Ig, Slovenia) | da/yes ne/no |
| --- | --- | --- |
|  | - installation of equipment | da/yes ne/no |
|  | - performing basic diagnostics of equipment operation | da/yes ne/no |
|  | - equipment warranty: 3 years | da/yes ne/no |
|  | - the equipment provider must provide an authorized official service in the Republic of Slovenia | da/yes ne/no |

| Common additional technical specifications of the equipment: | - Equipment must be from the same manufacturer and fully compatible with each other. |
| --- | --- |
|  | - The delivery must include the manufacturer's certificates of equipment conformity with the specified specifications. |
|  | - The equipment must allow all new software upgrade for a period of at least 1 year. |
|  | - The provider must provide training to work with the equipment (at least 3 hours, for at least 2 people), user instructions and free technical support by phone and e-mail for a period of at least 1 year. |