Project name: Nadgradnja nacionalnih raziskovalnih infrastruktur – RIUM

# Technical specifications for spectroelectrochemical equipment

# SUBJECT OF PUBLIC TENDER

LOT 1: X-ray photoelectron spectroscopy

X-ray photoelectron spectroscopy (XPS) is a surface analytical instrument that determines the elemental composition and the environment of an element - the oxidation state and the chemical bond of the elements. It will be used for advanced analyzes and studies of various materials. It is used for smooth and rough surfaces of thin films as well as for more complicated powders and fine particles, for the analysis of surface atomic layers as well as lamellar materials with a thickness of a few μm. It enables automatic angle-resolved XPS measurements, heating and cooling of the sample, and depth profiling of organic matter using cluster technology and monoatomic argon ions.

LOT 2: Multi-channel potentiostat/galvanostat

Possibility of performing various electrochemical techniques: amperometric, voltammetric, pulse, galvanostatic. Possibility of performing electrochemical impedance spectroscopy. All techniques can be performed separately on up to 5 or 8 samples per one instrument at a time. Possibility to perform impedance measurements in the frequency range between 0.01 mHz to 1 MHz.

# TECHNICAL REQUIEREMENTS

Delivery deadline:

LOT 1: The supplier must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 1.19 of the technical specifications at the location University of Maribor, Faculty of Chemistry and Chemical Technology, Smetanova ulica 17, 2000 Maribor, no later than 220 days from the entry into force of the contract.

LOT 2: The supplier must deliver the equipment that is the subject of this public procurement to the location University of Maribor, Faculty of Chemistry and Chemical Technology, Smetanova ulica 17, 2000 Maribor, no later than 60 days from the entry into force of the contract.

Warranty period:

LOT 1: For the delivered equipment, the supplier provides a 1 (one) year warranty for a flawless technical operation.

LOT 2: For the supplied equipment, the supplier provides a 3 (three) year warranty for a flawless technical operation.

LOT 1: X-ray photoelectron spectroscopy, XPS

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 1.1 | 1 | Large area performance on a clean Ag sample using the Ag 3d5/2 photoelectron peak must exceed: - peak resolution (full width at half maximum) of 0.50 eV with a sensitivity of at least 300,000 CPS, - peak resolution (full width at half maximum) of 0.60 eV with a sensitivity of at least 1,000,000 CPS. Small area performance on a clean Ag sample using the Ag 3d5/2 photoelectron peak must exceed: - minimum spot size: 20 µm or less, - peak resolution (full width at half maximum) of 0.60 eV with a sensitivity 2,000 CPS or more. Minimum charge neutraliser performance: - Ester peak resolution (full width at half maximum) < 0.82 eV with a sensitivity more than 20,000 CPS. XPS imaging: - Guaranteed spatial resolution of XPS images less than 10 µm. |
| 1.2 | 1 | Automated five axis (x, y z, rotation and tilting) precision manipulator system for translation of the sample holders. |
| 1.3 | 1 | Three sample holders: - dual height sample holder for samples with the thicknesses more than 15 mm, - sample holder for flat sample mounting with a claw for loading standard 15 mm diameter sample stubs or more, - Z-rotation sample holder (primarily used in depth profiling experiments). |
| 1.4 | 1 | Sample entry chamber must include: - pumping kit (turbo molecular pump, wide range ion gauge, automated solenoid valve, rotary vane pump, oil mist filter and other vacuum components), - fully computer-controlled pump/vent sequences, - auto-exchange magazine for automatic transfer of the sample holder from sample entry chamber into the sample analysis chamber, - an optical camera in the sample entry chamber which automatically records an image of the entire sample holder inside the vacuum chamber. It must be possible to use this image to define positions for analysis, - it must be possible to load 3 or more sample holders into the sample entry chamber, - the minimum area of the sample holder available for sample mounting must be 2000 mm2, - the instrument must be able to automatically transfer the sample holder to the sample analysis chamber without user intervention, - the instrument must be able to automatically exchange a sample holder in the sample analysis chamber for a second sample holder in the sample entry chamber without user intervention. |
| 1.5 | 1 | The main analysis chamber must include: - pumping kit (turbo molecular pump, titanium sublimation pump, wide range ion gauge, automated solenoid valve, rotary vane pump, oil mist filter and other vacuum components), - an optical microscope in the sample analysis chamber with a view of the sample analysis position, - the minimum pressure inside the sample analysis chamber must be 7 x 10-10 mbar or less, - the software must include an automated routine to set the sample height prior to analysis, - a set of standard samples including: Cu; Au; Ag; grids for alignment; and Faraday cup must be permanently mounted in the sample analysis chamber. |
| 1.6 | 1 | Source of standard monochromated Al X-rays and high energy X-rays (X-ray energy more than 2,000 eV). It must be possible to automatically switch between standard monochromated Al X-rays and the high energy source without user intervention. |
| 1.7 | 1 | Ion source The instrument must include an argon ion source with the following specifications: - ability to generate monatomic and cluster ions, - acceleration voltage of 5 kV or higher in monatomic mode, - acceleration voltage of 8 kV or higher in cluster ion mode, - minimum cluster size of 2000 atoms per cluster or more. |
| 1.8 | 1 | Licensed software for measuring and processing data for an unlimited number of users. The software must include an automated peak identification routine. |
| 1.9 | 1 | PC specifications: - MS Windows 10 Professional 64 bit operating system, - Intel® Core® i7 3.6 GHz processor or equivalent, - HP Z2 Tower G4 Workstation or equivalent, - 8 Gb RAM, DVD writer, 1 Tb hard disk, 2 x 24” IPS LED backlit monitors. |
| 1.10 | 1 | Cooled re-circulating water supply, water chiller for the entire instrument, water pressure boost pump. |
| 1.11 | 1 | Mains power distribution unit for the main instrument. |
| 1.12 | 1 | Device/vessel for transporting air sensitive samples from a glove box to the instrument without exposing the samples to the ambient atmosphere. |
| 1.13 | 1 | Direct heating and cooling of the sample in the main analysis chamber, at the sample analysis position. - Possibility to cool the sample to at least -100 degrees Celsius at the sample analysis position. - Possibility to heat the sample to at least 800 degrees Celsius at the sample analysis position. Heating and cooling kit should include: - Sample holder for heating and cooling of the sample in the main sample chamber, - PC software for control of heating and cooling devices. |
| 1.14 | At least 3 | Software for angle-resolved XPS measurements. |
| 1.15 | At least 20 | Licences for off-line data processing software. |
| 1.16 | 1 | Additional spare consumable items must include: - 3 x Mono anodes, - 3 x Mono filaments, - 3 x Ion source filaments, - 3 x Charge neutralizer filaments, - 3 x Detector consumables. |
| 1.17 | 1 | A set of special tools and spares required for initial maintenance of the instrument. |
| 1.18 | 1 | Warranty: 12 months or more. |
| 1.19 | 1 | Delivery and installation on the client's premises. Condition for successful execution of the item and signing of the acceptance report: - the supply and installation and commissioning of the equipment by the supplier, - introduction to equipment and basic education referred to item 1.20, - submission of operating instructions and other technical documentation. |
| 1.20 | 1 | Basic operation training Acquaintance with the equipment and basic training for 3 users lasting at least 5 working days (also includes the submission of all necessary data and passwords; training is carried out at the location FKKT UM, determined by the client). The item is completed after the delivery, installation and startup of the equipment. Acquaintance with the equipment and basic training must be performed by the engineer of the provider. |
| 1.21 | 1 | Additional advanced training for a period of 5 working days at the location of the delivered device, by a specialist provider after three months from the start of the equipment. Additional advanced training for 3 users is intended to educate users on new and advanced methods of using the instrument and follows only after the basic training on the use of the instrument, until the end of the warranty period. |
| 1.22 | 1 | Provided remote access to the device for the needs of calibration and control of the device at least for the duration of the warranty period (1 year). The instrument design must facilitate remote operation and fault diagnostics. |

LOT 2: Multi-channel potentiostat/galvanostat

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 2.1 | 2 | 5-channel potenciostat/galvanostat for electrochemical measurements. Each channel must allow electrochemical measurements and electrochemical impedance spectroscopy measurements. Each channel must be galvanically isolated and allow a frequency range measurements of electrochemical impedance spectroscopy up to 1 MHz. Each channel must allow bipotentiostatic measurements, IR compensation, potential measurement range from -10 V to +10 V, electric current range from 100 pA to at least 10 mA. |
| 2.2 | 1 | 8-channel potenciostat/galvanostat for electrochemical measurements. Each channel must allow electrochemical measurements and electrochemical impedance spectroscopy measurements. Each channel must be galvanically isolated and allow a frequency range of electrochemical impedance measurement up to 1 MHz. Each channel must allow bipotentiostatic measurements, IR compensation, potential measurement range from -10 V to +10 V, electric current range from 100 pA to at least 10 mA. |
| 2.3 | 40 | Graphite rod - counter electrode. |
| 2.4 | 15 | Coating evaluation test cell for flat samples. |
| 2.5 | 12 | Cell set for corrosion studies, including: - glass vessel (volume 1 L), - two graphite counter electrodes with holders, - one saturated calomel electrode with salt bridge, - one sample holder for flat coupons, - two platinum counter electrodes, - spare O-rings and Teflon parts, - parts for the Stern-Makrides arrangement. |
| 2.6 | 48 | Glassy carbon working electrode, diameter 3 mm. |
| 2.7 | 30 | Reference electrode with Ag/AgCl in aqueous KCl. |
| 2.8 | 30 | Counter electrode made of platinum wire. |
| 2.9 | 6 | Magnetic stirrer with appropriate electronics that enable the control of the stirrer by potentiostat. |
| 2.10 | 10 | Additional cables for use with the potentiostat. |
| 2.11 | 1000 | SPE electrodes (Ag/AgCl reference electrodes, carbon working electrode, platinum counter electrode). |
| 2.12 | 6 | Simple cell stands. |

The operation is co-financed by the Republic of Slovenia, the Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is implemented under the Operational Program for the Implementation of European Cohesion Policy 2014-2020 under Priority Axis 1: "International competitiveness of research, innovation and technological development in line with smart specialization for greater competitiveness and greening of the economy", priority investment 1.1 "Strengthening infrastructure for research and innovation and capacity to develop excellence in this field, as well as the promotion of competent centers, especially those of European importance ", specific objective 1.1.1" Effective use of research infrastructure and development of knowledge/competences for better national and international cooperation in the triangle knowledge ”.