Project title: Upgrading National Research Infrastructures – RIUM

Compliance declaration with minimum technical REquirements of the offered Equipment

Name of the tenderer:

Address of the tenderer:

Subject of the public procurement: Purchase of a 3D Laser Scanning Confocal Microscope

Designation of the public procurement: 302/50 - RIUM23-FERI03.7/2020

INSTRUCTIONS:

- This Annex No. 2 must be completed and signed. The tenderer must indicate the actual technical specifications of the equipment offered in line with all requirements.

- The tenderer shall upload this Annex No. 2 in Word and PDF format (including technical documentation proving the fulfillment of all technical requirements to the e-JN information system in the section “Other Annexes”.

We declare that the offered equipment fully meets the minimum technical requirements of the client for the public procurement titled "Purchase of a 3D Laser Scanning Confocal Microscope", and that we are ready to realize the order according to the following specification:

|  | EQUIPMENT |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  |  |  |  |  |  | DATASHEET |
| 2.1 2.2 2.3 | Microscope: |  |  |  |  |  |  |  |  |  |
|  | - Photomultiplier with 16-bit resolution for laser scanning with a single pinhole technique | - YES |  |  | - NO |  |  |  |  |
|  | - A confocal system with one or more pinholes | - YES |  |  | - NO |  |  |  |  |
|  | - Confocal optical system with two light sources: - laser source (with wavelength shorter than 410 nm) for 3D reconstruction of the observed sample's surface - broadband (white) light source for capturing 2D images and color composition of 3D images | - YES |  |  |  | - NO |  |  |  |
|  | - Set of included lenses, indicated in the table below: Objective lens; Numerical aperture - NA; Working distance [mm]. 5x; ≥ 0,13; ≥ 20. 10x; ≥ 0,3; ≥ 10. 20x; ≥ 0,4; ≥ 6 ≤ 15. 50x; ≥ 0,55; ≥ 5. 50x; ≥ 0,9; ≥ 0,3 ≤ 0,4. 100x; ≥ 0,8; ≥2 ≤ 4. | Set of included lenses: |  | - YES |  |  |  | - NO |  |
|  | - Motorized movement of X and Y axis | - YES |  |  | - NO |  |  |  |  |
|  | - X, Y axis travel range: equal or better than 100 mm | X-axis travel range: ________ mm |  |  |  |  |  |  |  |
|  |  | Y-axis travel range: ________ mm |  |  |  |  |  |  |  |
|  | - Z-axis travel range: equal or better than 25 mm | Z-axis travel range: ________ mm |  |  |  |  |  |  |  |
|  | - Sample height: equal or better than 50 mm | Sample height: ____________ mm |  |  |  |  |  |  |  |
|  | - Minimum field of view smaller or equal than 18 μm and maximum field of view larger or equal than 4000 μm | Minimum field of view: _____________ μm Maximum field of view: _____________ μm |  |  |  |  |  |  |  |
|  | - Height display resolution equal or better than 0.5 nm | Height display resolution: ___________ nm |  |  |  |  |  |  |  |
|  | - Width display resolution equal or better than 1 nm | Width display resolution: ___________ nm |  |  |  |  |  |  |  |
|  | - Height measurement repeatability: Magnification; Repeatability σ. 5x; equal or better than 500 nm. 10x; equal or better than 100 nm. 20x; equal or better than 50 nm. 50x; equal or better than 20 nm. | - YES |  |  |  | - NO |  |  |
|  | - Width measurement repeatability: Magnification; Repeatability σ. 5x; equal or better than 400 nm. 10x; equal or better than 400 nm. 20x; equal or better than 120 nm. 50x; equal or better than 50 nm. | - YES |  |  |  | - NO |  |  |
|  | - Microscope accuracy for Z-axis: 0.2 + L/100 μm or better | Microscope accuracy for Z-axis: ___________ μm |  |  |  |  |  |  |  |
|  | - Microscope accuracy for X and Y axis: ± 2,0 % or better | Microscope accuracy for X and Y axis: ___________ % |  |  |  |  |  |  |  |
|  | - 3D scanning of angles on the sample equal or larger than 80 degrees | - YES |  |  | - NO |  |  |  |  |
|  | - Laser 3D scanning and analysis of optically transparent samples | - YES |  |  | - NO |  |  |  |  |
|  | - Auto-Navigation for reliable positioning at high magnification | - YES |  |  | - NO |  |  |  |  |
|  | - Automatic measurement and analysis of inserted sample | - YES |  |  | - NO |  |  |  |  |
|  | - Configuration programming for automatic sample analysis | - YES |  |  | - NO |  |  |  |  |
|  | - Sample illumination using ring illumination or polarized light illumination | - YES |  |  | - NO |  |  |  |  |
|  | - Controller unit for the microscope | - YES |  |  | - NO |  |  |  |  |
|  | Microscope Control PC: |  |  |  |  |  |  |  |  |  |
|  | - Dedicated PC with preinstalled Microsoft Windows 10 Professional 64-bit or equivalent or newer | - YES |  |  | - NO |  |  |  |  |
|  | - Equal or better than 23” HD Monitor | - YES |  |  | - NO |  |  |  |  |
|  | - All necessary software for microscope operation and control preinstalled on the dedicated PC | - YES |  |  | - NO |  |  |  |  |
|  | - All necessary software for the analysis of captured images pre-installed on the dedicated PC | - YES |  |  | - NO |  |  |  |  |
|  | - Support and maintenance: - Free software support (for accompanying licenses) for at least ten (10) years for the offered and supplied dedicated PC configuration. - In the event of failure of the supplied dedicated computer after the expiration of the warranty period, the provider will provide free support in configuring and installing the software on a replacement dedicated computer, which will be provided by the customer according to the provider's specifications. If the upgrade of a replacement dedicated computer will require the replacement of hardware for communication with the microscope, the replacement communication component's price may not exceed the value of EUR 1,000. Any failure of a dedicated computer must not significantly cost the customer for over ten years. | - YES |  |  | - NO |  |  |  |  |
|  | Special requirements for the item Microscope Control PC, which ensure the fulfilment of the objectives in accordance with the Decree on green public procurement (Official Journal of the Republic of Slovenia RS, Nos. 51/17 and 64/19): (personal computer) |  |  |  |  |  |  |  |  |
|  | 1. Computer complies with the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. | - YES |  |  | - NO |  |  |  |  |
|  | 2. Computer is designed in such a way that the hard disk and memory are easily accessible using universally available tools (i.e. with a screwdriver, spatula, pliers, or tweezers) and are replaceable. | - YES |  |  | - NO |  |  |  |  |
|  | 3. The volume of the computer is measured in accordance with the SIST EN ISO 7779 standard and does not exceed 26 dB at the time of writing to the hard disk. The data must be labelled with the designation LpAm in dB. | - YES |  |  | - NO |  |  |  |  |
|  | 4. Screen meets the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. | - YES |  |  | - NO |  |  |  |  |
|  | Image processing and acquisition: |  |  |  |  |  |  |  |  |  |
|  | - Option for applying color texture from a 2D image to 3D measurement results (combining of laser scan with white light image) | - YES |  |  | - NO |  |  |  |  |
|  | - Various 2D and 3D measurements of the acquired images, such as thickness, cross section measurements, height difference, etc. | - YES |  |  | - NO |  |  |  |  |
|  | - Measurement of volume and step height of the sample | - YES |  |  | - NO |  |  |  |  |
|  | - Color reconstruction of 3D scanned samples | - YES |  |  | - NO |  |  |  |  |
|  | - Tilt compensation of tilted samples | - YES |  |  | - NO |  |  |  |  |
|  | - Cross section analysis at any location of the acquired image | - YES |  |  | - NO |  |  |  |  |
|  | - Automatic analysis of different samples (sample comparison) | - YES |  |  | - NO |  |  |  |  |
|  | - Analysis of spherical and cylindrical structures such as optical fibers | - YES |  |  | - NO |  |  |  |  |
|  | - Thin film thickness measurements | - YES |  |  | - NO |  |  |  |  |
|  | - Analysis of the roughness of the observed sample surfaces | - YES |  |  | - NO |  |  |  |  |
|  | - HDR (High Dynamic Range image) function | - YES |  |  | - NO |  |  |  |  |
|  | - Depth composition of the image | - YES |  |  | - NO |  |  |  |  |
|  | - Configuration for automatic measurement of the acquired image | - YES |  |  | - NO |  |  |  |  |

| Place and date: | Seal: | Authorized signature: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification. The tenderer shall attach to the tender the proof of fulfilling all technical requirements and in case the tenderer provides a link to the manufacturer's website, he must ensure their accuracy and functionality (active connection) in the Word format.

For the item Microscope Control PC the tenderer must provide proof which shows fulfilment of the objectives in accordance with the Decree on green public procurement (Official Journal of the Republic of Slovenia RS, Nos. 51/17 and 64/19), as written in Technical specification, which are a part of the documentation of this public tender.