- CORRECTION

Project name: Upgrading national research infrastructures – RIUM

Technical specifications for Surface Analytical Equipment for Spectroelectrochemical Studies and Inert Atmosphere Chamber - Glovebox

- SUBJECT OF PUBLIC TENDER

LOT 1: Time-of-Flight Secondary Ion Mass Spectrometry, TOF-SIMS

Secondary ion mass spectroscopy is a surface analysis technique used to measure the mass spectrum on the surface of solids. The spectrometer enables high-resolution analysis of elementary compositions and the molecular structure of surfaces, thin layers, and interphases of various materials (ceramics, semiconductors, polymers, biomaterials, coatings, etc.). TOF-SIMS is state-of-the-art equipment that enables high mass resolution, high spatial resolution, depth profiling with dual ion sources, for both conductive and insulating materials. The use of different ionic sources allows the analysis of both inorganic and organic materials (use of clusters) by acquiring high mass resolutions also for lighter elements such as nitrogen, hydrogen, and oxygen, and for acquiring high lateral resolution (<60 nm) of microanalysis in imaging.

LOT 2: Atomic force microscope, AFM

An atomic force microscope (AFM) is equipment for studying topography at the atomic resolution and interactions studies. It will be used for the analysis of new advanced materials for catalysts. It is suitable for the study of both conductive and non-conductive materials. The equipment will have the possibility of coupling with a potentiostat / galvanostat, which will enable the simultaneous extraction of electrochemical measurements while monitoring the topographic and morphological properties of the sample on the atomic scale. The equipment will also be used to determine topography on a larger scale for the analysis of new advanced materials, sensors, and components in microfluidics.

LOT 3: Inert atmosphere chamber - glovebox

Possibility of processing optoelectronic materials in oxygen-free and water-free conditions or to process anaerobic biological materials in sterile, oxygen-free conditions with maximum safety for the user. Possibility to perform measurements with various portable equipment in dry or sterile conditions.

- TECHNICAL REQUIEREMENTS

Delivery deadline:

LOT 1: The supplier must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 1.18 of the technical specifications at the location University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, no later than 240 days from the entry into force of the contract.

LOT 2: The supplier must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 2.23 of the technical specifications at the location University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, no later than 120 days from the entry into force of the contract.

LOT 3: The supplier must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 3.11 of the technical specifications at the location University of Maribor, Faculty of Chemistry and Chemical Technology, Smetanova ulica 17, 2000 Maribor, no later than 60 days from the entry into force of the contract.

Warranty period:

LOT 1: For the delivered equipment, the supplier provides a 3 (three) years warranty for a flawless technical operation.

LOT 2: For the supplied equipment, the supplier provides a 2 (two) years warranty for a flawless technical operation.

LOT 3: For the supplied equipment, the supplier provides a 3 (three) year warranty for a flawless technical operation.

LOT 1: Time-of-Flight Secondary Ion Mass Spectrometry, TOF-SIMS

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 1.1 | 1 | The system must include ion source producing Bin primary ions with the following specifications: - the source must provide large Bin clusters up to n=7 and high current of Bi1, - total DC current should exceed 40 nA, in pulsed mode (including bunching) a target current of 40 pA has to be obtained, - lateral resolution: less than 50 nm (with Bi clusters and Bi ions), - mass resolution: more than 18 000 at 29 atomic mass units (with Bi1 and Bi3), - operation in different modes: high-mass resolution and small spot size at the same time, i.e. a mass resolution of 9 000 at atomic mass units >50 atomic mass units and a spot size of <90 nm must be demonstrated simultaneously, - fast imaging of the features of size 90 nm or less, An internal Faraday cup must be included in the system. The time-of-flight (TOF) analyzer with the following characteristics must be included: - high secondary ion transmission and high mass resolution simultaneously (transmission higher than > 8.0 x 108 Al+/nC, mass resolution higher than 7 000), - parallel detection of all masses of one secondary ion polarity, - mass range: > 12 000 atomic mass units, - mass resolution at lower mass (29 atomic mass units) must exceed 18 000, and at higher mass (> 200 atomic mass units) a resolution of 26 000 or more has to be demonstrated, - the TOF analyzer including the secondary ion detector with data registration system must be capable to linearly detect and process effective count rates of more than 107 counts per second for a given secondary ion species. At the same time, a minimum count rate of <0.1 counts per second has to be linearly detected without any changes in primary ion current or analyzer settings, - in-column measurement of mass separated, pulsed primary ion currents down to 50 fA - a pumping system for an ultra-high vacuum with a demonstrated base pressure of less than 5 x 10–10 torr, - the noise exposure caused by the instrument should be less than 54 dB. A time-of-flight MS/MS package (tandem MS) for the TOF analyser. The option should allow for the diversion of individual precursor molecules into a high energy collision cell (collision-induced dissociation, CID). Precursor selection should be performed with a mass resolution of 3 000 or higher, so that different molecular species can be individually selected from the same integer mass. The time-of-flight MS/MS package should include all hardware and software for high-speed and mass-selected precursor selection. All ion sources included in items 1.1, 1.2 and 1.3 have to be installed on the main analytical chamber together with the TOF analyzer. Any sample treatment by the ion sources must be accomplished without movement of the sample under analysis. |
| 1.2 | 1 | System for (inorganic) sample analysis must include: - dual-source column with an electron impact gas ion source (for operation with oxygen and Ar, Ne, and Xe gases) and a thermal ionization cesium ion source for dual beam depth profiling in combination with the Bi ion cluster gun, with minimum acceleration energy of 100 eV or lower, - a current density of >48 mA/cm2 and >12 mA/cm2 for the oxygen and cesium sputter beams, respectively, - minimum sputter beam diameters: 10 μm for oxygen and 5 μm for cesium, - hardware and software for high-speed and mass-selected secondary ion imaging, - titanium sublimation pump, which must provide automated out-gassing of the filaments and switching between the Ti filaments, smart controller for ultra-high vacuum, - special sample holder and power supply for in-situ, high-speed sample rotation while the sample is under analysis. |
| 1.3 | 1 | System for organic samples analysis must include: - argon gas cluster ion source, - oxygen operation option for gas cluster source, - analysis option for the gas cluster source including additional pulsing systems and software, - operation in the dual-beam mode of depth profiling in combination with the liquid metal cluster ion gun as well as operation as a short pulse length analysis gun (possible operation also in the single-beam mode of depth profiling), - spot size: less than 50 μm (with an Ar cluster current of 10 nA or more), - minimum beam diameter: 5 μm or less. - minimum pulse length (using analysis option): < 20 ns |
| 1.4 | 1 | Internal bakeout of the main chamber including differential pumping of the TOF analyzer. |
| 1.5 | 1 | Internal bakeout of the sample introduction chamber. |
| 1.6 | 1 | High-resolution photo navigation system (reflex camera with 24.0 Megapixel or higher, integrated sample illumination, USB interface). |
| 1.7 | 1 | Motorized five-axes sample stage with travel in x, y, z directions, tilting and rotation. |
| 1.8 | 1 | Fully automated gas flooding system, including a regulative automated leak valve and nozzle, which allow the exposure of the sample to inert and reactive gases. The fully automated switching of gases by the vacuum control unit must be possible and included. |
| 1.9 | 1 | Licensed software for measuring and processing data for at least 8 users. |
| 1.10 | 1 | A secondary electron (SE) detector. The detection should allow the acquisition of SE images using different ion guns with and without analyzer extraction voltage. It should enable the user to set up operation parameters such as focus and field of view for both secondary ion polarities using SE images. |
| 1.11 | 1 | Computer-controlled heating and cooling of the sample in the sample introduction chamber and in the main analytical chamber while the sample is under analysis, with the following specifications: - possibility to cool the sample to at least −150 °C and the possibility to heat the sample to at least 600 °C, where full temperature range (from at least −150 °C to at least 600 °C) must be accessible on the same sample holder, - cooling from room temperature to −130 °C must be possible within less than 15 minutes, - additional dedicated sample holder must be included, with the possibility of cooling and heating and full movement of the sample in X, Y and Z directions, rotation and tilting of the sample. |
| 1.12 | 1 | A self-adjusting charge compensation system with the following specifications: - the analysis of insulating samples must be possible without special pre-treatment of the sample and without special sample mounting, - a low-energy electron flood gun with adjustable energy (1 to 20 eV) shall be used for charge compensation, - the charge compensation must work with a maximum repetition rate of up to 50 kHz independent of sample size and geometry, - the sample must be permanently held at ground potential throughout the complete measurement. |
| 1.13 | 1 | PC with specifications: - MS Windows 10 Professional 64 bit operating system, - Intel® Core® i9 3.6 GHz processor or equivalent, - HP Z2 Tower G4 Workstation or equivalent, - 8 Gb RAM, DVD writer, 1 Tb hard disk, 2 x 27” LED monitors with Full HD resolution or with higher resolution than Full HD. |
| 1.14 | 1 | Warranty: 36 months or more must cover: - the replacement costs of the defective parts, - the corresponding labor expenses together with travel and living expenses of the provider’s personnel. All parts and services necessary to maintain optimum condition and a high availability of the instrument should be provided. Everything that is necessary for successful surface analysis operation should be covered at no additional cost throughout the term of the warranty. This includes: - all spare parts needed, - all consumables needed for 1200 hours of operation per year or more including its exchange on-site performed by the provider’s expert, - one preventative maintenance visit per year (two working days on-site), - unlimited service visits, - free software updates. |
| 1.15 | 1 | Basic operation training Acquaintance with the equipment and basic training, performed by the engineer of the provider, for at least four users lasting at least 5 working days (also includes the submission of all necessary data and passwords, the cost of travel and accommodation of the provider’s engineer; training is carried out at the location of FKKT UM, determined by the client). The item is completed after the delivery, installation and startup of the equipment. |
| 1.16 | 1 | Two additional advanced training for a period of at least 3 working days each at the provider’s site, for at least 2 persons in April 2022 (3 working days, first advanced training) and April 2023 (3 working days, second advanced training), by a specialist of the provider. Additional advanced training is intended to educate users on new and advanced methods of using the instrument. |
| 1.17 | 1 | Provided remote access to the device for the needs of calibration and control of the device at least for the duration of the warranty period (3 years). The instrument design must facilitate remote operation and fault diagnostics. |
| 1.18 | 1 | Delivery and installation on the client’s premises (DAP FKKT UM, Maribor). Condition for the successful execution of the item and signing of the acceptance report: - the supply, installation and commissioning of the equipment on the client’s premises by the provider, - introduction to equipment and basic education referred to item 1.15, - submission of operating instructions and other technical documentation - all performance specifications (items from 1.1 to 1.13) must be demonstrated on-site (FKKT UM) before signing technical acceptance documents. |

Lot 2: Atomic force microscope, AFM

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 2.1 | 1 | The system must include the head of the instrument with a manual movement of Z-scanner with the following requirements: - minimum scan range of 15 µm or more in both closed-loop and open-loop operation, - the noise of the sensor less than 250 pm (Adev, bandwidth from 0.1 Hz to 1 kHz), - the scanner must be compatible with all supplied scan modes (item 2.5), in both air and liquid environments, - the cantilever holder must be compatible for operation in air and in liquid. |
| 2.2 | 1 | Piezo driven flexure stage XY-scanner with the following requirements: - scan range of more than 100 µm (operation in closed-loop), - the noise of the sensor less than 0.6 nm (Adev, bandwidth from 0.1 Hz to 1 kHz), - operation in open-loop, closed-loop, and hybrid mode, - possibility to accommodate sample sizes up to 80 mm in diameter or more, - possibility to accommodate sample sizes up to 15 mm or more in height without any modifications to the standard configuration. |
| 2.3 | 1 | An additional possibility to scan and analyze the sample with a height of 25 mm or more. |
| 2.4 | 1 | Illumination and camera system must include: - top-down optical view with the possibility to switch the field of view between 720 and 240 µm, - integrated sample positioning stage that must enable precise alignment of the probe over the desired region of analysis, - illuminator: LED fiber light, - integrated optical system for viewing the probe and sample both from above (top-down) and from below (bottom-view) the sample, - the infrared light source for the deflection measurements (wavelengths in the visible region are not acceptable). |
| 2.5 | 1 | The full complement of operating modes (for fluid and air operation), including: - contact mode, - tapping mode, - tapping mode with phase control and Q-control, - bimodal Dual AC mode, - electric force microscopy, - electrostatic force microscopy, - Kelvin probe microscopy, - surface potential microscopy, - magnetic force microscopy, - piezoresponse force microscopy, - dual AC mode, - dual AC Resonance Tracking, - contact resonance viscoelastic mapping mode, - loss tangent imaging, - force spectroscopy and force mapping, - phase imaging, - frequency modulation, - imaging in fluid (with the same setup as air imaging), - nanolithography, - nanomanipulation. |
| 2.6 | 1 | Licensed software for measuring and processing data on a minimum of 3 computers. |
| 2.7 | 1 | The system must include a single probe holder that supports all of the included operating modes and supports operation in both air and liquid for those modes. |
| 2.8 | 1 | Two additional sample holders: - holder for standard magnetic sample discs, - generic holder with spring clips. |
| 2.9 | 1 | Automatic calibration of the cantilever sensitivity (deflection sensitivity) and spring constant by selecting the probe type. The feature must actually calibrate the probe and not use nominal tabulated values for the sensitivity and spring constant. During the calibration, the tip should not touch the sample. |
| 2.10 | 1 | The beam used for deflection detection must approach the probe at an angle that is significantly (>20 degrees) off vertical relative to the sample. |
| 2.11 | 1 | Isolation enclosure must require: - acoustic and vibration isolation (without the need for compressed air), - the whole system (with optics and without the control system) must be contained within this enclosure, - the table for the microscope must be included. |
| 2.12 | 1 | Additional cell set for dedicated electrochemical experiments should be included with the following: - fluid cell assembly made of PEEK or similar inert nonmetallic material, - cantilever holder made of PEEK or similar inert nonmetallic material, for use with a customer supplied potentiostat, - working electrode sample mount suitable for both conducting and insulating sample substrates, - carbon (graphite) counter electrode - sealed Ag/AgCl reference electrode, - electrical junction box with connections and programmable jumpers to interface with a wide range of commercial potentiostats, - spare O-rings and assembly hardware, - additional sealing kit for an electrochemical cell, which provides operation of the cell in a fully sealed configuration. |
| 2.13 | 1 | The system must offer measurements in the conductive mode for current imaging and I/V spectroscopy capabilities at user-selectable voltages and points over a broad operating range. This package should include: - cantilever holder with an integral dual gain preamplifier for wide range current measurements at constant applied voltage. - two separate current output channels provide gains of 1x106 (1 µA/V) and 1x109 (1 nA/V), (4 pA to 10 µA), - conductive sample holder mount with sample bias contact, - additional probes for conductive mode (20 pieces or more). - silver paint, - tools and accessories, - the system must include current measurements from < 1pA to 10 μA. The full dynamic range of current sensing must be recorded simultaneously. |
| 2.14 | 1 | Extended head in Z-range that provides imagining of tall samples and long-range force measurements, in the range up to 40 µm or more. |
| 2.15 | In total 30 pieces or more | A standard pack of cantilevers or probes must include: - 10 x silicon probes with visible apex tip for tapping mode, f: 300 kHz, k: 26 N/m, no tip coating, - 10 x medium soft cantilevers for topography and viscoelastic properties of soft samples, f = 70 kHz, k= 2 N/m, no tip coating, - 10 x probes with triangular cantilevers (material: Pyrex-Nitride), gold coating. |
| 2.16 | In total 20 or more | Additional cantilevers or probes for tapping mode, operation in air, for soft/medium stiff samples, k = 2 N/m (at least 20 pieces or more). |
| 2.17 | In total 20 or more | Additional cantilevers of triangular shape (material: Pyrex-Nitride), for imaging in the liquid, gold coating (at least 20 pieces or more). |
| 2.18 | 1 | The additional accessory kit should include: - tweezers, - calibration grating, - power strip and - all necessary cords and cables. |
| 2.19 | 1 | PC specifications: - Dell Precision T5810 workstation or equivalent, - 3.5 GHz Intel four core Haswell Xeon or equivalent, - 8 GB DDR4 RAM, 2 x hard drives 2 TB, 1 x DVD-ROM, - AMD Radeon RX-470 4 GB video card or equivalent, - Hauppauge ImpactVCB-e capture card or equivalent, - 2 x 32” LCD screens, with a resolution of at least 3440 × 1440 or more, - 6 x USB 2.0 ports or more, 4 x USB 3.0 ports or better, - Microsoft Windows 10 Professional 64 bit. |
| 2.20 | 1 | Basic operation training Acquaintance with the equipment and basic training, performed by the engineer of the provider, for 3 users lasting at least 3 working days (also includes the submission of all necessary data and passwords, the cost of travel and accommodation of the engineer; training is carried out at the location FKKT UM, determined by the client). The item is completed after the delivery, installation and startup of the equipment. |
| 2.21 | 1 | Warranty: 24 months or more. |
| 2.22 | 1 | Delivery and installation on the client’s premises. Condition for the successful execution of the item and signing of the acceptance report: - the supply, installation and commissioning of the equipment by the provider, - introduction to equipment and basic education referred to item 2.20, - submission of operating instructions and other technical documentation, - all performance specifications (items from 2.1 to 2.18) must be demonstrated on-site (FKKT UM) before signing technical acceptance documents. |
| 2.23 | 1 | Provided remote access to the device for the needs of calibration and control of the device at least for the duration of the warranty period (2 years). The instrument design must facilitate remote operation and fault diagnostics. |

Lot 3: Inert atmosphere chamber – glovebox

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 3.1 | 1 | Inert atmosphere chamber made of stainless steel type 304 L or 304 or type of material that has better corrosion resistance: - vacuum antechamber for pre-cleaning the atmosphere with manual control with size of 400 mm or 390 mm in diameter and a length of 600 mm, - vacuum antechamber for pre-cleaning the atmosphere with manual control with size of 150 mm in diameter and a length of 400 mm with a movable drawer for sample transfer between ambient and regulated atmosphere - at least 1 safety valve, - at least 3 valves for gas input/output to and from inert atmosphere chamber (valve for input/output into and from atmosphere chamber, valve for input/output into and from large antechamber, valve for input/output into and from smaller antechamber), all three chambers must have capability of gas purging - at least 1 x 220 V / 16 A electrical socket, - at least 2 x USB ports for connection between external and internal equipment, compatible with standard USB 3.0. A, with female connector mounted inside atmosphere chamber and male connector or extension cord with minimum length of 1 m on outside of atmosphere chamber. - at least 4 electrical interfaces with width of 5 mm for connecting an external potentiostat / galvanostat, type: standard banana connector – 4 mm, there can be female connectors on both sides of the chamber - LED lighting of the work surface inside the inert atmosphere chamber - 3 or more shelves made of the same type of material as the inert atmosphere chamber, - at least 2 glove adjustments lines, each with position for single glove that enables a single person to work with both left and right hand, in such way, that a single person can work with both hands, where each individual glove adjustment line represents place for single glove. The first glove adjustment line is for left hand and second adjustment line is for right hand. - at least 1 HEPA filter for dust and small particle filtration. |
| 3.2 | 1 | Glove from Butyl or equivalent, 0,4-0,5 mm thick, left glove at least 500 mm or 0,5 m long. |
| 3.3 | 1 | Glove from Butyl or equivalent, 0,4-0,5 mm thick, right glove at least 500 mm or 0,5 m long. |
| 3.4 | 1 | System for automatic control of gas supply, chamber pressure and atmosphere purging with filter capacity of at least 3,5 kg for copper catalyst with ability to remove oxygen up to 1 ppm or equal adsorbent for oxygen gas removal and molecular sieves or equivalent adsorbent for water removal with filter mass capacity of 3,5 kg with reduction capacity of up to 1 ppm. |
| 3.5 | 1 | Oxygen analyzer capable of measuring in the range not exceeding 1 ppm and not less than 1000 ppm with a precision of no less than ± 1 ppm. |
| 3.6 | 1 | Water analyzer capable of measuring in the range not exceeding 1 ppm and not less than 500 ppm with a precision of no less than ± 1 ppm. |
| 3.7 | 1 | Vacuum pump with a minimum suction of 17 m3/h. |
| 3.8 | 1 | At least 1 HEPA filter. |
| 3.9 | 1 | Oxygen adsorbent, at least 3.5 kg. |
| 3.10 | 1 | Water adsorbent, at least 3.5 kg. |
| 3.11 | 1 | Delivery and installation in the client's premises. Condition for successful execution of the item and signing the acceptance report: - the supply and installation and commissioning of the equipment by the supplier, - submission of operating instructions and other technical documentation. |
| 3.12 | 1 | Warranty: At least 3 years or more. |

The operation is co-financed by the Republic of Slovenia, the Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is implemented under the Operational Program for the Implementation of European Cohesion Policy 2014-2020 under Priority Axis 1: "International competitiveness of research, innovation and technological development in line with smart specialization for greater competitiveness and greening of the economy", priority investment 1.1 "Strengthening infrastructure for research and innovation and capacity to develop excellence in this field, as well as the promotion of competent centers, especially those of European importance ", specific objective 1.1.1" Effective use of research infrastructure and development of knowledge/competences for better national and international cooperation in the triangle knowledge ”.