Name of the project: Upgrading national research infrastructures – RIUM

COMPLIANCE DECLARATION WITH MINIMUM TECHNICAL REQUIREMENTS OF THE OFFERED EQUIPMENT

Name of the tenderer:

Address of the tenderer:

Subject-matter of the public contract: Purchase of the Surface Analytical Equipment for Spectroelectrochemical Studies

Number of the public contract: 302/6 - RIUM32-FKKT02-1/2021

# INSTRUCTION:

- This Attachment No. 2 must be completed and signed. The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all technical requirements.

- The tenderer shall upload this Attachment No. 2, in Word and PDF format (including attached evidence of compliance with the technical requirements), to the information system e-JN under the section “Other Attachments”.

We declare that the offered equipment fully meets the minimum technical requirements of the contracting authority for the subject-matter of the public contract titled Purchase of the Surface Analytical Equipment for Spectroelectrochemical Studies and that we are ready to realize the public contract according to the following specification:

| EQUIPMENT |  |  |  |  |
| --- | --- | --- | --- | --- |
| Item number | No. of pieces | Required | Offered | Evidence |
| Time-of-Flight Secondary Ion Mass Spectrometry, TOF-SIMS with the following specifications: |  |  |  |  |
| 1.1 | 1 | The system must include ion source producing Bin primary ions with the following specifications: - the source must provide large Bin clusters up to n=7 and high current of Bi1, - total DC current should exceed 40 nA, in pulsed mode (including bunching) a target current of 40 pA has to be obtained, - lateral resolution: less than 50 nm (with Bi clusters and Bi ions), - mass resolution: more than 18 000 at 29 atomic mass units (with Bi1 and Bi3), - operation in different modes: high-mass resolution and small spot size at the same time, i.e. a mass resolution of 9 000 at atomic mass units >50 atomic mass units and a spot size of <90 nm must be demonstrated simultaneously, - fast imaging of the features of size 90 nm or less, An internal Faraday cup must be included in the system. The time-of-flight (TOF) analyzer with the following characteristics must be included: - high secondary ion transmission and high mass resolution simultaneously (transmission higher than > 8.0 x 108 Al+/nC, mass resolution higher than 7 000), - parallel detection of all masses of one secondary ion polarity, - mass range: > 12 000 atomic mass units, - mass resolution at lower mass (29 atomic mass units) must exceed 18 000, and at higher mass (> 200 atomic mass units) a resolution of 26 000 or more has to be demonstrated, - the TOF analyzer including the secondary ion detector with data registration system must be capable to linearly detect and process effective count rates of more than 107 counts per second for a given secondary ion species. At the same time, a minimum count rate of <0.1 counts per second has to be linearly detected without any changes in primary ion current or analyzer settings, - in-column measurement of mass separated, pulsed primary ion currents down to 50 fA - a pumping system for an ultra-high vacuum with a demonstrated base pressure of less than 5 x 10–10 torr, - the noise exposure caused by the instrument should be less than 54 dB. A time-of-flight MS/MS package (tandem MS) for the TOF analyser. The option should allow for the diversion of individual precursor molecules into a high energy collision cell (collision-induced dissociation, CID). Precursor selection should be performed with a mass resolution of 3 000 or higher, so that different molecular species can be individually selected from the same integer mass. The time-of-flight MS/MS package should include all hardware and software for high-speed and mass-selected precursor selection. All ion sources included in items 1.1, 1.2 and 1.3 have to be installed on the main analytical chamber together with the TOF analyzer. Any sample treatment by the ion sources must be accomplished without movement of the sample under analysis. |  |  |
| 1.2 | 1 | System for (inorganic) sample analysis must include: - dual-source column with an electron impact gas ion source (for operation with oxygen and Ar, Ne, and Xe gases) and a thermal ionization cesium ion source for dual beam depth profiling in combination with the Bi ion cluster gun, with minimum acceleration energy of 100 eV or lower, - a current density of >48 mA/cm2 and >12 mA/cm2 for the oxygen and cesium sputter beams, respectively, - minimum sputter beam diameters: 10 μm for oxygen and 5 μm for cesium, - hardware and software for high-speed and mass-selected secondary ion imaging, - titanium sublimation pump, which must provide automated out-gassing of the filaments and switching between the Ti filaments, smart controller for ultra-high vacuum, - special sample holder and power supply for in-situ, high-speed sample rotation while the sample is under analysis. |  |  |
| 1.3 | 1 | System for organic samples analysis must include: - argon gas cluster ion source, - oxygen operation option for gas cluster source, - analysis option for the gas cluster source including additional pulsing systems and software, - operation in the dual-beam mode of depth profiling in combination with the liquid metal cluster ion gun as well as operation as a short pulse length analysis gun (possible operation also in the single-beam mode of depth profiling), - spot size: less than 50 μm (with an Ar cluster current of 10 nA or more), - minimum beam diameter: 5 μm or less. - minimum pulse length (using analysis option): < 20 ns |  |  |
| 1.4 | 1 | Internal bakeout of the main chamber including differential pumping of the TOF analyzer. |  |  |
| 1.5 | 1 | Internal bakeout of the sample introduction chamber. |  |  |
| 1.6 | 1 | High-resolution photo navigation system (reflex camera with 24.0 Megapixel or higher, integrated sample illumination, USB interface). |  |  |
| 1.7 | 1 | Motorized five-axes sample stage with travel in x, y, z directions, tilting and rotation. |  |  |
| 1.8 | 1 | Fully automated gas flooding system, including a regulative automated leak valve and nozzle, which allow the exposure of the sample to inert and reactive gases. The fully automated switching of gases by the vacuum control unit must be possible and included. |  |  |
| 1.9 | 1 | Licensed software for measuring and processing data for at least 8 users. |  |  |
| 1.10 | 1 | A secondary electron (SE) detector. The detection should allow the acquisition of SE images using different ion guns with and without analyzer extraction voltage. It should enable the user to set up operation parameters such as focus and field of view for both secondary ion polarities using SE images. |  |  |
| 1.11 | 1 | Computer-controlled heating and cooling of the sample in the sample introduction chamber and in the main analytical chamber while the sample is under analysis, with the following specifications: - possibility to cool the sample to at least −150 °C and the possibility to heat the sample to at least 600 °C, where full temperature range (from at least −150 °C to at least 600 °C) must be accessible on the same sample holder, - cooling from room temperature to −130 °C must be possible within less than 15 minutes, - additional dedicated sample holder must be included, with the possibility of cooling and heating and full movement of the sample in X, Y and Z directions, rotation and tilting of the sample. |  |  |
| 1.12 | 1 | A self-adjusting charge compensation system with the following specifications: - the analysis of insulating samples must be possible without special pre-treatment of the sample and without special sample mounting, - a low-energy electron flood gun with adjustable energy (1 to 20 eV) shall be used for charge compensation, - the charge compensation must work with a maximum repetition rate of up to 50 kHz independent of sample size and geometry, - the sample must be permanently held at ground potential throughout the complete measurement. |  |  |
| 1.13 | 1 | PC with specifications: - MS Windows 10 Professional 64 bit operating system, - Intel® Core® i9 3.6 GHz processor or equivalent, - HP Z2 Tower G4 Workstation or equivalent, - 8 Gb RAM, DVD writer, 1 Tb hard disk, 2 x 27” LED monitors with Full HD resolution or with higher resolution than Full HD. |  |  |
| 1.14 | 1 | Warranty: 36 months or more must cover: - the replacement costs of the defective parts, - the corresponding labor expenses together with travel and living expenses of the provider’s personnel. All parts and services necessary to maintain optimum condition and a high availability of the instrument should be provided. Everything that is necessary for successful surface analysis operation should be covered at no additional cost throughout the term of the warranty. This includes: - all spare parts needed, - all consumables needed for 1200 hours of operation per year or more including its exchange on-site performed by the provider’s expert, - one preventative maintenance visit per year (two working days on-site), - unlimited service visits, - free software updates. |  |  |
| 1.15 | 1 | Basic operation training Acquaintance with the equipment and basic training, performed by the engineer of the provider, for at least four users lasting at least 5 working days (also includes the submission of all necessary data and passwords, the cost of travel and accommodation of the provider’s engineer; training is carried out at the location of FKKT UM, determined by the client). The item is completed after the delivery, installation and startup of the equipment. |  |  |
| 1.16 | 1 | Provided remote access to the device for the needs of calibration and control of the device at least for the duration of the warranty period (3 years). The instrument design must facilitate remote operation and fault diagnostics. |  |  |
| 1.17 | 1 | Delivery and installation on the client’s premises (DAP FKKT UM, Maribor). Condition for the successful execution of the item and signing of the acceptance report: - the supply, installation and commissioning of the equipment on the client’s premises by the provider, - introduction to equipment and basic education referred to item 1.15, - submission of operating instructions and other technical documentation - all performance specifications (items from 1.1 to 1.13) must be demonstrated on-site (FKKT UM) before signing technical acceptance documents. |  |  |

| Place and date: | Seal: | Signature of person responsible: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and page number(s) together with the relevant technical specification. Exceptionally, the tenderer may also provide a link to the manufacturer’s website with the relevant technical specification. Where a link to the online document is provided, the tenderer must indicate a page number indicating the relevant technical specification. The tenderer must also enclose evidence of compliance with the minimum technical requirements and, in the case that the tenderer provides links to the website, ensure their relevance and activity in Word format.