Name of the project: Upgrading national research infrastructures – RIUM

COMPLIANCE DECLARATION WITH MINIMUM TECHNICAL REQUIREMENTS OF THE OFFERED EQUIPMENT

Name of the tenderer:

Address of the tenderer:

Subject-matter of the public contract: Purchase of the Atomic force microscope, AFM and Inert Atmosphere Chamber - Glovebox

Ime sklopa: LOT 2:  Inert atmosphere chamber – glovebox

Number of the public contract: 302/7 – RIUM32-FKKT02-1/2021

# INSTRUCTION:

- This Attachment No. 2 must be completed and signed. The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all technical requirements.

- The tenderer shall upload this Attachment No. 2 to the information system e-JN under the section “Other Attachments”.

We declare that the offered equipment fully meets the minimum technical requirements of the contracting authority for the subject-matter of the public contract titled "Purchase of the Atomic force microscope, AFM and Inert Atmosphere Chamber - Glovebox", Lot 2: Inert atmosphere chamber – glovebox, and that we are ready to realize the public contract according to the following specification:

| EQUIPMENT |  |  |  |  |
| --- | --- | --- | --- | --- |
| Item number | No. of pieces | Required | Offered | Evidence |
| Inert atmosphere chamber – glovebox with the following specifications: |  |  |  |  |
| 2.1 | 1 | Inert atmosphere chamber made of stainless steel type 304 L or 304 or type of material that has better corrosion resistance: - vacuum antechamber for pre-cleaning the atmosphere with manual control with size of 400 mm or 390 mm in diameter and a length of 600 mm, - vacuum antechamber for pre-cleaning the atmosphere with manual control with size of 150 mm in diameter and a length of 400 mm with a movable drawer for sample transfer between ambient and regulated atmosphere - at least 1 safety valve, - at least 3 valves for gas input/output to and from inert atmosphere chamber, (valve for input/output into and from atmosphere chamber, valve for input/output into and from large antechamber, valve for input/output into and from smaller antechamber), all three chambers must have capability of gas purging - at least 1 x 220 V / 16 A electrical socket, - at least 2 x USB ports for connection between external and internal equipment, compatible with standard USB 3.0. A, with female connector mounted inside atmosphere chamber and male connector or extension cord with minimum length of 1 m on outside of atmosphere chamber - at least 4 electrical interfaces with width of 5 mm for connecting an external potentiostat / galvanostat, type: standard banana connector – 4 mm, there can be female connectors on both sides of the chamber - LED lighting of the work surface inside the inert atmosphere chamber - 3 or more shelves made of the same type of material as the inert atmosphere chamber, - at least 2 glove adjustments lines, each with position for single glove that enables a single person to work with both left and right hand, in such way, that a single person can work with both hands, where each individual glove adjustment line represents place for single glove. The first glove adjustment line is for left hand and second adjustment line is for right hand. - at least 1 HEPA filter for dust and small particle filtration. |  |  |
| 2.2 | 1 | Glove from Butyl or equivalent, 0,4-0,5 mm thick, left glove at least 500 mm or 0,5 m long. |  |  |
| 2.3 | 1 | Glove from Butyl or equivalent, 0,4-0,5 mm thick, right glove at least 500 mm or 0,5 m long. |  |  |
| 2.4 | 1 | System for automatic control of gas supply, chamber pressure and atmosphere purging with filter capacity of at least 3,5 kg for copper catalyst with ability to remove oxygen up to 1 ppm or equal adsorbent for oxygen gas removal and molecular sieves or equivalent adsorbent for water removal with filter mass capacity of 3,5 kg with reduction capacity of up to 1 ppm. |  |  |
| 2.5 | 1 | Oxygen analyzer capable of measuring in the range not exceeding 1 ppm and not less than 1000 ppm with a precision of no less than ± 1 ppm. |  |  |
| 2.6 | 1 | Water analyzer capable of measuring in the range not exceeding 1 ppm and not less than 500 ppm with a precision of no less than ± 1 ppm. |  |  |
| 2.7 | 1 | Vacuum pump with a minimum suction of 17 m3/h. |  |  |
| 2.8 | 1 | At least 1 HEPA filter. |  |  |
| 2.9 | 1 | Oxygen adsorbent, at least 3.5 kg. |  |  |
| 2.10 | 1 | Water adsorbent, at least 3.5 kg. |  |  |
| 2.11 | 1 | Delivery and installation in the client's premises. Condition for successful execution of the item and signing the acceptance report: - the supply and installation and commissioning of the equipment by the supplier, - submission of operating instructions and other technical documentation. |  |  |
| 2.12 | 1 | Warranty: At least 3 years or more. |  |  |

| Place and date: | Seal: | Signature of person responsible: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and page number(s) together with the relevant technical specification. Exceptionally, the tenderer may also provide a link to the manufacturer’s website with the relevant technical specification. Where a link to the online document is provided, the tenderer must indicate a page number indicating the relevant technical specification.