Project name: Upgrading national research infrastructures – RIUM

Declaration of compliance with the Contacting Authority’s minimum technical requirements and specifications of the offered equipment

Tenderer’s name:

Tenderer’s address:

Subject-matter of the public contract: Nabava mikroskopa na atomsko silo in komore za delo v inertni atmosferi (in Eng., Purchase of Atomic Force Microscope (AFM) and Inert Atmosphere Chamber - Glovebox

Lot name: LOT 1:  Mikroskop na atomsko silo (eng. Atomic Force Microscope, AFM)

Designation of the public contract: 302/7 – RIUM32-FKKT02-1/2021

INSTRUCTION:

- This Annex No. 2 has to be completed and signed. The tenderer has to specify the actual technical specifications of the offered equipment in compliance with all the requirements.

- The tenderer uploads this Annex No. 2 to the section “Other attachments” in the e-JN information system.

We declare that the offered equipment fully complies with the minimum technical requirements of the Contracting Authority for the subject matter of public procurement entitled Purchase of Atomic Force Microscope (AFM) and Inert Atmosphere Chamber – Glovebox for lot 1: Atomic Force Microscope (AFM), and that we are prepared to perform the contract according to the following specification:

| EQUIPMENT |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item number | Number of pieces | Required |  | Offered | Means of proof |  |
| Atomic force microscope (AFM) with the following characteristics: |  |  |  |  |  |  |
| 1.1 | 1 | The system must include the head of the instrument with a manual Z-scanner with the following requirements: - minimum scan range of 15 µm or more in both closed-loop and open-loop operation, - the noise of the sensor less than 250 pm (Adev, bandwidth from 0.1 Hz to 1 kHz), - the scanner must be compatible with all supplied scan modes (item 2.5), for operation in air or liquid, - the cantilever holder must be compatible for operation in air and in liquid. |  |  |  |  |
| 1.2 | 1 | Piezo driven flexure stage XY-scanner with the following requirements: - scan range of more than 100 µm (operation in closed-loop), - the noise of the sensor less than 0.6 nm (Adev, bandwidth from 0.1 Hz to 1 kHz), - operation in open-loop, closed-loop, and hybrid mode, - possibility to accommodate sample sizes up to 80 mm in diameter or more, - possibility to accommodate sample sizes up to 15 mm or more in height without any modifications to the standard configuration or restriction of the area of movement of the sample stage. |  |  |  |  |
| 1.3 | 1 | An additional possibility to scan and analyze samples with a height of 25 mm or more. |  |  |  |  |
| 1.4 | 1 | Illumination and camera system must include: - top-down optical view with the possibility to switch the field of view between 720 and 240 µm, - integrated sample positioning stage that must enable precise alignment of the probe over the desired region of analysis, - illuminator: LED fiber light, - integrated optical system for viewing the probe and sample both from above (top-down) and from below (bottom-view) the sample, - the infrared light source for the deflection measurements (wavelengths in the visible region are not acceptable). |  |  |  |  |
| 1.5 | 1 | The full complement of operating modes (for fluid and air operation), including: - contact mode, - tapping mode, - tapping mode with phase control and Q-control, - bimodal Dual AC-mode, - electric force microscopy, - electrostatic force microscopy, - Kelvin probe microscopy, - surface potential microscopy, - magnetic force microscopy, - piezoresponse force microscopy, - dual AC-mode, - dual AC- Resonance Tracking, - contact resonance viscoelastic mapping mode, - force spectroscopy and force mapping, - phase imaging, - frequency modulation, - imaging in fluid (with the same setup as air imaging), - nanolithography, - nanomanipulation. |  |  |  |  |
| 1.6 | 1 | Licensed software for measuring and processing data on a minimum of 3 computers. |  |  |  |  |
| 1.7 | 1 | The system must include a single probe holder that supports all of the included operating modes and supports operation in both air and liquid for those modes. |  |  |  |  |
| 1.8 | 1 | Two additional sample holders: - holder for standard magnetic sample discs, - generic holder with spring clips. |  |  |  |  |
| 1.9 | 1 | Automatic calibration of the cantilever sensitivity (deflection sensitivity) and spring constant by selecting the probe type. The feature must actually calibrate the probe and not use nominal tabulated values for the sensitivity and spring constant. During the calibration, the tip should not touch the sample. |  |  |  |  |
| 1.10 | 1 | The beam used for deflection detection must approach the probe at an angle that is significantly (>20 degrees) off vertical relative to the sample. |  |  |  |  |
| 1.11 | 1 | Isolation enclosure must enable: - acoustic and vibration isolation (without the need for compressed air), - the whole system (with optics and without the control system) must be contained within this enclosure, - the table for the microscope must be included. |  |  |  |  |
| 1.12 | 1 | Additional cell set for dedicated electrochemical experiments should be included with the following: - fluid cell assembly made of PEEK or similar inert nonmetallic material, - cantilever holder made of PEEK or similar inert nonmetallic material, for use with a customer supplied potentiostat, - working electrode sample mount suitable for both conducting and insulating sample substrates, - carbon (graphite) counter electrode, - sealed Ag/AgCl reference electrode, - electrical junction box with connections and programmable jumpers to interface with a wide range of commercial potentiostats, - spare O-rings and assembly hardware, - additional sealing kit for an electrochemical cell, which provides operation of the cell in a fully sealed configuration. |  |  |  |  |
| 1.13 | 1 | The system must offer measurements in the conductive mode for current imaging and I/V spectroscopy capabilities at user-selectable voltages and points over a broad operating range. This package should include: - cantilever holder with an integral dual gain preamplifier for wide range current measurements at constant applied voltage. - two separate current output channels provide gains of 1x106 (1 µA/V) and 1x109 (1 nA/V), (4 pA to 10 µA), - conductive sample holder mount with sample bias contact, - additional probes for conductive mode: 20 pieces or more, - silver paint, - tools and accessories, - the system must include current measurements from < 1pA to 10 μA. The full dynamic range of current sensing must be recorded simultaneously. |  |  |  |  |
| 1.14 | 1 | Extended head in Z-range that provides imagining of tall samples and long-range force measurements, in the range up to 40 µm or more. |  |  |  |  |
| 1.15 | In total 30 pieces or more | A standard pack of cantilevers or probes must include: - 10 x silicon probes with visible apex tip for tapping mode, f: 300 kHz, k: 26 N/m, no tip coating, - 10 x medium soft cantilevers for topography and viscoelastic properties of soft samples, f = 70 kHz, k= 2 N/m, no tip coating, - 10 x probes with triangular cantilevers (material: Pyrex-Nitride), gold coating. |  |  |  |  |
| 1.16 | In total 20 or more | Additional cantilevers or probes for tapping mode, operation in air, for soft/medium stiff samples, k = 2 N/m. |  |  |  |  |
| 1.17 | In total 20 or more | Additional cantilevers of triangular shape (material: Pyrex-Nitride), for imaging in the liquid, gold coating. |  |  |  |  |
| 1.18 | 1 | The additional accessory kit should include: - tweezers, - calibration grating, - power strip and - all necessary cables and accessories. |  |  |  |  |
| 1.19 | 1 | PC specifications: - Dell Precision T5810 workstation or equivalent, - 3.5 GHz Intel quad-core Haswell Xeon or equivalent, - 8 GB DDR4 RAM, 2 x hard drives 2 TB, 1 x DVD-ROM, - AMD Radeon RX-470 4 GB video card or equivalent, - Hauppauge ImpactVCB-e capture card or equivalent, - 2 x 32” LCD screens, with a resolution of at least 3440 × 1440 or more, - 6 x USB 2.0 ports or more, 4 x USB 3.0 ports or better, - Microsoft Windows 10 Professional 64 bit. |  |  |  |  |
| 1.20 | 1 | Basic operation training: Acquaintance with the equipment and basic training, performed by the engineer of the provider, for 3 users lasting at least 3 working days (also includes the submission of all necessary data and passwords, the cost of travel and accommodation of the engineer; training is carried out at the location FKKT UM, determined by the client). The item is completed after the delivery, installation and startup of the equipment. |  |  |  |  |
| 1.21 | 1 | Warranty: 24 months or more. |  |  |  |  |
| 1.22 | 1 | Delivery and installation in the Contracting Authority's premises. Condition for successful execution of the item and signing the handover record: - the supply, installation and start-up of the equipment by the provider, - introduction to equipment and basic education referred to item 2.20, - submission of operating instructions and other technical documentation, - all performance specifications (items from 2.1 to 2.18) must be demonstrated on-site (FKKT UM) before signing technical acceptance documents. |  |  |  |  |
| 1.23 | 1 | Provided remote access to the device for the needs of calibration and control of the device at least for the duration of the warranty period (2 years). The instrument design must facilitate remote operation and fault diagnostics. |  |  |  |  |
|  |  | Specific requirements: |  |  |  |  |
|  |  | - Personal computer meets the latest Energy Star energy efficiency standards applicable on the date of publication of the public procurement notice or on the date of the invitation to tender, or equivalent standards. |  |  |  |  |
|  |  | - Personal computer is designed in such a way that the hard drive and the memory are easily accessible using available universal tools (i.e., a screwdriver, spatula, pliers or tweezers) and replaceable. |  |  |  |  |
|  |  | - The volume of the PC has to be measured in accordance with the SIST EN ISO 7779 standard and must not exceed 26 dB when recording to the hard drive. The information has to be labelled with LpAm in dB. |  |  |  |  |
|  |  | - The monitor meets the latest Energy Star energy efficiency standards applicable on the date of publication of the public procurement notice or on the date of the invitation to tender. |  |  |  |  |

| Place and date: | Stamp: | Signature of responsible person: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1 of the Operational Program for Implementation of the European Cohesion Policy 2014-2020: “International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy”, priority investment 1.1 “Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance”, specific objective 1.1.1 “Efficient use of research infrastructure and development of knowledge/competences for better national and international cooperation in the knowledge triangle”.

The tenderer has to specify the actual technical specifications of the offered equipment in compliance with all the requirements.

The tenderer has to indicate the name of the accompanying equipment manufacturer’s catalogue or brochure and number of the page containing the relevant technical specifications. In exceptional cases, the tenderer may also provide a link to the website of the equipment manufacturer which includes an indication of the relevant technical specifications. Where a link to the online document is provided, the tenderer has to provide the number of the page containing the relevant technical specifications.

For item 1.19, the Contracting Authority requires the fulfilment of the objective set out in Point 7 of the second paragraph of Article 6 of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, Nos. 51/17, 64/19, 49/20 — ZIUZEEOP, 152/20 — ZZUOOP), hereinafter referred to as: Decree on GPP).

For item 1.19, PC specifications, the certificate or proof of compliance with the requirements of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17, 64/19, 49/20 – ZIUZEOP, 152/20 – ZZUOOP) is provided as proof.