Project name: Upgrading national research infrastructures – RIUM

Designation of the public contract: 302/7 – RIUM32-FKKT02-1/2021

Technical Specifications for the Atomic Force Microscope (AFM) and Inert Atmosphere Chamber - Glovebox

- SUBJECT-MATTER OF THE PUBLIC PROCUREMENT

Lot 1: Atomic force microscope (AFM)

Atomic force microscope (AFM), the equipment for studying the topography of the materials of atomic resolution and interactions. It should be applied to the analysis of new advanced materials for catalytic converters, it is suitable to study conducting and non-conducting materials. Equipment will have the possibility of coupling with a potentiostat/galvanostat, which will enable simultaneous implementation of electrochemical measurement while monitoring topographical and morphological properties of the sample at the atomic scale. Equipment will also be used to determine the topography on a larger scale for the analysis of new advanced materials, sensors and components in microfluidics.

Lot 2: Inert atmosphere chamber – glovebox

Possibility of processing optoelectronic materials in oxygen-free and water-free conditions or to process anaerobic biological materials in sterile, oxygen-free conditions with maximum safety for the user. Possibility to perform measurements with various portable equipment in dry or sterile conditions.

- TECHNICAL REQUIREMENTS

Delivery deadline:

LOT 1: The supplier must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 1.22 of the technical specifications at the location of the University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, no later than 120 days from the entry into force of the contract.

LOT 2: The supplier must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 2.11 of the technical specifications at the location of the University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, no later than 60 days from the entry into force of the contract.

Warranty period:

LOT 1: For the supplied equipment, the supplier provides a 2 (two) years warranty for a flawless technical operation.

LOT 2:  For the supplied equipment, the supplier provides a 3 (three) year warranty for a flawless technical operation.

Lot 1: Mikroskop na atomsko silo (eng. Atomic force microscope, AFM)

| Item number | Number of pieces | REQUIRED |
| --- | --- | --- |
| 1.1 | 1 | The system must include the head of the instrument with a manual Z-scanner with the following requirements: - minimum scan range of 15 µm or more in both closed-loop and open-loop operation, - the noise of the sensor less than 250 pm (Adev, bandwidth from 0.1 Hz to 1 kHz), - the scanner must be compatible with all supplied scan modes (item 2.5), for operation in air or liquid, - the cantilever holder must be compatible for operation in air and in liquid. |
| 1.2 | 1 | Piezo driven flexure stage XY-scanner with the following requirements: - scan range of more than 100 µm (operation in closed-loop), - the noise of the sensor less than 0.6 nm (Adev, bandwidth from 0.1 Hz to 1 kHz), - operation in open-loop, closed-loop, and hybrid mode, - possibility to accommodate sample sizes up to 80 mm in diameter or more, - possibility to accommodate sample sizes up to 15 mm or more in height without any modifications to the standard configuration or restriction of the area of movement of the sample stage. |
| 1.3 | 1 | An additional possibility to scan and analyze samples with a height of 25 mm or more. |
| 1.4 | 1 | Illumination and camera system must include: - top-down optical view with the possibility to switch the field of view between 720 and 240 µm, - integrated sample positioning stage that must enable precise alignment of the probe over the desired region of analysis, - illuminator: LED fiber light, - integrated optical system for viewing the probe and sample both from above (top-down) and from below (bottom-view) the sample, - the infrared light source for the deflection measurements (wavelengths in the visible region are not acceptable). |
| 1.5 | 1 | The full complement of operating modes (for fluid and air operation), including: - contact mode, - tapping mode, - tapping mode with phase control and Q-control, - bimodal Dual AC-mode, - electric force microscopy, - electrostatic force microscopy, - Kelvin probe microscopy, - surface potential microscopy, - magnetic force microscopy, - piezoresponse force microscopy, - dual AC-mode, - dual AC- Resonance Tracking, - contact resonance viscoelastic mapping mode, - force spectroscopy and force mapping, - phase imaging, - frequency modulation, - imaging in fluid (with the same setup as air imaging), - nanolithography, - nanomanipulation. |
| 1.6 | 1 | Licensed software for measuring and processing data on a minimum of 3 computers. |
| 1.7 | 1 | The system must include a single probe holder that supports all of the included operating modes and supports operation in both air and liquid for those modes. |
| 1.8 | 1 | Two additional sample holders: - holder for standard magnetic sample discs, - generic holder with spring clips. |
| 1.9 | 1 | Automatic calibration of the cantilever sensitivity (deflection sensitivity) and spring constant by selecting the probe type. The feature must actually calibrate the probe and not use nominal tabulated values for the sensitivity and spring constant. During the calibration, the tip should not touch the sample. |
| 1.10 | 1 | The beam used for deflection detection must approach the probe at an angle that is significantly (>20 degrees) off vertical relative to the sample. |
| 1.11 | 1 | Isolation enclosure must enable: - acoustic and vibration isolation (without the need for compressed air), - the whole system (with optics and without the control system) must be contained within this enclosure, - the table for the microscope must be included. |
| 1.12 | 1 | Additional cell set for dedicated electrochemical experiments should be included with the following: - fluid cell assembly made of PEEK or similar inert nonmetallic material, - cantilever holder made of PEEK or similar inert nonmetallic material, for use with a customer supplied potentiostat, - working electrode sample mount suitable for both conducting and insulating sample substrates, - carbon (graphite) counter electrode, - sealed Ag/AgCl reference electrode, - electrical junction box with connections and programmable jumpers to interface with a wide range of commercial potentiostats, - spare O-rings and assembly hardware, - additional sealing kit for an electrochemical cell, which provides operation of the cell in a fully sealed configuration. |
| 1.13 | 1 | The system must offer measurements in the conductive mode for current imaging and I/V spectroscopy capabilities at user-selectable voltages and points over a broad operating range. This package should include: - cantilever holder with an integral dual gain preamplifier for wide range current measurements at constant applied voltage. - two separate current output channels provide gains of 1x106 (1 µA/V) and 1x109 (1 nA/V), (4 pA to 10 µA), - conductive sample holder mount with sample bias contact, - additional probes for conductive mode: 20 pieces or more, - silver paint, - tools and accessories, - the system must include current measurements from < 1pA to 10 μA. The full dynamic range of current sensing must be recorded simultaneously. |
| 1.14 | 1 | Extended head in Z-range that provides imagining of tall samples and long-range force measurements, in the range up to 40 µm or more. |
| 1.15 | In total 30 pieces or more | A standard pack of cantilevers or probes must include: - 10 x silicon probes with visible apex tip for tapping mode, f: 300 kHz, k: 26 N/m, no tip coating, - 10 x medium soft cantilevers for topography and viscoelastic properties of soft samples, f = 70 kHz, k= 2 N/m, no tip coating, - 10 x (or more) probes with triangular cantilevers (material: Pyrex-Nitride), gold coating. |
| 1.16 | In total 20 or more | Additional cantilevers or probes for tapping mode, operation in air, for soft/medium stiff samples, k = 2 N/m. |
| 1.17 | In total 20 or more | Additional cantilevers of triangular shape (material: Pyrex-Nitride), for imaging in the liquid, gold coating (at least 20 pieces or more). |
| 1.18 | 1 | The additional accessory kit should include: - tweezers, - calibration grating, - power strip and - all necessary cables and accessories. |
| 1.19 | 1 | PC specifications: - Dell Precision T5810 workstation or equivalent, - 3.5 GHz Intel quad-core Haswell Xeon or equivalent, - 8 GB DDR4 RAM, 2 x hard drives 2 TB, 1 x DVD-ROM, - AMD Radeon RX-470 4 GB video card or equivalent, - Hauppauge ImpactVCB-e capture card or equivalent, - 2 x 32” LCD screens, with a resolution of at least 3440 × 1440 or more, - 6 x USB 2.0 ports or more, 4 x USB 3.0 ports or better, - Microsoft Windows 10 Professional 64 bit. |
| 1.20 | 1 | Basic operation training Acquaintance with the equipment and basic training, performed by the engineer of the provider, for 3 users lasting at least 3 working days (also includes the submission of all necessary data and passwords, the cost of travel and accommodation of the engineer; training is carried out at the location FKKT UM, determined by the client). The item is completed after the delivery, installation and startup of the equipment. |
| 1.21 | 1 | Warranty: 24 months or more. |
| 1.22 | 1 | Delivery and installation on the client’s premises. Condition for the successful execution of the item and signing of the handover record: - the supply, installation and start-up of the equipment by the provider, - introduction to equipment and basic education referred to item 1.20, - submission of operating instructions and other technical documentation, - all performance specifications (items from 1.1 to 1.18) must be demonstrated on-site (FKKT UM) before signing technical acceptance documents. |
| 1.23 | 1 | Provided remote access to the device for the needs of calibration and control of the device at least for the duration of the warranty period (2 years). The instrument design must facilitate remote operation and fault diagnostics. |

SPECIFIC REQUIREMENTS FOR ITEM 1.19

Equipment under item 1.19, Computer workstation and accessories, has to fully comply with the technical specifications ensuring the fulfilment of the objectives in accordance with the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, Nos. 51/17, 64/19, 49/20 - ZIUZEEP, 152/20 - ZZUOOP), namely:

PERSONAL COMPUTER

- Personal computers have to meet the latest Energy Star energy efficiency standards applicable on the date of publication of the public procurement notice or on the date of the invitation to tender, or equivalent standards. Means of proof: At the Contracting Authority’s request, the tenderer has to submit: - Energy Star certificate or - a certificate proving that the goods have been awarded the type I eco-label or - the manufacturer's technical documents demonstrating that the requirements have been met.

- Personal computers have to be designed in such a way that the hard drive and the memory are easily accessible using available universal tools (i.e., a screwdriver, spatula, pliers or tweezers) and replaceable. Means of proof: At the Contracting Authority’s request, the tenderer has to submit: - clear disassembly and repair instructions that allow products to be disassembled without damage when replacing key components or other parts for upgrading or repair (e.g., an exploded view showing accessible replaceable parts and necessary tools, video), and - a certificate proving that the goods have been awarded the type I eco-label and thus meet the requirements or - an appropriate proof demonstrating that the requirements are met.

- The volume of the PC has to be measured in accordance with the SIST EN ISO 7779 standard and must not exceed 26 dB when recording to the hard drive. The information has to be labelled with LpAm in dB. Means of proof: At the Contracting Authority’s request, the tenderer has to submit: - a certificate proving that the goods have been awarded the type I eco-label and thus meet the requirements or - An appropriate proof demonstrating that the requirements are met.

MONITORS

- Monitors have to meet the latest Energy Star energy efficiency standards applicable on the date of publication of the public procurement notice or on the date of the invitation to tender. Means of proof: At the Contracting Authority’s request, the tenderer has to submit: - Energy Star certificate or - a certificate proving that the goods have been awarded the type I eco-label or - the manufacturer's technical documents demonstrating that the requirements have been met.

Lot 2: Inert atmosphere chamber – glovebox

| Item number | Number of pieces | REQUIRED |
| --- | --- | --- |
| 2.1 | 1 | Inert atmosphere chamber made of stainless steel type 304 L or 304 or a type of material that has better corrosion resistance: - vacuum antechamber for pre-cleaning the atmosphere with manual control with size of 400 mm in diameter and a length of 600 mm, - vacuum antechamber for pre-cleaning the atmosphere with manual control with size of 150 mm in diameter and a length of 400 mm with a movable drawer for sample transfer between ambient and regulated atmosphere, - at least 1 safety valve, - at least 3 valves for gas input/output in the inert atmosphere chamber (gas input/output to the chamber, gas input/output to a larger antechamber and gas input/output to a smaller antechamber, with a gas purging option for all three spaces), - at least 1 x 220 V/16 A electrical socket, - at least 2 x USB ports for connection between external and internal equipment, compatible with USB 3.0, a with a female connector on the inside of the main chamber and a male connector or extension piece of at least 1 m in length on the outside of the main chamber, - at least 4 electrical interfaces with width of 5 mm for connecting an external potentiostat/galvanostat, type: standard 4 mm banana connector, female connectors may be located on both sides of the chamber, - LED lighting of the work surface inside the inert atmosphere chamber, - 3 or more shelves made of the same type of material as the inert atmosphere chamber, - at least 2 glove adjustments lines, each with position for single glove that enables a single person to work with both left and right hand in such a way that one person can work with both hands at the same time (1 line is meant as a port for 1 glove, 2 lines as ports for 2 gloves, where one port is for the left hand and the other for the right hand), - at least 1 HEPA filter for dust and small particle filtration. |
| 2.2 | 1 | Glove from Butyl or equivalent, 0,4–0,5 mm thick, left glove at least 500 mm or 0.5 m long. |
| 2.3 | 1 | Glove from Butyl or equivalent, 0,4–0,5 mm thick, left glove at least 500 mm or 0.5 m long. |
| 2.4 | 1 | System for automatic control of gas supply, chamber pressure and atmosphere purging with filter capacity of at least 3,5 kg for copper catalyst with ability to remove oxygen up to 1 ppm or equal adsorbent for oxygen gas removal and filter capacity of at least 3,5 kg for molecular sieves or equivalent adsorbent for water removal with reduction capacity of up to 1 ppm. |
| 2.5 | 1 | Oxygen analyzer capable of measuring in the range not exceeding 1 ppm and not less than 1000 ppm with a precision of no less than ± 1 ppm. |
| 2.6 | 1 | Water analyzer capable of measuring in the range not exceeding 1 ppm and not less than 500 ppm with a precision of no less than ± 1 ppm. |
| 2.7 | 1 | Vacuum pump with a minimum suction of 17 m3/h. |
| 2.8 | 1 | At least 1 HEPA filter. |
| 2.9 | 1 | Oxygen adsorbent, at least 3.5 kg. |
| 2.10 | 1 | Water adsorbent, at least 3.5 kg. |
| 2.11 | 1 | Delivery and installation in the client's premises. Condition for successful execution of the item and signing the handover record: - the supply, installation and start-up of the equipment by the provider, - submission of operating instructions and other technical documentation. |
| 2.12 | 1 | Warranty: at least 3 years or more. |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1 of the Operational Programme for Implementation of the European Cohesion Policy 2014-2020: “International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy”, priority investment 1.1 “Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance”, specific objective 1.1.1 “Efficient use of research infrastructure and development of knowledge/competences for better national and international cooperation in the knowledge triangle”.

ENERGY STAR standards: http://www.eu-energystar.org/. The version applicable at the time of issuing these environmental examples and criteria is 6.1 and updates are available at http://www.eu-energystar.org/specifications.htm.

Hard Drive Operating.

LpAm (Low Power AM) is used to express sound pressure emissions.

Energy Star standards: http://www.eu-energystar.org/. The version applicable at the time of issuing these environmental examples and criteria is 7.0 and updates are available at http://www.eu-energystar.org/specifications.htm.