Naziv projekta: Nadgradnja nacionalnih raziskovalnih infrastruktur – RIUM

Izjava o izpolnjevanju minimalnih tehničnih zahtev naročnika in specifikacije ponujene opreme

Naziv ponudnika:

Naslov ponudnika:

Predmet javnega naročila: Nabava visokozmogljivega vektorskega signalnega generatorja in visokozmogljivega vektorskega omrežnega analizatorja

Oznaka javnega naročila: 302/15 – RIUM37-FERI22-23/2021

NAVODILO:

- Ta priloga št. 2 se izpolni in podpiše. Ponudnik mora natančno navesti dejanske tehnične specifikacije ponujene opreme po vseh zahtevah.

- Ponudnik naloži to prilogo št. 2, v Word in PDF obliki (vključno s priloženimi dokazili o izpolnjevanju tehničnih zahtev), v informacijski sistem e-JN v razdelek »Druge priloge«.

Izjavljamo, da ponujena oprema v celoti izpolnjuje minimalne tehnične zahteve naročnika za predmet javnega naročila z naslovom Nabava visokozmogljivega vektorskega signalnega generatorja in visokozmogljivega vektorskega omrežnega analizatorja, in da smo javno naročilo pripravljeni realizirati po naslednji specifikaciji:

| OPREMA 1: Visokozmogljiv vektorski signalni generator |  |  |  |  |
| --- | --- | --- | --- | --- |
| Številka postavke | Število kosov | Zahtevano | Ponujeno | Dokazila (glej opis spodaj!) |
| To poglavje pokriva lastnosti strojne opreme zahtevane naprave |  |  |  |  |
| 1.1 |  | Vector signal generator with two RF paths (path A, path B) |  |  |
| 1.2 |  | Frequency range: at least from 100 kHz (or lower) to at least 44 GHz (or higher), RF path A and B |  |  |
| 1.3 |  | Path A: Baseband Bandwidth: at least 2 GHz (or more), 2 GS |  |  |
| 1.4 |  | Path B: Baseband Bandwidth: at least 500 MHz (or more), 256 MS |  |  |
| 1.5 |  | Wideband baseband I/Q paths to RF (path A, path B) |  |  |
| 1.6 |  | Wideband differential analog I/Q outputs |  |  |
| 1.7 |  | Differential analog I/Q inputs for RF path A |  |  |
| 1.8 |  | Digital baseband output for wideband baseband |  |  |
| 1.9 |  | Fading Simulator with at least 200 MHz (or more) fading bandwidth and MIMO fading simulation with at least 32 fading channels (or more) |  |  |
| 1.10 |  | Phase coherence option for the use of a common local oscillator (LO) for multiple RF signals, to generate phase-coherent RF carriers. Frequencies from at least 200 MHz (or lower) to at least 6 GHz (or higher) to be coupled |  |  |
| 1.11 |  | Ultra low phase noise for RF path A |  |  |
| To poglavje pokriva funkcijske lastnosti zahtevane naprave |  |  |  |  |
| 2.1 |  | Bit Error Rate Tester |  |  |
| 2.2 |  | White Gaussian noise generator |  |  |
| 2.3 |  | AM/AM, AM/PM Predistortion applied to the I/Q signal must be in real-time, any signal configuration changes must being effective immediately. Must be supported Flexible predistortion functions (e.g. predistortion table with up to at least 4000 entries and complex polynomial). |  |  |
| 2.4 |  | Enhanced noise generation with user definable phase noise shapes between 0 dBc and at least -110 dBc |  |  |
| 2.5 |  | Up to at least 25 individually settable notched noise signals generation (or more) |  |  |
| 2.6 |  | Pulse modulator a high-quality pulse modulator with on/off ratio at least >80 dB and very short rise/fall times at least <10 ns |  |  |
| 2.7 |  | High performance pulse generator with pulse widths from at least 20 ns to at least 1 s as well as generation of single pulse, double pulse, delayed pulse (external trigger) |  |  |
| 2.8 |  | Multi-function generator with at least 3 independent function generators and support for sine, pulse, triangle, trapezoid and noise functions |  |  |
| 2.9 |  | 100 MHz and 1GHz reference input/output |  |  |
| 2.10 |  | Additional flexible reference input in 0.1 Hz steps between 1.0 MHz and 100 MHz |  |  |
| 2.11 |  | Real-time Control Interface with streaming option via LAN with at least 2 Mpulses/s, and Agile radar signal generation within at least 2 GHz basebandbandwidth (or more) |  |  |
| 2.12 |  | Crest factor reduction to reduce crest factor of a baseband signal through clipping the high signal peaks and filtering the waveform afterwards |  |  |
| 2.13 |  | User-defined frequency response correction for real time frequency response compensation of internally generated baseband signals by user-defined data |  |  |
| 2.14 |  | Wideband Sequencing of arbitrary waveform segments, realtime change of amplitude, offset frequencies, relative phase, pulse width and off times, signal generation based on time of arrival information (ToA) embedded into the sequencing list, support of loops, nested loops and repetitions |  |  |
| 2.15 |  | AM / FM / PhiM modulation |  |  |
| 2.16 |  | OFDM signal generation |  |  |
| 2.17 |  | Multi carrier CW signal generation with user-definable carrier spacing |  |  |
| 2.18 |  | Baseband power sweep with linear slope ramp waveform |  |  |
| Podpora za naslednje standarde |  |  |  |  |
| 3.1 |  | DVB-H/T, DVB-S2/DVB-S2X |  |  |
| 3.2 |  | EUTRA/LTE, EUTRA/LTE Release 13/14/15, |  |  |
| 3.3 |  | Cellular IOT, Cellular IoT Release 14, Cellular IoT Release 15 |  |  |
| 3.4 |  | Verizon 5GTF Signals, 5G New Radio |  |  |
| 3.5 |  | OneWeb User-Defined Signal Generation, OneWeb Reference Signals |  |  |
| 3.6 |  | LoRa® |  |  |
| 3.7 |  | GPS, Galileo, Glonass (at least 24 channels), BeiDou (at least 24 channels) |  |  |
| Podpora za generiranje signalov iz osebnega računalnika |  |  |  |  |
| 4.1 |  | EUTRA/LTE |  |  |
| 4.2 |  | Multi carrier CW signal generation |  |  |
| 4.3 |  | AWGN |  |  |
| 4.4 |  | EUTRA/LTE Release 13/14/15 |  |  |
| 4.5 |  | 5G New Radio |  |  |
| 4.6 |  | OFDM Signal Generation |  |  |
| 4.7 |  | Cellular IoT Release 14 |  |  |
| 4.8 |  | LoRa® |  |  |
| To poglavje pokriva obvezne dodatke k napravi |  |  |  |  |
| 5.1 | 1 | Cable set for connecting digital IQ interfaces |  |  |
| 5.2 | 1 | Initial calibration |  |  |
| Dobava in namestitev opreme ter garancijsko obdobje |  |  |  |  |
| 6.1 |  | Oprema se dobavi in namesti na Fakulteto za elektrotehniko, računalništvo in informatiko, Koroška cesta 46, 2000 Maribor, v Laboratorij LEIS (prostor G125). |  |  |
| 6.2 |  | Garancijsko obdobje je najmanj tri (3) leta. |  |  |

| OPREMA 2: Visokozmogljiv vektorski omrežni analizator |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Številka postavke | Število kosov | Zahtevano | Ponujeno |  | Dokazila (glej opis spodaj!) |  |  |
| To poglavje pokriva lastnosti strojne opreme zahtevane naprave Vector Network Analyzer |  |  |  |  |  |  |  |
| 1.1 |  | Vector Network Analyzer frequency range from at least and not more than 10 MHz and to at least or more than 120 GHz |  |  |  |  |  |
| 1.2 |  | Single sweep frequency range from at least and not more than 10 MHz and to at least or more than 43.5 GHz |  |  |  |  |  |
| 1.3 |  | Number of ports: at least 4 individual ports (port 1, port 2, port 3, and port 4) with single sweep frequency range given by the reference 1.2, and at least 2 individual ports with extended frequency range above single sweep frequency range |  |  |  |  |  |
| 1.4 |  | Precision frequency reference |  |  |  |  |  |
| 1.5 |  | Timing resolution of at least 8 ns at 20 MHz resolution bandwidth and at least 16 ns at 10 MHz resolution bandwidth |  |  |  |  |  |
| 1.6 |  | IF bandwidth of at least 30 MHz (or more) |  |  |  |  |  |
| 1.7 |  | Frequency resolution of at least 1 mHz (or less) |  |  |  |  |  |
| 1.8 |  | Source step attenuator with at least 70 dB (or more) for port 1 and port 3 |  |  |  |  |  |
| 1.9 |  | Receiver step attenuator with at least 35dB (or more) for port 2 and port 4 |  |  |  |  |  |
| 1.10 |  | Internal pulse modulator for port 1 and port 3 |  |  |  |  |  |
| 1.11 |  | Direct source/receiver access |  |  |  |  |  |
| 1.12 |  | Direct source monitor access for port 1 and 3 |  |  |  |  |  |
| 1.13 |  | Second source for simultaneous measurement of several S-parameters at the same time and ability to test two 2-port devices in parallel on 4-port configuration |  |  |  |  |  |
| 1.14 |  | Direct IF access for external converters |  |  |  |  |  |
| 1.15 |  | LO/IF connections for external mixers |  |  |  |  |  |
| 1.16 | 2* | Frequency converter for frequency range from at least and not more than 50 GHz and to at least or more than 75 GHz with waveguide flange for test port and applicable connectors with cabling for connection with base unit (* at least number of units required to achieve 2 fully individual and fully functional ports in extended frequency range given by the frequency converter) |  |  |  |  |  |
| 1.17 | 2* | Frequency converter for frequency range from at least and not more than 95 GHz and to at least or more than 120 GHz with waveguide flange for test port and applicable connectors with cabling with base unit (* at least number of units required to achieve 2 fully individual and fully functional ports in extended frequency range given by the frequency converter) |  |  |  |  |  |
| Signal and Spectrum Analyzer |  |  |  |  |  |  |  |
| 1.18 |  | Signal and Spectrum Analyzer frequency range from at least and not more than 10 MHz and to at least or more than 120 GHz |  |  |  |  |  |
| 1.19 |  | Single sweep frequency range from at least and not more than 2 Hz and to at least or more than 50 GHz |  |  |  |  |  |
| 1.20 |  | OCXO precision frequency reference with initial calibration accuracy of +-5 x 10-9 |  |  |  |  |  |
| 1.21 |  | External generator control |  |  |  |  |  |
| 1.22 |  | RF preamplifier 30dB from at least 100 kHz (or lower) to at least 50 GHz (or higher) |  |  |  |  |  |
| 1.23 |  | Real-time spectrum analyzer with bandwidth of at least 800 MHz (or more) |  |  |  |  |  |
| 1.24 |  | LO/IF connections for external mixers |  |  |  |  |  |
| 1.25 | 1* | Frequency converter for frequency range from at least and not more than 50 GHz and to at least or more than 75 GHz with waveguide flange for test port and applicable connectors with cabling for connection with base unit (* can be shared with vector network analyzer if applicable) |  |  |  |  |  |
| 1.26 | 1* | Frequency converter for frequency range from at least and not more than 95 GHz and to at least or more than 120 GHz with waveguide flange for test port and applicable connectors with cabling for connection with base unit (* can be shared with vector network analyzer if applicable) |  |  |  |  |  |
| To poglavje pokriva funkcijske lastnosti zahtevane naprave Vector Network Analyzer |  |  |  |  |  |  |  |
| 2.1 |  | Spectrum analyzer operation mode for all VNA ports |  |  |  |  |  |
| 2.2 |  | Time domain analysis (TDR) with support for versatile signal integrity measurements, compliance verification with relevant requirements, and transmission characteristics of signals with jitter or noise, including calculations of the rise time, skew and eye diagrams for different bit patterns. |  |  |  |  |  |
| 2.3 |  | Scalar mixer measurements (with and without access to LO) and arbitrary frequency-converting measurements |  |  |  |  |  |
| 2.4 |  | Vector corrected converter measurements |  |  |  |  |  |
| 2.5 |  | Integrated solution (no additional external hardware) for measurements on pulsed signals, pulse configuration, pulse profile and pulse train measurements with IF bandwidth of at least 30 MHz (or more) |  |  |  |  |  |
| Signal and Spectrum Analyzer |  |  |  |  |  |  |  |
| 2.6 |  | Pulse measurement application with rapid spurious identification, low phase noise, extensive pulse analysis functions. |  |  |  |  |  |
| 2.7 |  | Analog modulation analysis for AM, FM, PM |  |  |  |  |  |
| 2.8 |  | Multi carrier group delay measurements |  |  |  |  |  |
| 2.9 |  | Power amplifier measurements |  |  |  |  |  |
| 2.10 |  | Frequency Response Measurements |  |  |  |  |  |
| 2.11 |  | Noise figure measurements |  |  |  |  |  |
| 2.12 |  | Phase noise measurements |  |  |  |  |  |
| 2.13 |  | Spurious measurements |  |  |  |  |  |
| 2.14 |  | EMI measurements for diagnostics of RF interference |  |  |  |  |  |
| 2.15 |  | Transient measurements |  |  |  |  |  |
| 2.16 |  | Vector signal analysis with possibility to analyze digitally modulated single carriers down to the bit level. |  |  |  |  |  |
| 2.17 |  | Bit Error Rate measurements with Pseudo Random Bit Sequence data |  |  |  |  |  |
| 2.18 |  | Time-side-lobe measurements with automatic measurement of the compression parameters of modulated pulses. Results such as the main lobe vs. side lobe level and the time differences between the main lobe and the side lobes must be displayed in the results summary table. The equipment must support to upload reference pulse waveforms in I/Q format and compare phase and frequency within a pulse with the measured values. |  |  |  |  |  |
| 2.19 |  | Support for direct DPD measurements to determine the extent to which predistortion can compensate for nonlinear effects in an amplifier design |  |  |  |  |  |
| 2.20 |  | Noise Power Ratio Measurement |  |  |  |  |  |
| 2.21 |  | Transient measurements and chirp analysis |  |  |  |  |  |
| 2.22 |  | Multi-Modulation Analysis must support signals containing multiple modulations. It must run in Continuous mode without requiring the external software tool for every individual measurement. |  |  |  |  |  |
| 2.23 |  | OneWeb Reverse Link Measurements to perform measurements on reverse link OneWeb signals |  |  |  |  |  |
| 2.24 |  | User defined frequency response correction using user provided files |  |  |  |  |  |
| Podpora za naslednje standarde |  |  |  |  |  |  |  |
| 3.1 |  | EUTRA/LTE NB-IoT Downlink measurements, EUTRA/LTE FDD Downlink measurement application, EUTRA/LTE FDD Uplink measurement, EUTRA/LTE Downlink MIMO measurement application, EUTRA/LTE Advanced Uplink measurement application, EUTRA/LTE TDD Downlink measurement application, EUTRA/LTE TDD Uplink measurement application |  |  |  |  |  |
| 3.2 |  | 3GPP 5G-NR DL Measurements, 3GPP 5G-NR UL Measurements |  |  |  |  |  |
| 3.3 |  | Verizon 5GTF Measurement Application, Verizon 5GTF UL Measurements |  |  |  |  |  |
| 3.4 |  | OneWeb Reverse Link Measurements to perform measurements on reverse link OneWeb signals |  |  |  |  |  |
| To poglavje pokriva obvezne dodatke k napravi |  |  |  |  |  |  |  |
| 4.1 | 1 | Torque wrench for 3.5/2.92/2.4/1.85 mm connectors with 0.9 Nm force |  |  |  |  |  |
| 4.2 | 1 | Rugged 50 Ohm test cable with applicable connectors and length of at least 600 mm and frequency range of at least or more than 50 GHz |  |  |  |  |  |
| 4.3 | 4 | 50 Ohm RF cables with applicable connectors and frequency range from at least and not more than 0 Hz to at least or more than 50 GHz and length of at least 600 mm |  |  |  |  |  |
| 4.4 | 2 | Power cables for frequency converters |  |  |  |  |  |
| 4.5 | 1 | Calibration kit for frequency range from and not more than 0 Hz to at least or more than 50 GHz |  |  |  |  |  |
| 4.6 | 1 | Calibration kit with sliding match for frequency range from at least and not more than 50 GHz and to at least or more than 75 GHz |  |  |  |  |  |
| 4.7 | 1 | Calibration kit with sliding match for frequency range from at least and not more than 95 GHz and to at least or more than 120 GHz |  |  |  |  |  |
| 4.8 |  | Initial calibration |  |  |  |  |  |
| Dobava in namestitev opreme ter garancijsko obdobje |  |  |  |  |  |  |
| 5.1 |  | Oprema se dobavi in namesti na Fakulteto za elektrotehniko, računalništvo in informatiko, Koroška cesta 46, 2000 Maribor, v Laboratorij LEIS (prostor G125). |  |  |  |  |
| 5.2 |  | Garancijsko obdobje je najmanj tri (3) leta. |  |  |  |  |

| Kraj in datum: | Žig: | Podpis odgovorne osebe: |
| --- | --- | --- |

Operacijo sofinancirata Republika Slovenija, Ministrstvo za izobraževanje, znanost in šport in Evropska unija iz Evropskega sklada za regionalni razvoj. Operacija se izvaja v okviru Operativnega programa za izvajanje evropske kohezijske politike 2014-2020 v okviru prednostne osi 1: »Mednarodna konkurenčnost raziskav, inovacij in tehnološkega razvoja v skladu s pametno specializacijo za večjo konkurenčnost in ozelenitev gospodarstva«, prednostne naložbe 1.1 »Krepitev infrastrukture za raziskave in inovacije ter zmogljivosti za razvoj odličnosti na tem področju, pa tudi spodbujanje pristojnih centrov, zlasti takšnih, ki so evropskega pomena«, specifičnega cilja 1.1.1 »Učinkovita uporaba raziskovalne infrastrukture ter razvoj znanja/kompetenc za boljše nacionalno in mednarodno sodelovanje v trikotniku znanja«.

Ponudnik mora natančno navesti dejanske tehnične specifikacije ponujene opreme po vseh zahtevah.

Ponudnik mora navesti ime priloženega kataloga ali prospekta (dokumenta) proizvajalca opreme in številko strani z navedbo relevantne tehnične specifikacije. Izjemoma lahko ponudnik navede tudi povezavo do spletne stani proizvajalca opreme, na kateri je navedba relevantne tehnične specifikacije. V primeru navedbe povezave do spletnega dokumenta mora ponudnik navesti številko strani z navedbo relevantne tehnične specifikacije. V kolikor se ponudnik posluži uporabe povezave do spletne strani, priporočamo, da ob oddaji kljub temu odda zaslonski posnetek z označeno tehnično specifikacijo iz katerega se jasno vidi da gre za ponujeno opremo, saj se spletna stran lahko do pregleda dokumentacije tudi spremeni ali postane nedostopna. V takem primeru ponudnik sam nosi riziko neizpolnjevanja pogojev.

Ponudnik mora natančno navesti dejanske tehnične specifikacije ponujene opreme po vseh zahtevah.

Ponudnik mora navesti ime priloženega kataloga ali prospekta (dokumenta) proizvajalca opreme in številko strani z navedbo relevantne tehnične specifikacije. Izjemoma lahko ponudnik navede tudi povezavo do spletne stani proizvajalca opreme, na kateri je navedba relevantne tehnične specifikacije. V primeru navedbe povezave do spletnega dokumenta mora ponudnik navesti številko strani z navedbo relevantne tehnične specifikacije. V kolikor se ponudnik posluži uporabe povezave do spletne strani, priporočamo, da ob oddaji kljub temu odda zaslonski posnetek z označeno tehnično specifikacijo iz katerega se jasno vidi da gre za ponujeno opremo, saj se spletna stran lahko do pregleda dokumentacije tudi spremeni ali postane nedostopna. V takem primeru ponudnik sam nosi riziko neizpolnjevanja pogojev.