# CONTRACT AWARD CRITERIA

The contract award criteria are:

Price

Price (relevant is total tender price from APPLICATION FORM) – in consideration is taken the price to two decimal places

The applicant, who offers the lowest price, receives 40 points; other applications receive points calculated using this formula:

(the lowest offered price / offered price) x 40 = number of points

Chosen technical characteristics

1. Small area performance Ag sample using the Ag 3d5/2 photoelectron peak should exceed:

| Minimum spot size 15 µm or less with following performance: - peak resolution (full width at half maximum) at this area of 0.6 eV with a sensitivity equal to or less than 6,000 CPS. - peak resolution (full width at half maximum) at this area of 1.0 eV with a sensitivity equal to or less than 6,000 CPS. | 0 points |
| --- | --- |
| Minimum spot size 10 µm or less with following performance: - peak resolution (full width at half maximum) at this area of 0.6 eV with a sensitivity 25,000 CPS or more. - peak resolution (full width at half maximum) at this area of 1.0 eV with a sensitivity of at least 55,000 CPS or more. | 4 points |

2. Large area performance Ag sample using the Ag 3d5/2 photoelectron peak should exceed:

| - peak resolution (full width at half maximum) of 0.6 eV with a sensitivity of at least 2,500,000 CPS, - peak resolution (full width at half maximum) of 1.0 eV with a sensitivity of at least 6,000,000 CPS. | 4 points |
| --- | --- |
| - peak resolution (full width at half maximum) of 0.6 eV with a sensitivity equal to or less than 1,500,000 CPS, - peak resolution (full width at half maximum) of 1.0 eV with a sensitivity equal to or less than 5,000,000 CPS. | 0 points |

3. Sample exposure to the X-rays during small spot analysis options:

| Large area of the sample is exposed at all times, the analysis area is chosen by the analyzer. | 0 points |
| --- | --- |
| Only analyzed area of the sample is exposed at any analysis condition, the analysis area is chosen by the X-ray beam size and positioning. | 3 points |

4. Electrical 4 contact hot/cold in situ stage and mount options:

| 4 electrical contacts stage has the capability of compucentric rotation for sputter depth profiling and angular resolved analysis, which can be performed during an electrochemical experiment at a user-defined sample temperature within the specified temperature range. | 1 point |
| --- | --- |
| 4 electrical contacts stage is capable of in situ heating and cooling in the range of -120⁰ C to +500⁰ C. | 1 point |
| 4 electrical contacts stage is capable of in situ heating and cooling in the range of -50⁰ C to +400⁰ C. | 0 points |

5. Direct visualization of the analysis area:

| Is performed via either optical camera imaging or via collection of X-ray induced secondary electrons from the analysis area. | 2 points |
| --- | --- |
| Is performed via either optical camera imaging or via rapid collection of X-ray induced photoelectrons from the analysis area. | 0 points |

6. Drift correction of analysis position for prolonged unattended acquisition from multiple separated small areas of the sample:

| Is performed automatically when needed via provided software by adjusting an X-ray beam position on the small area of the sample using prerecorded images of this area correlated with the recorded mechanical positioning of the stage. | 1 point |
| --- | --- |
| No option to correct the analysis position drift during the unattended acquisition from multiple separated small areas of the sample. | 0 points |

7. Multiple sample holders loading options:

| 3 or more sample holders can be loaded simultaneously into the sample entry chamber. | 1 point |
| --- | --- |
| Less than 3 sample holders can be loaded simultaneously into the sample entry chamber. | 0 points |