TECHNICAL DOCUMENTATION

The offered goods must comply with all further technical requirements (all requirements in the table are submitted cumulatively):

State-of-the-art X-ray photoelectron spectrometer capable of analyzing solid specimens to provide surface chemical analysis. The instrument must be equipped with monochromatic Al X-ray source, a 5-axis motorized specimen stage, an electron energy analyzer and associated electron optics, an ultra-high vacuum pumping system, an introduction chamber for introducing specimens into the instrument quickly and easily, an argon ion gun for sputter etching the sample, a separate ultra-high vacuum suitcase compatible load lock, directly connected glove box for inserting air-sensitive samples, and a computer that operates the completely integrated system as well as processes and stores data. The instrument shall have remote operation capability. The subject of the contract is also other services, supplies and obligations listed in this tender documentation.

| Large Area Performance: clean Ag sample using the Ag 3d5/2 photoelectron peak |  |
| --- | --- |
| 1 | Peak energy resolution (full width at half maximum) of 0.6 eV or less with sensitivity 1,500,000 CPS or more. |
| 2 | Peak energy resolution (full width at half maximum) of 1.0 eV or less with sensitivity 5,000,000 CPS or more. |
| Small Area Performance: clean Ag sample using the Ag 3d5/2 photoelectron peak |  |
| 3 | Minimum spot size of 15 m or less, with peak energy resolution (full width at half maximum) of 0.6 eV or less and sensitivity 6,000 CPS or more at this spot size. |
| 4 | Minimum spot size of 15 m or less, with peak energy resolution (full width at half maximum) of 1.0 eV or less and sensitivity 6,000 CPS or more at this spot size. |
| Charge Neutralization and XPS Imaging |  |
| 5 | Minimum Ester peak resolution (full width at half maximum) < 0.85 eV or less with a sensitivity more than 4,000 CPS. |
| 6 | Guaranteed spatial resolution of XPS images 10 µm or less. |
| General Sample Entry Chamber |  |
| 7 | Appropriate pumping kit (turbo molecular pump, ion-sputter pump, wide range ion gauge, automated solenoid valve, rotary vane pump, baking pump, dry pump, oil mist filter and other vacuum components). Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 8 | Fully computer-controlled pump/vent sequences. |
| 9 | Venting system shall be optimized for use of argon gas. |
| 10 | A sample introduction mechanism that permits samples to be introduced into the main chamber without breaking vacuum. The sample introduction mechanism shall be equipped with an air- or water-cooled turbomolecular pump, separate from the main chamber pump, capable of achieving a pressure of less than 2.67 x 10-6 mbar within 15 minutes or less of actuating pump down of the mechanism (no sample on mount). |
| 11 | An optical digital high resolution camera in the sample entry chamber which automatically records an image of the entire sample holder inside the entrance vacuum chamber. It must be possible to use this image to define positions for analysis. In case there are more introduction lines/entries that are not interconnected (such as separate UHV-suitcase entry connection) each shall have an equal separate digital camera. |
| 12 | The minimum area of the sample holder available for sample mounting must be 2000 mm2 or more. |
| 13 | Cooling and heating possibility. |
| 14 | Compatible with the inert gas transfer vessel for contamination free sample introduction, aside from the connected glove box entrance. |
| Main Analysis Chamber |  |
| 15 | Appropriate pumping kit (turbo molecular pump, ion-sputter pump, wide range ion gauge, automated solenoid valve, rotary vane pump, baking pump, dry pump, oil mist filter and other vacuum components). The vacuum system shall be completely automated and shall include automated controls for each of the valves and pumps in the instrument. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 16 | Camera with optical zoom in the sample analysis chamber with a view of the sample analysis position. |
| 17 | Minimum pressure inside the sample analysis chamber of 7 x 10-10 mbar or less. The system shall include an overpressure interlock. |
| 18 | Software with an automated routine to set the sample height prior to analysis. |
| 19 | A set of all necessary standard samples including Cu, Au, Ag, Si, alignment grids and a Faraday cup for alignment and calibration. |
| 20 | Mu metal chamber for better magnetic shielding capability for the analysis chamber. |
| 21 | An automated bake out system capable of baking out the instrument that allows the bake out time to be operator selectable, and include automatic shutoff of the bake out system when finished. The instrument shall be equipped with interlocks to terminate the bake in the event of pressure overload, main chamber pump failure, or excessive temperatures. |
| 22 | An internal lighting to allow viewing of the samples and internal components. The internal lighting system shall not interfere with data collection. |
| 23 | Viewing ports that allow the operator to view sample insertion and proper orientation of the sample at the analysis position. The viewing ports windows shall have sufficient shielding capability to absorb any X-ray flux produced by the instrument. |
| X-ray Source |  |
| 24 | Source of focused monochromated Al Kα X-rays. |
| 25 | Capable of 15 m or less multi-point analysis within the field of view without moving the stage. Types of analysis shall include survey spectra, high-resolution spectra, maps, line scans, and multi-point depth profiling (single sputter crater). |
| Electron Energy Analyzer |  |
| 26 | A 180° hemispherical electron analyzer designed for both small and large area sensitivity for X-ray induced photoelectrons. |
| 27 | Equal or less than 5° angular acceptance for Angle Resolved XPS analysis. |
| 28 | Must possess a multi-channel detector capable of collecting both spectral and imaging data. |
| 29 | Must be capable of collecting unscanned 128 channel XPS maps and depth profiles. In the unscanned maps, each pixel shall contain a complete photoelectron spectrum. |
| Charge Neutralization |  |
| 30 | Charge neutralization capability that permits the analysis of electrically insulating materials. |
| 31 | The neutralization scheme shall effectively stabilize the surface potential of non-conductive specimens. The supplier shall demonstrate that by performing an XPS measurement on a polyethylene terephthalate (PET) polymer film, furnished by the supplier, prior to acceptance of the instrument. Charge neutralization measured at non-conductive PET polymer film specimen at 100 µm x-ray beam size: -peak resolution (full width half maximum) less than 0.85 eV for the O=C-O component, with more than 4,000 CPS for C-C peak. |
| Sputter Ion Source and Argon Gas Cluster Ion Source |  |
| 32 | The instrument must include a combination of argon ion sputtering capabilities that include: -monoatomic argon sputtering, -argon gas cluster ion sputtering, available either as a single compound gun or two separated guns. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 33 | Performance: -acceleration voltage of 1 to 5 kV or higher in monatomic mode, -acceleration voltage of 2.5 to 20 kV or higher in cluster ion mode, -maximum cluster size of 2500 atoms per cluster or more, -minimum cluster size of 600 atoms per cluster or less. |
| 34 | The instrument operating software shall be capable of automatically turning off the ionizer filament and the flow of the argon gas at the completion of an unattended XPS acquisition. |
| 35 | The ion gun column shall contain a mass filter where smaller clusters and monoatomic argon are removed, as well as a physical bend to eliminate neutral particles from the ion beam. |
| Ultraviolet He Source for UPS measurements |  |
| 36 | A fully computer controlled windowless high intensity bakeable discharge UPS source that can produce He(I) and He(II) photons, including two stage differential pumping with automated UV lamp ignition and gas delivery. |
| 37 | -UPS sensitivity of more than 1,000,000 CPS on the Ag 4d peak, -energy resolution of the Fermi edge of 80/20 slope width of less than 120 meV at room temperature. |
| 38 | The Ultraviolet He Source option shall have a differential pumping system comprised of a rotary pump and a turbomolecular pump. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 39 | Under computer control, the instrument must have the ability to bias the sample for obtaining work function measurements. |
| 40 | Possibility of reducing the UV source power for sensitive materials (polymers) applications. |
| Stage and General Sample Holders |  |
| 41 | Five (5) axis (X, Y, Z, rotation, tilt) high precision motorized stage. Its operation shall be fully integrated into the control software for navigation and sample positioning and allow compucentric rotation of at least 170⁰ and compucentric automated tilt of at least 90⁰. |
| 42 | The stage must also be capable of automatically alternating the tilt and/or rotate between the sputter and acquire portions of a depth profile cycle. |
| 43 | The stage shall be capable of motorized azimuthal rotation ± 170° during ion sputtering, and the center of rotation shall be selectable at any point on the specimen that is within 8 mm or more radius from the center of the sample mount. |
| 44 | The stage shall be capable of accommodating a sample holder to which flat specimens up to 13 mm or more in height are attached. All positions of the whole sample holder must be accessible. |
| 45 | 5 general sample holders for mounting flat samples. In case of different sample holder sizes, 5 larger area and 4 smaller area ones are required. At least one sample holder plate shall have a possibility to mount the standard stub SEM sample holders with diameter of 12.7 mm and pin length of at least 6 mm. |
| Hot/Cold (H/C) Stage |  |
| 46 | Direct heating and cooling of the sample in the main analysis chamber, at the sample analysis position. -Possibility to cool the sample to at least -100⁰C or more at the sample analysis position. -Possibility to heat the sample to at least +800⁰C or more at the sample analysis position. Heating and cooling kit shall include: -Sample holder for heating and cooling of the sample in the main sample chamber, -The heater element and thermocouple used to obtain temperature measurements must be located within the sample mounts. -PC software for control of heating and cooling devices. |
| 47 | Stage full range of motion on at least 4 axes or more must be available during heating and cooling experiments. |
| 4 (four) Electrical Contact Hot/Cold (H/C) Stage and Holder |  |
| 48 | Stage shall have 4 electrical contacts for in situ electrochemical experiments, with a maximum of 100 V DC and more and 1 A DC applied to each of the 4 contact electrodes. |
| 49 | Stage shall be capable of in situ heating and cooling within the temperature range of -50⁰C or less to +400⁰C or more with the available 4 contacts for in situ electrochemical experiments. |
| 50 | Heater and thermocouple integrated in the sample holder for enhanced reliability and temperature control. |
| 51 | Stage full range of motion on at least 4 axes or more must be available during heating and cooling experiments. |
| 52 | It shall be possible to load under inert atmosphere. |
| 53 | Electrochemical cell shall be equipped with a holder suitable for cross-section samples with the height of the sample of up to 8.5 mm, which can be then properly connected to the holder's contacts and cycled. |
| Controlled Atmosphere Transfer Vessel |  |
| 54 | Vessel for transporting air sensitive samples from a glove box to the instrument without exposing the samples to the ambient atmosphere. |
| Glove Box with Built-in Interface to XPS Entry Chamber |  |
| 55 | Fully functional argon gas atmosphere glove box with two or three glove entries with technical specifications listed in Table 1. |
| 56 | Hardware connection interface to the XPS instrument introduction chamber from the glove box including all the needed integrations for pumping and vacuum protection of the interlock. |
| 57 | The glove box entry line into the main chamber must be physically separated from the UHV entry line. Additionally, it shall be also possible to insert the sample holder directly into the general sample entry chamber outside of the glove box. |
| 58 | All necessary tools, pumping kits, overpressure interlocks, and interconnecting mechanisms needed for flawless transfer of sample holder from the glove box into the main entry chamber and back under uninterrupted argon gas atmosphere conditions. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 59 | Venting system shall be optimized for use of argon gas. |
| Built-in Hardware Interface for UHV Flag Type Samples Introduction, Compatible with the UHV Suitcase |  |
| 60 | Interface to the UHV suitcase for introduction of UHV samples including all needed hardware including transfer arm with grabber to move the flag type sample, support frame for chamber and for UHV suite case, lifting trolley for UHV suitcase support, etc. including the modified sample holder with the proper receiver for flag type samples |
| 61 | Separate chamber entry setup for connecting a UHV suitcase equipped with separate bake out system, UHV pumping kit, and adapted with all necessary tools and mechanisms for proper insertion of flag-type sample holders. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 62 | Viewing ports that allow the operator to view sample insertion and proper orientation of the sample at the analysis position. The viewport windows shall have sufficient shielding capability to absorb any X-ray flux produced by the instrument. |
| 63 | A digital camera capable of taking high resolution image of the whole sample holder in the UHV introduction chamber that can be later recognized and used by the large area analysis software for taking spectra from the user defined areas of the sample holder. |
| Peripheral Appliances, Software, PCs, Monitors, Remote Control |  |
| 64 | Mains power distribution unit for the main instrument. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 65 | Cooled re-circulating water supply, water chiller for the entire instrument. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 66 | Uninterruptable power supply system with suitable power characteristics and power autonomy of at least 30 min. System must have a bypass switch that allows the XPS system to operate while the UPS system is offline for service or repairs. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. |
| 67 | State-of-the-art electronics with reliable USB communication and data streaming directly to the system PC featuring software energy scale calibration; voltage readout in software for remote diagnostics; low system overhead during data acquisition; real-time display of data and acquisition status. |
| 68 | A state-of-the art PC including a recordable/erasable DVD drive, with at least 16 Gb RAM or more, at least 1 Tb or more hard disk, 2 x 26” IPS LED backlit monitors. |
| 69 | Latest version of acquisition software with full analytical capability including surveys, high-resolution multiplexes, sputter depth profiles, line scans, chemical images, angle dependent XPS, automated analyses and user-defined settings, with the option of free software upgrade within the warranty period after acceptance. |
| 70 | The system shall support fully remote operation and fault diagnostics. |
| 71 | At least 15 licenses of the latest versions of post-acquisition analysis software with an extensive library of post-analysis data processing algorithms including: atomic concentration quantification, background subtraction, smoothing, peak identification, linear least squares fitting, target factor analysis, curve and peak analysis and separation of multiple chemical states in maps, line scans and profiles. |
| 72 | At least 4 licenses of software shall allow the thin film structure estimation using single angle spectral data as well as calculation of discrete layers’ thicknesses. |
| 73 | A set of special tools and spares required for initial maintenance of the instrument. |
| Delivery and Installation |  |
| 74 | Delivery DDP National Institute of Chemistry, Ljubljana, Slovenia, to the instruments' installation room. |
| 75 | Installation of the main instrument and related appliances, glove box and connecting hardware, and UHV entry and connecting hardware. Initial setup of hardware and software, which is included in the offer price, verification and confirmation of all specifications. |
| 76 | Technical specifications shall be confirmed by performing testing of the instrument at the customer’s site with the standard testing samples provided by the supplier. The testing shall be conducted by the supplier after the installation prior acceptance signoff by the customer, and the results shall comply with the official instrument’s technical specifications and the factory acceptance tests of the instrument performed upon production. |
| 77 | Delivery and installation procedure until the full functionality is accepted including acceptance tests at the customer site shall be completed by the supplier and signed off by the customer by the latest 10.12.2021. |

Table 1: The glove box for integrated XPS configuration shall comply with the following characteristics:

- Single-sided operated glovebox with main antechamber, 5° to 7° sloped front window, and workspace dimensions of 1250 mm width x 900 mm height x 780 mm depth (+/-2% measures discrepancy allowed).

- Specifications:

- 1 x glovebox, Stainless Steel 1.4301

- 1 x stand, height customized to match the XPS loadlock height; inkl. castors and machine feet with adjustable height of at least 800 mm (+/-5%)

- Windows with chemicals and scratches resistant polycarbonate with SAPHIR coating

- 2 or 3 gloveport feedthroughs of d=220 mm, incl. gloves (butyl-kautschuk, 0.4 mm thickness, size L)

- LED lighting of the workspace

- 3 x shelves, mounted at backside, height adjustable

- 2 x dust filter HEPA H13 inside box at gas inant outlet

- 4 x flanges for feedthroughs DN40KF: 230V, 1ph

- Main antechamber, diameter at least 390 mm, length at least 600 mm with sliding tray, positioned on the sidewall opposite to the XPS loadlock connection

- Mini antechamber, diameter at least 150 mm, length 400 mm 1/3 inside, 2/3 outside box. Hinged cover shall be present inside and outside, with manual operation via 3-way valve; it shall include sliding tray.

- Underpressure relief valve with opening pressure <-12,5 mbar, closing pressure >-12,5 mbar +/-1-2 mbar, max. pressure -20mbar with max. flow rate 200 l/min

- Overpressure relief valve with opening pressure >12,5mbar, closing pressure <12,5mbar +/-1-2mbar, max. pressure 20mbar with max. flow rate

- Sensor-set: H2O and O2 with following or equivalent, specifications:

- H2O-analyzer: moisture probe, PLC-controlled, with measuring range 0-500ppm, transducer, 0-10V output, cable shall be included.

- O2-analyzer: Oxygen probe, PLC-controlled, with measuring range 0-1000ppm, transducer, solid state device, 0-10V output, cable shall be included.

- Gas purification system for closed cycle circulation:

- Removal: oxygen and moisture, <1 ppm (referring to gloveboxes with max. leak rate O2 = 0.05 Vol%/h, according ISO 10648-2)

- Filter: 1-filter column gas purifier

- Filter capacity: at least 19 l oxygen and 900 g water in case of 2 gloveports, or at least 36 l oxygen and 1350 g water in case of 3 gloveports

- Circulation blower: up to 45 m³/h (2 gloveports) or 88 m³/h (3 gloveports), water cooled, frequency controlled, vibration dampened

- Filter column valves: electro-pneum. main valves

- Filter column regeneration shall be performed via automatic sequencing program

- Control unit: PLC, touch panel 7 inches (Siemens or equivalent), colour

- Automatic box pressure control in adjustable range between at least -14 to +14 mbar or more (box can be operated at over- or under pressure), pressure sensor for glovebox shall be included

- Additional foot switch to change box pressure during operation including graphical display (24h- and 180 day trend shall be possible) of box pressure

- Vacuum pump: at least 17 m³/h performance, oil-sealed with gas ballast and oilmist filter

- Control function of main antechamber semi-automatic evacuation/refill shall be operated on Touch Panel

- Alarm interface (potential free contact) for glovebox alarm to be used to stop evacuation of the XPS loading chamber if the box pressure drops below the alarm limit

- Unit for inertizing (purging) the workspace with inert gas with operation via operation panel of the gas purifier including PLC-controlled valves, max. flow rate of at least 200 l/min; incl. manual regulation valve for reduction of the gas flow, function of Purging ON / OFF, max. purging time programmable from to at least 999 min. The O2 and H2O concentrations shall be readable during purge.

- All necessary hardware modifications for direct connection to XPS entry chamber for flawless incertion of air-sensitive samples including all flanges, valves and cabling

- Economical and power saving operation mode of glovebox workstation and gas purifier that can reduce the power consumption by up to 90% and minimize the noise level; positive pressure regulation shall be possible without vacuum pump, power consumption shall be changeable to optimize blower speed, and automatic light switch off function shall be available.

| REQUIRED CHARACTERISTIC Large Area Performance: clean Ag sample using the Ag 3d5/2 photoelectron peak |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| --- | --- | --- | --- | --- |
| 1 | Peak energy resolution (full width at half maximum) of 0.6 eV or less with sensitivity 1,500,000 CPS or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 2 | Peak energy resolution (full width at half maximum) of 1.0 eV or less with sensitivity 5,000,000 CPS or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Small Area Performance: clean Ag sample using the Ag 3d5/2 photoelectron peak |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 3 | Minimum spot size of 15 m or less, with peak energy resolution (full width at half maximum) of 0.6 eV or less and sensitivity 6,000 CPS or more at this spot size. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 4 | Minimum spot size of 15 m or less, with peak energy resolution (full width at half maximum) of 1.0 eV or less and sensitivity 6,000 CPS or more at this spot size. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Charge Neutralization and XPS Imaging |  | CHARACTERISTIC OF OFFERED GOODS |  |  |
| 5 | Minimum Ester peak resolution (full width at half maximum) < 0.85 eV or less with a sensitivity more than 4,000 CPS. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 6 | Guaranteed spatial resolution of XPS images 10 µm or less. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC General Sample Entry Chamber |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 7 | Appropriate pumping kit (turbo molecular pump, ion-sputter pump, wide range ion gauge, automated solenoid valve, rotary vane pump, baking pump, dry pump, oil mist filter and other vacuum components). Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 8 | Fully computer-controlled pump/vent sequences. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 9 | Venting system shall be optimized for use of argon gas. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 10 | A sample introduction mechanism that permits samples to be introduced into the main chamber without breaking vacuum. The sample introduction mechanism shall be equipped with an air- or water-cooled turbomolecular pump, separate from the main chamber pump, capable of achieving a pressure of less than 2.67 x 10-6 mbar within 15 minutes or less of actuating pump down of the mechanism (no sample on mount). | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 11 | An optical digital high resolution camera in the sample entry chamber which automatically records an image of the entire sample holder inside the entrance vacuum chamber. It must be possible to use this image to define positions for analysis. In case there are more introduction lines/entries that are not interconnected (such as separate UHV-suitcase entry connection) each shall have an equal separate digital camera. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 12 | The minimum area of the sample holder available for sample mounting must be 2000 mm2 or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 13 | Cooling and heating possibility. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 14 | Compatible with the inert gas transfer vessel for contamination free sample introduction, aside from the connected glove box entrance. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Main Analysis Chamber |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 15 | Appropriate pumping kit (turbo molecular pump, ion-sputter pump, wide range ion gauge, automated solenoid valve, rotary vane pump, baking pump, dry pump, oil mist filter and other vacuum components). The vacuum system shall be completely automated and shall include automated controls for each of the valves and pumps in the instrument. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 16 | Camera with optical zoom in the sample analysis chamber with a view of the sample analysis position. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 17 | Minimum pressure inside the sample analysis chamber of 7 x 10-10 mbar or less. The system shall include an overpressure interlock. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 18 | Software with an automated routine to set the sample height prior to analysis. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 19 | A set of all necessary standard samples including Cu, Au, Ag, Si, alignment grids and a Faraday cup for alignment and calibration. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 20 | Mu metal chamber for better magnetic shielding capability for the analysis chamber. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 21 | An automated bake out system capable of baking out the instrument that allows the bake out time to be operator selectable, and include automatic shutoff of the bake out system when finished. The instrument shall be equipped with interlocks to terminate the bake in the event of pressure overload, main chamber pump failure, or excessive temperatures. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 22 | An internal lighting to allow viewing of the samples and internal components. The internal lighting system shall not interfere with data collection. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 23 | Viewing ports that allow the operator to view sample insertion and proper orientation of the sample at the analysis position. The viewing ports windows shall have sufficient shielding capability to absorb any X-ray flux produced by the instrument. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC X-ray Source |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 24 | Source of focused monochromated Al Kα X-rays. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 25 | Capable of 15 m or less multi-point analysis within the field of view without moving the stage. Types of analysis shall include survey spectra, high-resolution spectra, maps, line scans, and multi-point depth profiling (single sputter crater). | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Electron Energy Analyzer |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 26 | A 180° hemispherical electron analyzer designed for both small and large area sensitivity for X-ray induced photoelectrons. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 27 | Equal or less than 5° angular acceptance for Angle Resolved XPS analysis. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 28 | Must possess a multi-channel detector capable of collecting both spectral and imaging data. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 29 | Must be capable of collecting unscanned 128 channel XPS maps and depth profiles. In the unscanned maps, each pixel shall contain a complete photoelectron spectrum. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Charge Neutralization |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 30 | Charge neutralization capability that permits the analysis of electrically insulating materials. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 31 | The neutralization scheme shall effectively stabilize the surface potential of non-conductive specimens. The supplier shall demonstrate that by performing an XPS measurement on a polyethylene terephthalate (PET) polymer film, furnished by the supplier, prior to acceptance of the instrument. Charge neutralization measured at non-conductive PET polymer film specimen at 100 µm x-ray beam size: -peak resolution (full width half maximum) less than 0.85 eV for the O=C-O component, with more than 4,000 CPS for C-C peak. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Sputter Ion Source and Argon Gas Cluster Ion Source |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 32 | The instrument must include a combination of argon ion sputtering capabilities that include: -monoatomic argon sputtering, -argon gas cluster ion sputtering, available either as a single compound gun or two separated guns. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 33 | Performance: -acceleration voltage of 1 to 5 kV or higher in monatomic mode, -acceleration voltage of 2.5 to 20 kV or higher in cluster ion mode, -maximum cluster size of 2500 atoms per cluster or more, -minimum cluster size of 600 atoms per cluster or less. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 34 | The instrument operating software shall be capable of automatically turning off the ionizer filament and the flow of the argon gas at the completion of an unattended XPS acquisition. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 35 | The ion gun column shall contain a mass filter where smaller clusters and monoatomic argon are removed, as well as a physical bend to eliminate neutral particles from the ion beam. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Ultraviolet He Source for UPS measurements |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 36 | A fully computer controlled windowless high intensity bakeable discharge UPS source that can produce He(I) and He(II) photons, including two stage differential pumping with automated UV lamp ignition and gas delivery. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 37 | -UPS sensitivity of more than 1,000,000 CPS on the Ag 4d peak, -energy resolution of the Fermi edge of 80/20 slope width of less than 120 meV at room temperature. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 38 | The Ultraviolet He Source option shall have a differential pumping system comprised of a rotary pump and a turbomolecular pump. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 39 | Under computer control, the instrument must have the ability to bias the sample for obtaining work function measurements. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 40 | Possibility of reducing the UV source power for sensitive materials (polymers) applications. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Stage and General Sample Holders |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 41 | Five (5) axis (X, Y, Z, rotation, tilt) high precision motorized stage. Its operation shall be fully integrated into the control software for navigation and sample positioning and allow compucentric rotation of at least 170⁰ and compucentric automated tilt of at least 90⁰. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 42 | The stage must also be capable of automatically alternating the tilt and/or rotate between the sputter and acquire portions of a depth profile cycle. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 43 | The stage shall be capable of motorized azimuthal rotation ± 170° during ion sputtering, and the center of rotation shall be selectable at any point on the specimen that is within 8 mm or more radius from the center of the sample mount. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 44 | The stage shall be capable of accommodating a sample holder to which flat specimens up to 13 mm or more in height are attached. All positions of the whole sample holder must be accessible. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 45 | 5 general sample holders for mounting flat samples. In case of different sample holder sizes, 5 larger area and 4 smaller area ones are required. At least one sample holder plate shall have a possibility to mount the standard stub SEM sample holders with diameter of 12.7 mm and pin length of at least 6 mm. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Hot/Cold (H/C) Stage |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 46 | Direct heating and cooling of the sample in the main analysis chamber, at the sample analysis position. -Possibility to cool the sample to at least -100⁰C or more at the sample analysis position. -Possibility to heat the sample to at least +800⁰C or more at the sample analysis position. Heating and cooling kit shall include: -Sample holder for heating and cooling of the sample in the main sample chamber, -The heater element and thermocouple used to obtain temperature measurements must be located within the sample mounts. -PC software for control of heating and cooling devices. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 47 | Stage full range of motion on at least 4 axes or more must be available during heating and cooling experiments. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC 4 (four) Electrical Contact Hot/Cold (H/C) Stage and Holder |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 48 | Stage shall have 4 electrical contacts for in situ electrochemical experiments, with a maximum of 100 V DC and more and 1 A DC applied to each of the 4 contact electrodes. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 49 | Stage shall be capable of in situ heating and cooling within the temperature range of -50⁰C or less to +400⁰C or more with the available 4 contacts for in situ electrochemical experiments. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 50 | Heater and thermocouple integrated in the sample holder for enhanced reliability and temperature control. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 51 | Stage full range of motion on at least 4 axes or more must be available during heating and cooling experiments. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 52 | It shall be possible to load under inert atmosphere. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 53 | Electrochemical cell shall be equipped with a holder suitable for cross-section samples with the height of the sample of up to 8.5mm, which can be then properly connected to the holder's contacts and cycled. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Controlled Atmosphere Transfer Vessel |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 54 | Vessel for transporting air sensitive samples from a glove box to the instrument without exposing the samples to the ambient atmosphere. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Glove Box with Built-in Interface to XPS Entry Chamber |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 55 | Fully functional argon gas atmosphere glove box with two or three glove entries with technical specifications listed in Table 1. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 56 | Hardware connection interface to the XPS instrument introduction chamber from the glove box including all the needed integrations for pumping and vacuum protection of the interlock. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 57 | The glove box entry line into the main chamber must be physically separated from the UHV entry line. Additionally, it shall be also possible to insert the sample holder directly into the general sample entry chamber outside of the glove box. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 58 | All necessary tools, pumping kits, overpressure interlocks, and interconnecting mechanisms needed for flawless transfer of sample holder from the glove box into the main entry chamber and back under uninterrupted argon gas atmosphere conditions. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 59 | Venting system shall be optimized for use of argon gas. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Built-in Hardware Interface for UHV Flag Type Samples Introduction, Compatible with the UHV Suitcase |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 60 | Interface to the UHV suitcase for introduction of UHV samples including all needed hardware including transfer arm with grabber to move the flag type sample, support frame for chamber and for UHV suite case, lifting trolley for UHV suitcase support, etc. including the modified sample holder with the proper receiver for flag type samples | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 61 | Separate chamber entry setup for connecting a UHV suitcase equipped with separate bake out system, UHV pumping kit, and adapted with all necessary tools and mechanisms for proper insertion of flag-type sample holders. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 62 | Viewing ports that allow the operator to view sample insertion and proper orientation of the sample at the analysis position. The viewport windows shall have sufficient shielding capability to absorb any X-ray flux produced by the instrument. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 63 | A digital camera capable of taking high resolution image of the whole sample holder in the UHV introduction chamber that can be later recognized and used by the large area analysis software for taking spectra from the user defined areas of the sample holder. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| REQUIRED CHARACTERISTIC Peripheral Appliances, Software, PCs, Monitors, Remote Control |  | CHARACTERISTIC OF OFFERED GOODS |  | DESCRIBED IN (document), PAGE NUMBER: |
| 64 | Mains power distribution unit for the main instrument. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 65 | Cooled re-circulating water supply, water chiller for the entire instrument. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 66 | Uninterruptable power supply system with suitable power characteristics and power autonomy of at least 30 min. System must have a bypass switch that allows the XPS system to operate while the UPS system is offline for service or repairs. Supplied cables and hoses must be adjustable in length up to at least 8 m or more. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 67 | State-of-the-art electronics with reliable USB communication and data streaming directly to the system PC featuring software energy scale calibration; voltage readout in software for remote diagnostics; low system overhead during data acquisition; real-time display of data and acquisition status. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 68 | A state-of-the art PC including a recordable/erasable DVD drive, with at least 16 Gb RAM or more, at least 1 Tb or more hard disk, 2 x 26” IPS LED backlit monitors. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 69 | Latest version of acquisition software with full analytical capability including surveys, high-resolution multiplexes, sputter depth profiles, line scans, chemical images, angle dependent XPS, automated analyses and user-defined settings, with the option of free software upgrade within the warranty period after acceptance. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 70 | The system shall support fully remote operation and fault diagnostics. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 71 | At least 15 licenses of the latest versions of post-acquisition analysis software with an extensive library of post-analysis data processing algorithms including: atomic concentration quantification, background subtraction, smoothing, peak identification, linear least squares fitting, target factor analysis, curve and peak analysis and separation of multiple chemical states in maps, line scans and profiles. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 72 | At least 4 licenses of software shall allow the thin film structure estimation using single angle spectral data as well as calculation of discrete layers’ thicknesses. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |
| 73 | A set of special tools and spares required for initial maintenance of the instrument. | YES | NO |  |
|  |  | The applicant must circle the suitable option |  |  |