Project title: Upgrading National Research Infrastructures – RIUM

COMPLIANCE DECLARATION WITH MINIMUM TECHNICAL REQUIREMENTS OF THE OFFERED EQUIPMENT

Correction No. 3

Name of the tenderer:

Address of the tenderer:

Subject of the public procurement: Purchase of Thin Film Deposition System

Designation of the public procurement: 302/13 – RIUM36-FERI03.2/2021

INSTRUCTIONS:

- This Annex No. 2 must be completed and signed. The tenderer must indicate the actual technical specifications of the equipment offered in line with all requirements.

- The tenderer shall upload this Annex No. 2 to the e-JN information system in the section “Other Annexes”.

We declare that the offered equipment fully meets the minimum technical requirements of the client for the public procurement titled "Purchase of Thin Film Deposition System", and that we are ready to realize the order according to the following specification:

|  | EQUIPMENT |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  |  |  |  |  |  |  | DATASHEET |
| 1 | Chamber and vacuum subsystem: |  |  |  |  |  |  |  |  |  |
|  | - Chamber dimensions: diameter minimal 800 mm; height minimal 1000 mm. | Diameter: _______________ mm |  |  |  |  |  |  |  |  |
|  |  | Height: _______________ mm |  |  |  |  |  |  |  |  |
|  | - Stainless steel electro polished chamber (304) UHV. | - YES |  |  |  | - NO |  |  |  |  |
|  | - View ports: minimal 3 | Number of viewports: _____________ |  |  |  |  |  |  |  |  |
|  | - Rotary drive fixture for substrates up to 150 mm. | Rotary drive fixture: |  | - YES |  |  |  |  | - NO |  |
|  |  | Substrate size: _________ mm |  |  |  |  |  |  |  |  |
|  | - Water cooled chamber wall and chamber door. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Chamber shielding by UHV Aluminum foil to shield the chamber walls and deck plates for shielding the chamber floor. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Silver plated hardware against elevated temperatures inside the chamber. | - YES |  |  |  | - NO |  |  |  |  |
|  | - Shield foil mounted in front of the turbo pump port. | - YES |  |  |  | - NO |  |  |  |  |
|  | - Port for monitoring the partial oxygen pressure in the chamber. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Wide range turbo pumping system with capability at least 6500 l/s, compatible with O3. Two smaller pumps are acceptable. | Turbo molecular pump compatible with O3 | - YES |  |  |  |  | - NO |  |  |
|  |  | Turbo molecular pump capability: _______________ l/s |  |  |  |  |  |  |  |  |
|  | - ID of the turbo pump: minimal 320 mm (ISO). | ID of the turbo pump (ISO): ___________ mm |  |  |  |  |  |  |  |  |
|  | - Roughing pump: dry roughing pump with a speed minimum of 25 l/s, compatible with O3. | Roughing pump compatible with O3 | - YES |  |  |  |  | - NO |  |  |
|  |  | Roughing pump capability: _______________ l/s |  |  |  |  |  |  |  |  |
|  | - Pump down time: less than 5x10-6 Torr in less than 45 min. | Pump down vacuum: _________________ Torr |  |  |  |  |  |  |  |  |
|  |  | Pump down time: ___________________ minutes |  |  |  |  |  |  |  |  |
|  | - System ultimate pressure: better than 3x10-7 Torr. | System ultimate pressure: _________________ Torr |  |  |  |  |  |  |  |  |
|  | - Time to achieve the ultimate pressure: in less than 12 h. | Time to achieve the ultimate pressure: _________________ h (hour) |  |  |  |  |  |  |  |  |
|  | - Pressure sensor at the roughing line. | - YES |  |  | - NO |  |  |  |  |  |
|  | - High vacuum gauge designed with three (3) sensors: hot ion, convection enhanced Pirani and vacuum switch. | - YES |  |  | - NO |  |  |  |  |  |
|  | - High vacuum gauge measurement range: from 1100 Torr to 3.8x10-10 Torr. | High vacuum gauge measurement range: from ______________ Torr to _______________ Torr |  |  |  |  |  |  |  |  |
|  | - Absolute pressure measurement gauge: range 100mTorr FS | Absolute pressure measurement gauge range: od _________________ mTorr |  |  |  |  |  |  |  |  |
|  | - Substrate heating system with IR heater arrays. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Substrate heating temperature: minimum 190 °C. | Substrate heating temperature: ___________________ °C |  |  |  |  |  |  |  |  |
|  | - Chamber illumination system. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Chamber illumination must be possible also during deposition process. | - YES |  |  | - NO |  |  |  |  |  |
|  | Sample loading subsystem: |  |  |  |  |  |  |  |  |  |
|  | - Manual or motorized load lock system. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Possibility to load samples/substrates with dimensions: minimal diameter: 150 mm and minimal height: 75 mm. | Diameter: ______________ mm |  |  |  |  |  |  |  |  |
|  |  | Height: ______________ mm |  |  |  |  |  |  |  |  |
|  | - Load lock chamber equipped with a fiber optic feedthrough, that can be mounted on the sample which is transferred into the deposition chamber. | - YES |  |  |  | - NO |  |  |  |  |
|  | - Sample loading platform must allow rotation (by a stepper motor) of the sample connected with an optical fiber in CV and CCV directions. | - YES |  |  |  | - NO |  |  |  |  |
|  | - Rotation of the fiber sample while holding one end in the fixed position in the load lock. | - YES |  |  |  | - NO |  |  |  |  |
|  | Primary deposition system: |  |  |  |  |  |  |  |  |  |
|  | - Dual cathode magnetron reactive sputtering system: minimum two (2) pairs of special low pressure operation sputtering cathodes (in total four cathodes), for sequential or simultaneous sputtering of two different oxides or nitrides from metal targets. | - YES |  |  | - NO |  |  |  |  |  |
|  | - The subsystem should provide the possibility to deposit at least the following oxides: SiO2, TiO2, HfO2, GeO2, and nitrides: TiN, SiN, AlN. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Diameter of sputtering cathodes: minimum 4,00” | Diameter of sputtering cathodes: _____________ inch |  |  |  |  |  |  |  |  |
|  | - Minimum two output power supplies for operating either sequentially or in common exciter mode. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Power per power supply: more than 4 kW. | Power per power supply: ________________ kW |  |  |  |  |  |  |  |  |
|  | - Output frequency per power supply compatible with sputtering cathodes, pure sine wave or equivalent. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Cathodes should have built in gas rings for direct process gas injection. | - YES |  |  | - NO |  |  |  |  |  |
|  | - High target utilization of the cathodes at high pressure: 40% or greater at pressures bellow 1mTorr. | Target utilization of the cathodes at pressures bellow 1mTorr: ___________ % |  |  |  |  |  |  |  |  |
|  | Secondary deposition subsystem: |  |  |  |  |  |  |  |  |  |
|  | - An RF sputtering system. | - YES |  |  | - NO |  |  |  |  |  |
|  | - RF sputtering cathode with a diameter: minimal 4”. | RF sputtering cathode diameter: ________________ inch |  |  |  |  |  |  |  |  |
|  | - Cathode designed to operate as a pure dielectric material sputtering source. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Power supply: minimum 1200 W. | Power supply: __________________ W |  |  |  |  |  |  |  |  |
|  | - Automatic matching network. | - YES |  |  | - NO |  |  |  |  |  |
|  | - The secondary deposition subsystem must be able to operate in co-sputter mode with all cathodes from the primary deposition subsystem. | - YES |  |  |  | - NO |  |  |  |  |
|  | Gas controls subsystem: |  |  |  |  |  |  |  |  |  |
|  | - Gas control system for the following gases N2, O2 and Ar. | - YES |  |  |  | - NO |  |  |  |  |
|  | - Gas lines equipped with an appropriate particle filter for the gas type. | - YES |  |  |  | - NO |  |  |  |  |
|  | - Each gas line equipped with separate MFC |  |  |  |  |  |  |  |  |  |
|  | 1 MFC for Ar feeding into a gas ring on the cathode pair #1. | - YES |  |  |  | - NO |  |  |  |  |
|  | 1MFC for Ar feeding into a gas ring on the cathode pair #2. | - YES |  |  |  | - NO |  |  |  |  |
|  | 1 MFC for Ar feeding into a gas ring on the 5th cathode. | - YES |  |  |  | - NO |  |  |  |  |
|  | 1 MFC for Ar for the ion source. | - YES |  |  |  | - NO |  |  |  |  |
|  | 1 MFC for Ar, for Hollow Cathode Electron source. | - YES |  |  |  | - NO |  |  |  |  |
|  | 1 MFC for O2, for the Ion source. | - YES |  |  |  | - NO |  |  |  |  |
|  | 1 MFC for N2, for the Ion source. | - YES |  |  |  | - NO |  |  |  |  |
|  | Deposition control and measurement subsystems: |  |  |  |  |  |  |  |  |  |
|  | - A) Film layer control: |  |  |  |  |  |  |  |  |  |
|  | Quartz micro balance or equivalent system. | - YES |  |  |  | - NO |  |  |  |  |
|  | System must provide at least 2x12 quartz crystal holders, dedicated to each pair of 1&2, 3&4, cathodes from the primary deposition system. | Number of quartz crystal holders, dedicated to each pair of 1&2, 3&4, cathodes from the primary deposition system: ________________ |  |  |  |  |  |  |  |  |
|  | At least one crystal sensor for cathode #5 | Number of crystal sensors for cathode #5: _____________________ |  |  |  |  |  |  |  |  |
|  | The crystals must be monitored with an appropriate monitoring system fully compatible with the main system software. | - YES |  |  |  | - NO |  |  |  |  |
|  | Crystal monitoring system supports co-deposition. | - YES |  |  |  | - NO |  |  |  |  |
|  | Crystal frequency resolution: >± 0.004 Hz at frequency 6 MHz. | Crystal frequency resolution at frequency 6 MHz: ______________ Hz |  |  |  |  |  |  |  |  |
|  | Thickness rate resolution: >± 0.005 Å. | Thickness rate resolution: __________________ Å |  |  |  |  |  |  |  |  |
|  | - B Optical film thickness and layer growth control: | - YES |  |  |  | - NO |  |  |  |  |
|  | Optical control system with range: minimal between 400 nm and 1650 nm | - YES |  |  |  | - NO |  |  |  |  |
|  | Spectral resolution at specified wavelength ranges: minimal 1 nm | - YES |  |  |  | - NO |  |  |  |  |
|  | Wavelength repeatability: minimal 0,1 nm at wavelength range 270 nm to 1050 nm; minimal 0,5 nm at wavelength range 1050 nm to 1650 nm | - YES |  |  |  | - NO |  |  |  |  |
|  | Optical control system must allow importing filter/film designs from all major commercial thin film design software equivalent to: OptiLayer, Film Star, Essential MacLeod, TFCalc. | - YES |  |  |  | - NO |  |  |  |  |
|  | Thin film deposition system control and thin film layer control: |  |  |  |  |  |  |  |  |  |
|  | - Control software must enable automatically control of the deposition process. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Programable automated control system for deposition sequence. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Main control system must be fully compatible with the proprietary software from the film thickness and layer growth control system. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Remote control provision. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Logging of various parameters, such as bias voltages and currents for sputter cathodes, ion source, electron source, gas flows, temperature, pressure, and film thickness. | - YES |  |  | - NO |  |  |  |  |  |
|  | Moving subsystem for moving the entire thin film deposition system: |  |  |  |  |  |  |  |  |  |
|  | - Integrated wheels or an adequate moving subsystem, which should not, cumulatively with the thin film deposition system, exceed the clean room ceiling height of 2100 mm. | - YES |  |  |  |  | - NO |  |  |  |
|  | Acceptance criteria: |  |  |  |  |  |  |  |  |  |
|  | - Demonstration of a narrow bandpass filter with FWHM 3,5 nm and the maximum transmission greater than 80 % at wavelength between 500 nm and 600 nm. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Demonstration and production of a multilayer film consisted of GeO2 and SiO2 layers. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Demonstration and production of a dielectric reflector on a fused silica substrate with average reflection greater than 99 % for S- and P- polarization for angles of incidence from 0 to 45°. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Demonstration of an anti-reflective coating at the end of a standard optical fiber, with reflectance minimal 1 % in the wavelength range between 900 nm and 2300 nm. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Sputtering targets for demonstrating the samples must be included in the system scope. | - YES |  |  | - NO |  |  |  |  |  |
|  | - Demonstration of operation of the secondary deposition system by deposition of Al2O3 on a wafer with diameter minimum 100 mm. | - YES |  |  | - NO |  |  |  |  |  |

| Place and date: | Seal: | Authorized signature: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.