Project title: Upgrading National Research Infrastructures – RIUM

COMPLIANCE DECLARATION WITH MINIMUM TECHNICAL REQUIREMENTS OF THE OFFERED EQUIPMENT

Name of the tenderer: Servohydraulic machines for uniaxial and biaxial flexible standalone actuators for static and dynamic testing system at low and higher temperature

Address of the tenderer:

Subject of the public procurement: Purchase of servohydraulic machines for uniaxial and biaxial flexible standalone actuators for static and dynamic testing system at low and higher temperature

Designation of the public procurement: 302/62 – RIUM33-FS11-1/2020

INSTRUCTIONS:

- This Annex No. 2 must be completed and signed. The tenderer must indicate the actual technical specifications of the equipment offered in line with all requirements.

- The tenderer shall upload this Annex No. 2 to the e-JN information system in the section “Other Annexes”.

We declare that the offered equipment fully meets the minimum technical requirements of the client for the public procurement titled "Purchase of servohydraulic machines for uniaxial and biaxial flexible standalone actuators for static and dynamic testing system at low and higher temperature”, and that we are ready to realize the order according to the following specification:

|  | EQUIPMENT |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  |  | DATASHEET |
| 1 | ±250kN Axial Dynamic Testing System |  |  |  |  |
|  | The load frame shall be of a two-column construction and be fatigue rated, capable of testing in tensile and compressive modes with actuator with capacity of at least ±250kN and stroke of 150 mm (or ±75 mm) | Loading capacity: _______________ kN |  |  |  |
|  |  | Stroke: _______________ mm |  |  |  |
|  | The load frame must have a stiffness of at least 585kN/mm (measured at 1m crosshead height) | Stiffness: _______________kN/ mm |  |  |  |
|  | Maximum high of frame must be lower than 3500 mm | Maximum high of frame: _____________mm |  |  |  |
|  | Vertical distance between the crosshead and platen to be at least 1970 mm | Vertical distance: _________ mm |  |  |  |
|  | Distance between columns should be at least 660mm | Horizontal distance: _________ mm |  |  |  |
|  | The load frame shall be equipped with a hydraulically operated system for raising, lowering, and clamping the moving crosshead | - YES | - NO |  |  |
|  | The load frame must feature a frame-mounted jog control handset with both coarse and fine actuator position adjustment | - YES |  | - NO |  |
|  | The load frame must feature a set up mode, restricting the actuators velocity to at least 10mm/s during test set-up | Actuator velocity during test set-up at least: _________ mm/s |  |  |  |
|  | The system must meet the requirements of EN ISO 13849-1 performance level D (PLd) | - YES | - NO |  |  |
|  | Anti rotating device for actuator shoul be supplied | - YES | - NO |  |  |
|  | An alignment fixture with load capacity of 400 kN must be included with the machine along with software to measure and adjust the alignment of the system | - YES | - NO |  |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 1 | Actuator, Manifold and Servo-valve for 250kN |  |  |  |
|  | The linear actuator shall have a static & dynamic rating of at least 250kN at 210bar and be fatigue rated for testing in tensile, compressive, and through-zero modes where 75mm stroke displacement starts | Loading capacity: _______________ kN |  |  |
|  |  | Stroke: _______________ mm Maximum pressure: ____________ bar |  |  |
|  | The linear actuator must be fitted with an integrated LVDT for actuator position measurement with a measurement accuracy at least of ±0.2% of LVDT full travel | Accuracy at least of _______________% |  |  |
|  | The linear actuator should be fitted with a manifold and supplied with servo-valves rated for oil flows of at least 40 lpm and pressure 210 bars | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 2 | Load Cell for ±250kN |  |  |  |
|  | The system must be supplied with a linear fatigue rated load cell with a static and fatigue rating of ±250kN as integral part of system | Loading capacity of cell: _______________ kN |  |  |
|  | Load cell should be possible to use as an integral part of the force measurement system and also in load mode control test of a testing machine. | - YES | - NO |  |
|  | The load cell must feature an accelerometer mounted directly on the linear load axis to measure inertial loads generated due to the moving fixture mass during a cyclic test | - YES | - NO |  |
|  | The control system must be capable of reducing inertia-induced loads from the load signal in real time and must feature a fully automatic dynamic load-compensation set-up routine | - YES | - NO |  |
|  | The load measurement system shall meet ISO 7500-1 part 1 Class 0.5 down to 1/250th of loadcell full scale | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 3 | Load Cell for ± 5kN |  |  |  |
|  | The system must be supplied with a linear fatigue rated load cell with a static and fatigue rating of ±5 kN | - YES | - NO |  |
|  | Load cell should be possible to use as an integral part of the force measurement system and also in load mode control test of a testing machine. Maximum load of load cell is 30% of overload (e.g. up to 6,5 kN) | - YES | - NO |  |
|  | Loadcell should be supplied with accessories for mounting in grips | - YES | - NO |  |
|  | The load measurement system shall meet ISO 7500-1 part 1 Class 0.5 down to 1/250th of loadcell full scale | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 4 | Grips for ±250kN |  |  |  |
|  | The system must be supplied with hydraulic wedge action grips suitable for round and flat specimens | - YES | - NO |  |
|  | Grips for gripping round (from dia 5mm to 34mm) and flat specimens (from 1mm to 42mm), for static tensile and compression as well as full reverse-stress testing up to the full load capacity of the system ±250kN | Grips for gripping round from dia from _____to ______ mm |  |  |
|  |  | Grips for gripping flat specimens from _____to ______ mm |  |  |
|  | The control system must prevent the operation of the grips (opening or closing) either whilst a test is running or whilst in load or strain control | - YES | - NO |  |
|  | The mechanical grips for fatigue rated testing up to ±100kN must supplied. Grips must be suitable for use in a temperature chamber from -70 to +300°C | - YES | - NO |  |
|  | The system should be supplied with mechanical wedge action grips suitable for gripping round (from dia 6 mm to 20mm) and flat specimens (from 0 mm to 16 mm) for static tensile and compression as well as full reverse-stress testing up to ±100kN | - YES | - NO |  |
|  | Two of water-cooled pullrods must be supplied for testing at high temperature up to +300°C | - YES | - NO |  |
|  | Grips suitable for fracture mechanics testing of CT specimens up to a width of 25mm must be supplied | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 5 | Control System – control unit |  |  |  |
|  | The controller should have at least a 6 term control loop. (i.e. P, I, D, Lag, Feed Forward, Notch filter) | - YES | - NO |  |
|  | The controller must have channels for: position, axial load, and two optional channels (e.g. elongation or strain). These must be available for control and data acquisition | - YES | - NO |  |
|  | Data from all control and aquisition channels should be recorded synchronously and should be capable of at least the full 10kHz loop update rate across all channels, irrespetive of how many channels are used. | - YES | - NO |  |
|  | The controller must have at least one analogue input for an external analogue waveform drive signal, and at least four assignable analogue outputs | - YES | - NO |  |
|  | All supplied transducers shall be automatically recognised and calibrated by the controller | - YES | - NO |  |
|  | The controller must prevent any test proceeding with an uncalibrated transducer to prevent incorrect data being gathered | - YES | - NO |  |
|  | The controller must have a minimum of two limit detectors per connected transducer and must have a detection time of 1ms or better. The limit detectors must not be reliant in any way on the connection of a personal computer | - YES | - NO |  |
|  | The control system must include both manual and automatic tuning functions | - YES | - NO |  |
|  | The automatic tuning function should automatically determine load string stiffness based on a singular tuning value calculated from a single quasi-static ramp to optimise the control loop for a particular specimen. | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 6 | Software and PC Hardware for ±250kN |  |  |  |
|  | The system should have software supplied with the system to cover Dynamic Fatigue, Low Cycle Fatigue, Quasi-static, and Fracture mechanics testing | - YES | - NO |  |
|  | The PC user interface should clearly indicate to the user the current control mode, and clearly indicate from a distance the test status, e.g. running, finished, limit tripped or actuator power off | - YES | - NO |  |
|  | Factory integration of supplied Personal Computor with intel core i7 CPU, Keyboard and wide Screen Flat Panel Monitor at least 22'' size | - YES | - NO |  |
|  | Any PC offered with the system shall utilise the Windows 10 or newer Professional operating system and must feature a 22-inch LCD flat panel monitor or larger | - YES | - NO |  |

|  | Especial requirements: |  |  |  |
| --- | --- | --- | --- | --- |
|  | 1. Personal computers shall comply with the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards | - DA | - NE |  |  |
|  | 2. Personal computers must be designed in such a way that the hard disk and memory are easily accessible using universally available tools (i.e. with a screwdriver, spatula, pliers, or tweezers) and are replaceable | - DA | - NE |  |  |
|  | 3. The volume of the personal computer must be measured in accordance with the SIST EN ISO 7779 standard and must not exceed 26 dB at the time of writing to the hard disk. The data must be labelled with the designation LpAm in dB | - DA | - NE |  |  |
|  | 4. Screens shall meet the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. | - DA | - NE |  |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 7 | Fatigue Test Software |  |  |  |
|  | Systems should come with independent test software for different test types (e.g. quasi-static, fatigue, fracture mechanics etc.) | - YES | - NO |  |
|  | Test software should communicate with temperature chamber and be able to adjust and control the temperature via the test method where required | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 8 | Dynamic test software |  |  |  |
|  | Shall be able to generate the following waveforms and provide looping capability: -Sine, haversine -Square, haversquare -Triangle, havertriangle -Ramp -Hold -User defined waveforms as either turning points or sampled data -Trapezoid waveforms with user defined ramp rates and hold times | - YES | - NO |  |
|  | -Test software should communicate with temperature chamber and be able to adjust and control the temperature via the test method where required | - YES | - NO |  |
|  | Dynamic software should have a quick test functionality, to automatically create a fatigue test method in minutes | - YES | - NO |  |
|  | During a test the software shall have user customisable graphs of test data capable of real time auto-scaling and manual zoom in/out | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 9 | Static test software |  |  |  |
|  | Static Software should have quick test functionality to allow running a testing in seconds with only one parameter | - YES | - NO |  |
|  | Static software should contain a methods library of pre-built complying to international standards. Including ASTM E8, E21, E646, E517 and ISO 6891-1 | - YES | - NO |  |
|  | Static software should come with the ability to view and record live video during a test and save it to the test results | - YES | - NO |  |
|  | Static software should have different user levels and the ability to set up guided or prompted test workflows for different operators | - YES | - NO |  |
|  | Static software should have its own analysis & reporting functionality | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 10 | Fracture mechanics test software |  |  |  |
|  | Fracture software should have the capability to run pre-cracking, da/dN K1c and J1c tests | - YES | - NO |  |
|  | Fracture software must have methods to comply to the latest standards: K1c ASTM E399, B645, ISO 12135, BS7448-1 and for J1c ASTM E1820 | - YES | - NO |  |
|  | Fracture software should have the ability to update the just the standards compliance without having to update the main software package | - YES | - NO |  |
|  | Fracture software must support the following specimen types: Compact Tension (CT), Single Edge Notch Bend (SENB), Disc Shape CT, Arc-Shape Tension | - YES | - NO |  |
|  | Fracture software should have its own analysis & reporting functionality | - YES | - NO |  |
|  | Fracture software should run validity checks of the specimen parameters and test results against the test standard | - YES | - NO |  |
|  | Fracture software should have different user levels and the ability to set up guided test workflows | - YES | - NO |  |
|  | Fracture softwere should be compatible with the Matelect 2DP potential-based availability growth measurement system by present Matelect DCM-2 device for J1C and da/dN tests | - YES | - NO |  |
|  | Test cam for video recording and playback module | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 11 | Hydraulic Supply & Hoses |  |  |  |
|  | A hydraulic power unit must be provided with a minimum flow rate of 78lpm at 210bar. Supplied by 400 V, 50Hz and 3 phase operation | Miximum flow rate __________ lpm |  |  |
|  |  | Maximum pressure of HPU _________ bar |  |  |
|  | The unit must be air cooled and fitted with an interface that include a digital display of oil pressure and temperature and include protection devices for oil temperature, oil pressure, oil level, oil filter condition and motor temperature | - YES | - NO |  |
|  | Noise levels must not exceed 63 dB (A) | Maximum noise level _____________dB |  |  |
|  | HPU should be supplied with fully integrated acoustic attenuation system | - YES | - NO |  |
|  | The unit must be fitted with a 2 micron pressure and return line filter | Minimum size of particles filtering by pressure and return filters _____ microns |  |  |
|  | All hoses for interfacing between the machine and HPU must be supplied, as follows 6 m hoses from the HPU to the Distribution mainfold, 3 m hoses between distribution mainfold to testing system. | - YES | - NO |  |
|  | Oil and filters for hydraulic power unit must be sustained | - YES | - NO |  |
|  | Floor standing hydraulc distribution mainfold with a single inlet ad up to three outlets | - YES | - NO |  |
|  | Hydraulic Hydralic power unit should be equiped with pressure accumulator | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 12 | Extensometry |  |  |  |
|  | System must come with a clip-on extensometer, clip-on COD gauges and a non-contacting extensometer | - YES | - NO |  |
|  | The clip-on extensometer should have a 25 mm gauge length with -2.5 mm to +12,5 travel | - YES | - NO |  |
|  | The clip-on extensometer should come with gauge length extenders for 25mm and 37.5mm | - YES | - NO |  |
|  | The clip-on extensometer should be suitable for temperatures from -80°C to +200°C | Minimum testing temperature ___________ °C |  |  |
|  |  | Maximum testing temperature ___________ °C |  |  |
|  | There should be COD gauges provided with 5mm gauge length with 2mm travel | - YES | - NO |  |
|  | There should be COD gauges provided with 10mm gauge length and 4mm travel | - YES | - NO |  |
|  | The COD gauges must comply with ASTM E399 and be suitable for temperatures from -80°C to +200°C | Minimum testing temperature ___________ °C |  |  |
|  |  | Maximum testing temperature ___________ °C |  |  |
|  | The non-contacting extensometer should be supplied. A non-contacting video extensometer based on Digital Image Correlation (DIC) measurement system for Field of Viwe (FOV) 240 mm. | - YES | - NO |  |
|  | Standard Field of View Lens with 16mm focal length should be used for FOV at least of 240 mm | - YES | - NO |  |
|  | The non-contacting video extensometer should be able to automatically measure the gauge length of specimens | - YES | - NO |  |
|  | The non-contacting video extensometer should be able to perform strain control for fatigue tests, up to 20Hz | Maximum frequencies for fatigue control test ____Hz |  |  |
|  | The non-contacting video extensometer must have an accuracy and resolution of at least 2µm when used for fatigue testing | An accuracy and resolution of fatigue testing at least _____µm |  |  |
|  | The non-contacting video extensometer must have an accuracy and resolution of at least 0.5µm when used for quasi-static testing | An accuracy and resolution of quasi-static testing at least ____µm |  |  |
|  | All extensometers above should be capable of strain measurement and control for both quasi-static and fatigue testing. Please indicate if any are only suitable for quasi-static testing or testing below 5Hz | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 13 | Temperature Chamber |  |  |  |
|  | The system must be supplied with an environmental chamber capable of reaching temperatures from -80 to +350°C | Minimum testing temperature ___________ °C |  |  |
|  |  | Maximum testing temperature ___________ °C |  |  |
|  | The chamber must be controllable by temperature controller from the test software including changing temperature, setting dwell and soak times, and triggering events | - YES | - NO |  |
|  | Using Liquid N2 as cooling liquid | - YES | - NO |  |
|  | The chamber internal dimensions should be at least 760x400x400mm (HxWxD) | Minimum inside high ______mm |  |  |
|  |  | Minimum inside width ______ mm |  |  |
|  |  | Minimum inside depth ______mm |  |  |
|  | The chamber external dimensions must not exceed 950x550x900mm (HxWxD) | Maximum outside high ______mm |  |  |
|  |  | Maximum outside width ______ mm |  |  |
|  |  | Maximum outside depth ______mm |  |  |
|  | The chamber must sit on a roller carriage to allow easy insertion and removal from the test space | - YES | - NO |  |
|  | Frame Mounting Bracket and support legs | - YES | - NO |  |
|  | LN2 coolant Hose set | - YES | - NO |  |
|  | Piston rod (250 mm length) extented for use on testing machine with actuator in paltten, enables at lest of ±250kN fatigue rate hydraulic grips | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 14 | SYSTEM WITH ±100 kN for z and x-y axial loading Load Frame |  |  |  |
|  | The load frame shall be of a two-column construction and be fatigue rated, capable of testing in tensile and compressive modes with a capacity of at least ±100kN and 150mm of usable stroke | Loading capacity: _______________ kN |  |  |
|  |  | Stroke: _______________ mm |  |  |
|  | The load frame must have a stiffness of at least 390kN/mm (measured at 1m crosshead height) | Stiffness: _______________kN/ mm |  |  |
|  | Vertical distance between the crosshead and platen to be at least 1630mm | Maximum vertical distance ______ mm |  |  |
|  | Distance between columns should be at least 560mm | Minimum distance between columns ______ mm |  |  |
|  | The load frame shall be equipped with a hydraulically operated system for raising, lowering, and clamping the moving crosshead | - YES | - NO |  |
|  | The load frame shall detect the upper crosshead being unclamped and prevent a test being started | - YES | - NO |  |
|  | The load frame must feature a frame-mounted, emergency stop in full compliance with current regulations. The emergency stop should not be on a cable that allows it to be moved away from the test-space | - YES | - NO |  |
|  | The load frame must feature a frame-mounted jog control handset with both coarse and fine actuator position adjustment | - YES | - NO |  |
|  | The load frame must feature a set up mode, restricting the actuators velocity to at least 10mm/s during test set-up | - YES | - NO |  |
|  | The ±100kN system must meet the requirements of EN ISO 13849-1 performance level D (PLd). Systems which meet a lower performance level will not be considered | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 15 | Actuator, Manifold and Servo-valve for 100kN |  |  |  |
|  | The linear actuator shall have a static & dynamic rating of at least 100kN at 210bar and be fatigue rated for testing in tensile, compressive, and through-zero modes | Loading capacity: _______________ kN |  |  |
|  |  | Stroke: _______________ mm |  |  |
|  | The linear actuator shall have a minimum of 150 mm total stroke (75mm) and feature hydrostatic bearings to aid with protection from side loads | accuracy: _______________% |  |  |
|  | The linear actuator must be fitted with an integrated LVDT for actuator position measurement with a measurement accuracy of at least ±0.2% of LVDT full travel | accuracy: _______________% |  |  |
|  | The linear actuator should be fitted with a manifold and supplied with servo-valves rated for oil flows of at least 40 lpm | Flow _______lpm |  |  |
|  | The manifold shall be provided with an integrated hydraulic shut-off, so the machine can be isolated from the hydraulic circuit without the need for additional hydraulic control modules | - YES | - NO |  |
|  | Piston rod extender for use on servohydraulic machine to reduce the minimum daylight between load cell and actuator | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 16 | Software and PC Hardware for ±250kN |  |  |  |
|  | The system should have software supplied with the system to cover Dynamic Fatigue, Low Cycle Fatigue, Quasi-static, and Fracture mechanics testing | - YES | - NO |  |
|  | The PC user interface should clearly indicate to the user the current control mode, and clearly indicate from a distance the test status, e.g. running, finished, limit tripped or actuator power off | - YES | - NO |  |
|  | Factory integration of supplied Personal Computor with intel core i7 CPU, Keyboard and wide Screen Flat Panel Monitor at least 22'' size | - YES | - NO |  |
|  | Any PC offered with the system shall utilise the Windows 10 or newer Professional operating system and must feature a 22-inch LCD flat panel monitor or larger | - YES | - NO |  |

|  | Especial requirements: |  |  |  |
| --- | --- | --- | --- | --- |
|  | 1. Personal computers shall comply with the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards | - DA | - NE |  |  |
|  | 2. Personal computers must be designed in such a way that the hard disk and memory are easily accessible using universally available tools (i.e. with a screwdriver, spatula, pliers, or tweezers) and are replaceable | - DA | - NE |  |  |
|  | 3. The volume of the personal computer must be measured in accordance with the SIST EN ISO 7779 standard and must not exceed 26 dB at the time of writing to the hard disk. The data must be labelled with the designation LpAm in dB | - DA | - NE |  |  |
|  | 4. Screens shall meet the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. | - DA | - NE |  |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 17 | Load Cell for ±100kN |  |  |  |
|  | The system and both actuators must be supplied with a linear fatigue rated load cell with a static and fatigue rating of ±100kN as integral part | - YES | - NO |  |
|  | The control system must be capable of reducing inertia-induced loads from the load signal in real time and must feature a fully automatic dynamic load-compensation set-up routine | - YES | - NO |  |
|  | The load measurement system shall meet ISO 7500-1 part 1 Class 0.5 down to 1/250th of loadcell full scale | - YES | - NO |  |
|  | The load cell must have a minimum 300% overload capacity without mechanical damage | Overload capacity ___________ % |  |  |
|  | The load cell must feature an accelerometer mounted directly on the linear load axis to measure inertial loads generated due to the moving fixture mass during a cyclic test | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 18 | Grips for ±100kN |  |  |  |
|  | The system must be supplied with hydraulic wedge action grips suitable for round and flat specimens | - YES | - NO |  |
|  | Grips for gripping round (dia 6 to 20mm) and flat specimens (from 1 to 21mm), for static tensile and compression as well as full reverse-stress testing up to the full load capacity ±100kN of the system | Grips for gripping round from dia from _____to ______ mm |  |  |
|  |  | Grips for gripping flat specimens from _____to ______ mm |  |  |
|  | The control system must prevent the operation of the grips (opening or closing) either whilst a test is running or whilst in load or strain control | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 19 | Hydraulic power unit, supply & Hoses |  |  |  |
|  | A hydraulic power unit must be provided with a minimum flow rate of 48lpm at 210bar. | Maximum flow rate __________ lpm |  |  |
|  |  | Maximum pressure of HPU _________ bar |  |  |
|  | The unit must be air cooled and fitted with an interface that include a digital display of oil pressure and temperature and include protection devices for oil temperature, oil pressure, oil level, oil filter condition and motor temperature | - YES | - NO |  |
|  | Noise levels must not exceed 73 dB (A) | Maximum noise level _____________dB |  |  |
|  | The unit must be fitted with a 2 micron pressure and return line filter | Minimum size of particles filtering by pressure and return filters _____ microns |  |  |
|  | All hoses for interfacing between the machine and HPU must be supplied in length of 12 m between HPU and testing actuator | - YES | - NO |  |
|  | Oil and filters for HPU must be supplied | - YES | - NO |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 20 | Software and PC Hardware for flexible actuators 2 x for ±100kN |  |  |  |
|  | The integratl controller for both actuators must have its own dedicated PC system. | - YES | - NO |  |
|  | The PC user interface should clearly indicate to the user the current control mode, and clearly indicate from a distance the test status, e.g. running, finished, limit tripped or actuator power off | - YES | - NO |  |
|  | Factory integration of supplied Personal Computor with intel core i7 CPU, Keyboard and wide Screen Flat Panel Monitor at least 22'' size | - YES | - NO |  |
|  | Any PC offered with the system shall utilise the Windows 10 or newer Professional operating system and must feature a 22-inch LCD flat panel monitor or larger | - YES | - NO |  |

|  | Especial requirements: |  |  |  |
| --- | --- | --- | --- | --- |
|  | 1. Personal computers shall comply with the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards | - DA | - NE |  |  |
|  | 2. Personal computers must be designed in such a way that the hard disk and memory are easily accessible using universally available tools (i.e. with a screwdriver, spatula, pliers, or tweezers) and are replaceable | - DA | - NE |  |  |
|  | 3. The volume of the personal computer must be measured in accordance with the SIST EN ISO 7779 standard and must not exceed 26 dB at the time of writing to the hard disk. The data must be labelled with the designation LpAm in dB | - DA | - NE |  |  |
|  | 4. Screens shall meet the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. | - DA | - NE |  |  |

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 21 | Flexible 2 actuators with at least of ±100kN for 2D x-y testing system |  |  |  |
|  | Each actuator with at least 40 litre/min and equiped with servo valve capacity ±100kN and stroke ±75mm pedestal based actuator must share the same controller as for 100 kN testing machine | Capacity of flow ________ litre/min |  |  |
|  |  | Capacity of servo valve ±_______kN |  |  |
|  |  | Stoke ± _______mm |  |  |
|  | Each actuator should be supplied with at least ±100 kN fatigue rated mechanical wedge grips | - YES | - NO |  |
|  | Each actuator has consists ped base actuator rig system | - YES | - NO |  |
|  | Floor standing distribution mainfold must be supplied | - YES | - NO |  |
|  | Hoses set of 12 m length from mainfold to floor standing distribution mainfold have to be supplied | - YES | - NO |  |
|  | 2 hoses 6 m length from distribution mainfold to each standalone actuator | - YES | - NO |  |

| Place and date: | Seal: | Authorized signature: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

For the Computer Software and Hardware (PC) item, the tenderer shall indicate a proof for meeting the requirements of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No 51/17 and 64/19), as stated in Point 3 of the Technical Specifications that are part of these procurement documents.

ENERGY STAR standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 6.1, and updates are available at http://www.eu-energystar.org/specifications.htm.

Hard Drive Operating.

LpAm (Low power AM) is the designation for the sound emission pressure.

Energy Star Standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 7.0, and updates are available at http://www.eu-energystar.org/specifications.htm.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

For the Computer Software and Hardware (PC) item, the tenderer shall indicate a proof for meeting the requirements of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No 51/17 and 64/19), as stated in Point 3 of the Technical Specifications that are part of these procurement documents.

ENERGY STAR standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 6.1, and updates are available at http://www.eu-energystar.org/specifications.htm.

Hard Drive Operating.

LpAm (Low power AM) is the designation for the sound emission pressure.

Energy Star Standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 7.0, and updates are available at http://www.eu-energystar.org/specifications.htm.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.

For the Computer Software and Hardware (PC) item, the tenderer shall indicate a proof for meeting the requirements of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No 51/17 and 64/19), as stated in Point 3 of the Technical Specifications that are part of these procurement documents.

ENERGY STAR standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 6.1, and updates are available at http://www.eu-energystar.org/specifications.htm.

Hard Drive Operating.

LpAm (Low power AM) is the designation for the sound emission pressure.

Energy Star Standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 7.0, and updates are available at http://www.eu-energystar.org/specifications.htm.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.