Naziv projekta: Nadgradnja nacionalnih raziskovalnih infrastruktur – RIUM

Project title: Upgrading National Research Infrastructures – RIUM

Tehnične specifikacije ZA SERVO-HIDRAvLIČNA STROJA ZA enoosno in FLEKSIBILNO dvoOSNO preizkuševališče s samostoječima aktuatorjema za STATIČNA IN dinamičnA testiranjA na nizki in povišini temperaturi

Technical specifications of SERVOHYDRAULIC machines for uniaxial and BIAXIAL flexible standalone actuators for STATIC AND dynamic testing system at low and higher temperature

- PREDMET JAVNEGA NAROČILA 						1. SUBJECT-MATTER OF THE PUBLIC CONTRACT

| Predmet javnega naročila je nabava opreme za servohidravlično enoosno 250 kN in fleksibilno dvoosno 100 kN dinamično preizkuševališče na nizki in povišani temperaturi, vključno z dostavo, vnosom, namestitvijo, montažo, konfiguracijo, inštalacijo in zagonom naprave ter usposabljanjem, skladno s tehničnimi specifikacijami. Lokacija dostave in vnos opreme: Univerza v Mariboru, Fakulteta za strojništvo, Smetanova 17, 2000 Maribor, Slovenija, Laboratorij za strojne elemente in konstrukcije, objekt A, prostor A-002. Rok dobave: najpozneje 240 dni od dneva pričetka veljavnosti pogodbe. Rok namestitve in izvedbe usposabljanja: najkasneje nadaljnjih 30 dni od roka dobave. Garancijski rok: najmanj 2 leti od dneva podpisa prevzemnega zapisnika. Za vso ponujeno opremo velja, da mora zanjo v času garancijskega obdobja proizvajalec opreme zagotavljati dostop do vseh posodobitev programske opreme brez dodatnih stroškov. Vsi popravki, servisni paketi ter nadgradnje in podpora morajo biti od istega proizvajalca strojne oz. programske opreme. V času garancijskega obdobja mora ponudnik za vso opremo nuditi tehnično podporo. Tehnična podpora mora vključevati svetovanje pri nastavitvi parametrov, konfiguraciji sistema in pomoč pri izvedbi posameznih procesov. Tehnična podpora se lahko izvaja preko oddaljenega dostopa. Ponudnik mora za opremo zagotoviti: - da so vsi ponujeni moduli/komponente združljivi in medsebojno povezljivi, - da je mogoče vse zahtevane module/komponente povezati v enoten in brezhibno delujoč sistem v skladu s tehničnimi specifikacijami, - odpraviti vse napake, ki bi se pojavile zaradi morebitne medsebojne nezdružljivosti ali nepovezljivosti oz. zaradi neustrezne povezave v enoten, brezhibno delujoč sistem, - servis in vzdrževanje, vključno z oskrbo z nadomestnimi deli za obdobje najmanj 5 let od montaže strojne oz. programske opreme. Usposabljanje: Dobavitelj mora izvesti osnovno začetno usposabljanje, ki vključuje strojne in programske module na dobavljenem predmetu pogodbe za najmanj 4 (štiri) osebe s strani uporabnika v obsegu 5 (pet) delovnih dni. V ponudbi mora biti zajet ves droben material, ki je potreben za priključitev naprav, tudi če ni izrecno naveden. | Subject-matter of the public contract is the purchase of an Equipment for servohydraulic uniaxial 250kN and 2 standalone 100kN actuators at lower and higher temeprature, including the delivery, evaluation, placement, assembly, configuration, installation, and starting of the systems, in line with technical specifications. Delivery location and placement: University of Maribor, Faculty of Mechanical Engineering, Smetanova 17, 2000 Maribor, Slovenia, Laboratory for Machine parts and structures, Building A, room A-002. Delivery time: at the latest 240 days from the day of entry into force of the contract. Deadline for installation and training: no later than 30 days from the delivery date The guarantee period: At least 2 years from the day of signature of the take-over protocol. The equipment manufacturer is obliged to provide access to all software updates at no additional cost during the guarantee period. All patches, service packs, upgrades and support must be provided by the hardware or software of same manufacturer. During the warranty period the supplier must provide technical support. Technical support must include consulting on setting parameters, use of the machine, configuring the system, and assisting in the implementation of individual processes, etc. Technical support can be provided through remote access. The tenderer of the equipment must guarantee: - that all modules/components offered are compatible and interconnected, - that all required modules/components can be integrated into a single and seamlessly functioning system in line with technical specifications, - to eliminate all errors resulting from possible incompatibility or mismatch, or from improper connection to a single, seamlessly functioning system, - Service support and maintance with spare parts must be provided at least 5 years after the warranty period. Training: The supplier must carry out basic initial training, which includes hardware and software modules on the delivered subject of the contract for at least 4 (four) persons within the scope of 5 (five) working days. The tender must include all the fine material, necessary for connecting the devices, even if it is not explicitly stated. |
| --- | --- |

| - TEHNIČNE ZAHTEVE # Servo-hidravličnI STROJI ZA enoosno in FLEKSIBILNO dvoOSNO preizkuševališče s samostoječima aktuatorjema za STATIČNA IN dinamična testiranjA na nizki in povišini temperaturi Univerza v Mariboru naroča opremo za servohidravlično enoosno ±250 kN in dvoosno ±100 kN dinamično preizkuševališče na nizki in povišani temperaturi, ki omogoča: - Izvajanje enosnih statičnih in dinamičnih testov na samostoječem servohidravličnem stroju s kapaciteto do ±250kN, hodom do 150 mm na sobni temperaturi kot tudi na temperaturi od -80°C do +315°C do kapacitete ±100 kN in z hidravličnimi čeljustmi od -40° do +177°C. - Izvajanje dvoosnih statičnih in dinamičnih testov na 2D preizkuševališču z dvema aktuatorjema s servohidraličnim krmiljenjem s kapaciteto do ±100 kN in hodom do 150 mm na sobni temperaturi. - Izvajanje dvoosnega statičnega in dinamičnega testa v x-y ravnini z dvema aktuatorjema z dodatnim obremenjevanjem v z-smeri (npr. prebojnim testom) na servohidravličnem stroju s kapaciteto do ±100 kN in hodom do 150 mm. - Sistem mora biti sposoben izvajati teste z računalniškim reguliranjem in zajemanjem podatkov ter obdelovati izmerjene merilne vrednosti za določevanje nateznih mehanskih lastnosti v skladu s standardom ASTM E8M, lomne žilavosti za krhke in duktilne materiale v skladu s standardom ASTM E 399, ASTM E 1820 in hitrosti rasti utrujenostne razpoke v skladu s standardom ASTM E 647. Sistem mora biti hkrati sposoben izvajti tudi brezkontaktno meritev rastezka. Oprema za servohidravlično enoosno ±250 kN in dvoosno ±100 kN dinamično preizkuševališče na nizki in povišani temperaturi mora biti modulno sestavljena iz več medseboj avtonomno delojočih sistemov, ki morajo biti medseboj kompatibilni glede upravljanja, testiranja in programske opreme kot tudi zamenjave posameznih komponent. Čeprav lahko delujejo posamezni aktuatorji avtonomno, morajo omogočiti krmiljenje strojev ter zajemanje podatkov in obdelavo na isti programski platformi. Osnovo preizkuševališča predstavlja: - en servohidravlični preizkuševalni stroj s dvosteberskim okvirjem in sicer s kapaciteto vsaj ±250kN in hodom vsaj 150 mm (±75 mm) prostim razmikom med stebri vsaj 650 mm ali več in višino med zgornjo ploščo in jarmom vsaj 1970 mm, tako da je prosta višina med čeljustmi vsaj 930 mm, ko je aktuator v nultem položaju, - en servohidravlični preizkuševalni stroj z dvosteberskim okvirjem s kapaciteto vsaj ±100kN in hodom vsaj 150 mm z medsebojnim razmikom med stebri vsaj 550 mm ali več in višino med zgornjo ploščo in jarmom vsaj 1650 mm, - dva fleksibilna servohidravlična aktuatorja s kapaciteto vsaj ±100kN in hodom vsaj 150 mm (±75 mm) s sondami za silo ±100kN !!in mehanskimi čeljustmi z enotnim kontrolerjem in računalnikom za zajemanje podatkov dvoosnega testa. - temperaturno ambientna komora za temperaturno območje od -120°C do +315°C s notranjim prostorom vsaj 760x400x400 mm (višinaxglobinaxširna) - hidravlični agregat z zračnim hlajenjem z močjo vsaj 30 kW in pretokom olja vsaj 78 l/min in pritiskom vsaj 205 barov. - hidravlični agregat z zračnim hlajenjem z močjo vsaj 18 kW in pretokom olja vsaj 48 l/min in pritiskom vsaj 205 barov. - videoekstenziometer za avtomatično brezkontaktno merjenje raztezka znotraj vidnega območja do 240 mm, ki se lahko kasneje razširi do 500 mm vidnega območja s pomočjo DIC (Digital Image Correlation) metode. - dodatna sonda za merjenje sile do ±5kN z možnostjo povezave na programsko platformo za krmiljenje in obdelavo meritev. - dinamični ekstenziometer z začetno merilno razdaljo 25mm in hodom od -2.5mm do +12.5mm za meritve v temperaturnem območju od -80˘C do +177°C. - Računalniška in programska oprema za krmiljenje preizkusov in zajemanje meritev sile in premikov ter za avtomatsko obdelavo podatkov in izdalavo poročil. Programska oprema mora vsebovati modul za aplikacijo na kovinah v skladu s standardi za statična in dinamična testiranja za testiranja na utrujanje in lomno žilavost, tako da zajema KIC, JIC, CTOD; DCOD; FCG ter lomno in utrujenostno analizo v skladu s standardi: -ASTM E606 in D3479, za nizkoziklična utrujanja LCF; -ASTM E466 in D3479 ta visokociklična utrujanja HCF; -podloge a LCF in prehod do HCF; -ASTM E399 (podpira FFC(T) in SE(B) vzoce, ter - ISO 12135 (podpira FFC(T), M(T) in SE(B) vzorce, - ASTM E1820, podpira FFC(T), M(T) in SE(B) vzorce. - Aparatura in program za natačno nastavitev centričnosti in kotnih odstopanj z merilnimi lističi in vzorcem za tlačni preizkus do sile vsaj 250kN. - Tovarniško integrirana računalniška oprema z operacijskim sistemom MS Windows 10 Professional, tipkovnico in LED zaslonom vsaj 22”. V nadaljevanju so podane minimalne tehnične zahteve. Ponudba mora izpolnjevati ali presegati vse minimalne tehnične zahteve. |  |  | - TECHNICAL SPECIFICATIONS # Servohydraulic machines for uniaxial and BIAXIAL flexible standalone actuators for STATIC AND dynamic testing system at low and higher temperature University of Maribor is purchasing equipment for uniaxial ±250 kN and 2 standalone ±100kN actuators testing system at low and higher temperature which facilitates: - Performing single static and dynamic tests on a stand-alone servo-hydraulic machine with a capacity of up to ± 250 kN, stroke up to 150 mm at room temperature as well as at a temperature of -80 ° C to + 315 ° C up to a capacity of ± 100 kN and with hydraulic grips from -40° do +177°C. - Carrying out biaxial static and dynamic tests on a 2D test site with two actuators with servo-hydraulic control with a capacity up to ± 100 kN and a stroke up to 150 mm at room temperature. - Carrying out a two-axis static and dynamic test in the x-y plane with two actuators with additional loading in the z-direction (e.g. punch test) on a servo-hydraulic machine with a capacity of up to ± 100 kN and a stroke of up to 150 mm. - The system must be able to perform tests with computer control and data acquisition and process the measured quantities to determine tensile mechanical properties in accordance with ASTM E8M, fracture toughness for brittle and ductile materials in accordance with ASTM E 399, ASTM E 1820 and fatigue crack growth rate according to ASTM E 647. At the same time, the system must be able to perform non-contact elongation measurement. Equipment for servo-hydraulic uniaxial ± 250 kN and biaxial ± 100 kN dynamic test stations at low and elevated temperatures must consist of several modular autonomously operating systems, which must be compatible with each other in terms of control, testing and software as well as replacement of individual components. Individual actuators can operate at the same time from the same software. The servohydraulic testing system consists: - an one servo-hydraulic test machine with a two-column frame, one with a capacity at least ±250 kN and a stroke at least of 150 mm (± 75 mm) with a clearance spacing of at least 650 mm or more and a height between the top plate and the yoke of at least 1970 mm height between jaws at least 930 mm when the actuator is in the zero position, - a second servo-hydraulic testing machine with a two-column frame with a capacity at least of ±100 kN and a stroke of at least 150 mm with a spacing between the columns of at least 550 mm or more and a height between the top plate and the yoke of at least 1650 mm, - two flexible servo-hydraulic actuators with a capacity of at least ±100 kN and a stroke of at least 150 mm (± 75 mm) and force probes of ±100 kN and mechanical jaws with one controler and PC system for data aquistion. - ambient temperature chamber for temperature range from -120°C to +315°C with internal space at least 760x400x400 mm (heightxdeepxwidth) - air-cooled hydraulic power unit with a power of at least 30 kW and an oil flow of at least 78 l / min and a pressure of at least 205 bar. - air-cooled hydraulic power unit with a power of at least 18 kW and an oil flow of at least 48 l / min and a pressure of at least 205 bar. - Video extensometer for automatic non-contact measurement of elongation within the visible range 240mm Field of View, which can be upgraded for up to 500 mm Field of View (FOV) with integrated DIC (Digital Image Correlation) technique at a later date. - Additional load cell for measuring force up to 5kN with the possibility of connection to the software platform for control and processing of measurements. - a dynamic extensometer with an initial measuring distance of 25 mm and a trafel of -2.5 mm to + 12.5 mm for measurements in the temperature range from -80 ° C to + 177° C. - Hardware equipement and software platform for controlling tests and capturing force and displacement measurements, and for automatic data processing and reporting. The software must include a specific software for static applications, a specific one for dynamic applications and dynamic packages for Fatigue Testing and for Fracture Toughness, covering: KIC; JIC; CTOD; DCPD; FCG and Fracture and Fatigue Analyzer according to the following standards: -ASTM E606 and D3479, for LCF - ASTM E466 and D3479, for HCF - Including templates for LCF and transition to HCF - ASTM E399 (support FFC(T) and SE(B) specimen) - ISO 12135 (support FFC(T), M(T) and SE(B) specimen) - ASTM E1820 (support FFC(T), M(T) and SE(B) specimen) - An alignment fixture and software for precise adjustment of centricity and angular deviations with measuring sheets and sample for pressure test up to a force of at least 250 kN. - Factory-integrated computers equipment with MS Windows 10 Professional or newer operating system, keyboard and LED screen at least 22 ”. Minimum technical requirements are given below. The offer must match or exceed all minimum technical requirements. |  |
| --- | --- | --- | --- | --- |
| Številka postavke Item No. | Število kosov No. of pieces | ZAHTEVANO |  | REQUIRED |
| 2.1 2.2 2.3 2.4 2.5 2.6 2.7 2.8 2.9 2.10 2.11 2.12 2.13 2.14 2.15 2.16 2.17 2.18 2.19 2.20 2.21 2.22 2.23 2.24 2.25 2.26 2.27 2.28 2.29 2.30 2.31 2.32 2.33 2.34 2.35 2.36 2.37 2.38 2.39 2.40 2.41 2.42 2.43 2.44 2.45 2.46 2.47 2.48 2.49 2.50 2.51 2.52 2.53 2.54 | 1 1 1 1 1 1 1 1 set for flat 1 for round 1 set 1 set for flat 1 for round 1 set 1 set 1 1 1 1 1 1 1 1 1 set 1 set 1 1 1 1 1 1 1 1 1 1 set 1 1 1 1 1 set for flat 1 for round 1 1 1 1 1 1 1 1 set 1 set 2 1 1 2 2 set 1 1 set 1 set | Minimalne tehnične zahteve: I. Servohidravlični stroj za dinamična testiranja do ±250kN I.a) Ogrodje stroja - Nosilni okvir stroja mora biti izdelan kot dvoosna konstrukcija za dinamična testiranja, s katerim je možno opraviti natezna in tlačna testiranja do sile vsaj ±250kN in pri pomiku bata cilindra do 150 mm (ali ±75 mm) - Nosilni okvir stroja mora imeti togost vsaj 470 kN/mm, ki se izmeri na višini jarma 0,9 m od spodnje plošče) - Maksimalna višina ogrodja ne sme preseči 3700 mm. - Spodnja osnova stroja (velikost temelja) mora biti manjša od 1200x1200 mm. - Nosilni okvir mora biti samostoječa pokončna konstrukcija, ki je fiksno pritrjena v temelj z vijaki, ali ima lastno protivibracijsko izolacijsko podlago. - Višinska razlika med zogrnjim jarmom in osnovno pritrdilno ploščo distance mora biti vsaj 1970 mm - Razmik med stebroma mora biti vsaj 635 mm - Nosilni okvir stroja mora omogočiti hidravlično dvigovanje in spuščanje jarma ter zagotoviti hidravlično pritrditev za preprečitev pomika jarma. - Nosilni okvir stroja mora zaznati, da zgornji jarem ni pritrjen in preprečiti začetek izvajanja testa - Nosilni okvir stroja mora biti opremljen z vgrajenim stikalom za zasilno ustavitev v skladu z veljavnimi predpisi. Stikalo za ustavitev v sili ne sme biti na kablu, ki omogoča odmikanja od preizkusnega prostora. - Nosilni okvir stroja mora imeti na ogrodju stroja nameščeno prenosno ročico za fino in grobo nastavitev položaja aktuatorja. - Nosilni okvir stroja mora imeti možnost omejitve hitrosti bata aktuatorja na najmanj 10 mm/s - Sistem za obremenjevanje s silo ±250 kN mora izpolnjevati zahteve standarda EN ISO 13849-1 na ravni zmogljivosti D (PLd). Sistemi z nižjo ravnijo zmogljivosti ne bodo upoštevani - Dobavljena in montirana mora biti naprava, ki preprečuje rotacijo aktuatorja tudi, ko je montirana temperaturna komora. - K servohidravličnem stroju mora biti priložena naprava z obremenitveno kapaciteto do 400 kN za poravnavo osnosti skupaj s programsko opremo za merjenje in prilaganjanje poravnave sistema. I.b) Aktuator, hidravlični razdelilec in servo-ventili - Na nosilni okvir stroja je na spodnjo nosilno ploščo vgrajen servohidravlični aktuator s linearnim pogonom - Osnovna spodnja plošča in telo aktuatorja morata biti integrirana kot en element. - Linearno-proporcionalni aktuator mora imeti statično in dinamično vrednost pri sili najmanj 250kN pri 210 barih. Preverjen mora biti na utrujenje, natezno in tlačno testiranje v primeru, ko je bat aktuatorja v nultem položaju, iz katerega se premika 75mm - Linearno proporcionalni aktuator mora imeti vsaj 150 mm celotnega hoda bata ali 75mm - Linearno-proporcionalni aktuator mora biti opremljen z integriranim merilcem pomika bata aktuatorja, ki ima natančnost v klasi 0.5 po ISO 9513. - Linearno-proporcionalni aktuator mora biti opremljen z razdelilnikom in servoventili, ki omogočajo pretok olja vsaj 40 l/min in za pritisk 210 barov - Hidravlični razdelilnik mora biti opremljen z vgrajenim hidravličnim izklopom, tako da je stroj možno izolirati in ustvariti hidravlični krog brez potrebe po dodatnih hidravlično krmilnih modulih. I.c) Sonda za silo ±250kN - Servohidravlični stroj mora biti integralno opremljen s sondo za linearno merjenje sile v statičnem in dinamičnem načinu obremenjevanja do ±250kN - Sonda za silo se uporablja kot integralni del preizkuševalnega stroja, ki meri silo in krmili eksperiment v kontroli sile - Sonda za merjenje sile mora imeti merilnik pospeška, ki je nameščen neposredno v osi obremenitve, za merjenje vztrajnostnih obremenitev, ki nastanejo zaradi osnega pomikanja mas med cikličnim testiranjem - Merilni sistem mora biti sposoben v realnem času zmanjšati vztrajnostne obremenitve iz signala sile in mora popolnoma avtomatično nastaviti dinamično kompezacijo sile ali nosilni okvir in priključki morajo biti posebej dizajnirani da minimizirajo vztrajnostni učinek na premikajoče se mase ter s tem zagotovijo dinamično togost. - Sistem za merjenje sile mora ustrezati standaru ISO 7500-1 del 1 razred 0.5 do 1/50 obsega merjene sile - Sonda za silo mora prenesti vsaj 150 % nominalno preobremenitev brez mehanskih poškodb I.d) Sonda za silo ±5 kN - Servohidravlični stroj mora biti opremljen s sondo za linearno merjenje sile v statičnem in dinamičnem načinu obremenjevanja do ±5kN - Sonda za silo 5 kN se uporablja kot integralni del preizkuševalnega stroja, ki meri silo in krmili eksperiment v kontroli sile. Maksimalna preobremenitev je vsaj 30 % nad nominalo oz. vsaj 6,5 kN! - Sonda mora biti dobavljena z dodatnimi elementi za pritrditev v čeljusti I.e) Čeljusti za prejem - Sistem za statična in dinamična testiranja v nategu in tlaku do maksimalne sile ±250kN mora biti dobavljen s hidravličnimi koničastimi čeljustmi, ki so primerne za pritrjevanje okroglih (s premerom od 6,5 do vsaj 30 mm) in ploščatih vzorev (debeline od 1,1 do vsaj 35 mm) - Nadzorni sistem mora preprečiti odpiranje ali zapiranje čeljusti med testiranjem bodisi v kontroli sile ali pomika, ali v primeru ročne kontrole čeljusti, ki je locirana na nosilnem okvirju, samo namenoma je lahko uporabnik lahko aktivira ali dezaktivira - Mehanske ali hidravlične čeljusti za dinamična testiranja do ±50 kN ali 100kN morajo biti dobavljene. Čeljusti morajo biti primerne za uporabo v temperaturni komori na temperaturi od -40° to +175°C - Sistem za statična in dinamična testiranja v nategu in tlaku do maksimalne sile ±50kN ali 100kN mora biti dobavljen tudi z mehanskimi koničastimi čeljustmi, ki so primerne za pritrjevanje okroglih s premerom od 6,5 do 26 mm in ploščatih vzorev (debeline od 1,1 do 25,5 mm) za statična natezna in tlačna testiranja ko - Dve vodno hlajeni prijemali morata biti primerni za testiranje na temperaturi do +177°C - Dobavljene morajo biti čeljusti za izvajanje testiranja lomne žilavosti na kompaktnih preizkušancih do širine 25±0.5 mm I.f) Krmilno-kontrolni sistem - upravljalna konzola - Sistem mora imeti digitalni nadzorno krmilni sistem, ki lahko izvaja testiranje v kontroli položaja aktuatorja, v kontroli sile ali deformacije s pomočjo merilca. - Krmilna zanka mora omogočiti generiranje oblike obremenjevanja z najmanjšo hitrostjo v zaprti zanki od 6 kHz - Krmilnik mora imeti možnost avtomatskega posodabljanja pogojev krmilne zanke s kompezacijo za spremembo togosti vzorca med obremenitvenim ciklom. Ta značilnost mora delovati pri vsaj 1 kHz - Krmilnik mora imeti vsaj 6 kontrolnih zank (kot so sila, pomik, deformacija, zaostajanje, naprej, zarezni filter) - Krmilnik mora imeti kanale za pozicijo, aksialno silo in dva opcijska kanala, ki se povežejo z merilci (npr. raztezka, deformacij). Kanali morajo biti dostopni za kontrolo in zajemanje merjenih podatkov. - Podatke iz vseh nadzornih kanalov za zajem podatkov je potrebno zapisovati sinhronizirano in omogočiti polno hitrost posodabljanja s frekvenco vsaj 6 kHz po vseh kanalih ne glede na število kanalov v uporabi - Krmilnik mora imeti vsaj en analogni vhod za zunanji analogni krmilni signal v obliki valovne krivulje in vsaj 4 analogne izhode, ki jih uporabnik poljubno določi. - Vsi priključeni pretvorniki (merilci) morajo biti samodejno prepoznani in kalibrirani s strani kontrolorja - Krmilnik mora preprečiti kakršenkoli test z nekalibriranim merilcem, da prepreči zajemanje nepravilnih rezultatov. Torej, brez kalibracije sistem ne more delovati. - Krmilnik mora imeti vsaj dva detektorja mejnega stanja in mora imeti čas zaznamovanja vsaj 1ms ali manj. Detektorji mej ne smejo biti odvisni od povezave s osebnim računalnikom. - V primeru izgube pobezave, krmilnik mora zaključiti cikel potem se ustaviti. - Nadorni sistem mora vključevati funkcije ročnega in samodejnega nastavljanja - Sistem mora delovati s merodajnim profilom, da doseže optimalno krmilno zanko. - Funkcija samodejnega uglaševanja mora hkrati nastaviti vse nadzorne kanale (pozicija, sila, pomik) z uporabo samo enega parametra (npr. togosti vzorca) - Krmilni sistem mora imeti funkcijo, ki zagotavlja, da je obremenitev med preizkusom znotraj vnaprej določenega območja obremenjevanja preizkušanca. I.g) Programska in računalniška oprema - Sistem mora imeti priloženo programsko opremo za dinamično utrujanje, nizkociklično obremenjevanje, kvazistatično obremenjevanje in testiranje za določitev parametrov mehanike loma. - Vmesnik uporabnikovega osebnega računalnika mora uporabniku jasno podati trenutni način izvedbe testa in jasno podati informacijo o stanju preizkusa npr. v teku, zaključen, dosežena meja ali aktuator je izključen - Dobavljen mora biti tovarniško integriran računalnik s procesorem i7, s tipkovnico in LCD zaslonom vsaj 22'' - Vsak računalnik, ki se uporablja v okviru preizkuševalnega sistema, mora imeti operacijski sistem Windows 10 Professional ali novejši - V nekaterih primerih preizkusna programska oprema lahko dovoli prekrivanje preizkusnih zaslonov, če velikost okna dopušča, da se še vidi nadzorni zaslon. - Zaželeno je da zaslonski nadzorni zaslon bo vedno dobro viden, zato se tega območja zaslona ne sme zmanjšati ali zakriti. I.h) Programska oprema za utrujanje - Sistemi morajo biti opremljeni z neodvisno preizkusno programsko opremo za različne vrste preizkusov (npr. kvazistatične teste, utrujenja in lomno-mehanske teste) - Program za izvajanje testa mora komunicirati s temperaturno komoro in imeti možnost prilagajanja in nadzora temperature med samim testom. - Vsa programska oprema mora imeti funkcijo pomoči G1) Programska oprema za dinamična testiranja - Program mora biti zmožen generirati predpisano dinamično karakteristiko obremenjevanja: - Sinus, utripni sinus - Oglato, utripno oglato - Trikotno, utripno trikotno - Linearno - mirovanje - obremenitvena krivulja, določena s strani uporabnika na osnovi nastavitve ali posameznih izmerjenih točk - Trapezna obremenitvena krivulja s hitrostjo naraščanja obremenitve, ki jo določi uporabnik - Program mora zagotavljati prilagodljive parametre beleženja podatkov, ki jih je mogoče spremeniti za vsak preizkusni korak testa: - Inkrementalno beleženje podatkov - shrani podatke, če se spremenijo za količino, ki je večja od določene - Neprekinjeno beleženje podatkov, linearni ali logaritemski cikel poveča beleženje podatkov - Inkrementalno zapisovanje amplitudne zanke-beleži podatke o porušitvi vzorca, tako da zazna spremembo amplitude v zaprti zanki - Dinamična programska oprema mora imeti funkcijo hitrega preizkusa, da v kratkem času oblikuje karakteristiko utrujenostnega testa. - Dinamična programska oprema mora imeti možnost izvajanja krmiljenja za poljubno obliko obremenjevanja v zaprtem krogu - Dinamična programska oprema mora biti opremljena s knjižnico primerov metod za pomoč pri usposabljanju in učenju uporabnikov - Za natančno sledljivost podatkov je zaželeno, da za vsako izvedbo preizkusnega niza v mapi z rezultati vedno shranjena samo kopija preizkusne metode - Med preizkusom mora imeti programska oprema uporabniku prilagodljive grafe testnih podatkov, ki omogočajo samodejno spreminjanje v realnem času in ročno povečavo / pomanjšavo G2) Programska oprema za statična testiranja - Statična programska oprema mora imeti funkcionalnost hitrega preizkusa, ki omogoča izvajanje preizkusa v nekaj sekundah z enim samim parametrom - Statična programska oprema mora vsebovati vnaprej izdelano knjižnico metod, ki ustreza mednarodnim standardom. Vključno s ASTM E606, D3479, ASTM E466, ASTM E517 in ISO 12135 - Statična programska oprema mora imeti možnost ogleda in snemanja video posnetkov v živo med preizkusom in ga shraniti med rezultate testa - Statična programska oprema mora imeti različne uporabniške ravni in možnost nastavitve vodenih ali pozvanih preizkusnih delovnih tokov za različne operaterje - Statična programska oprema mora imeti lastno funkcionalnost za analizo in poročanje G3) Programska oprema za lomno mehaniko - Programi za lomno mehaniko morajo imeti sposobnost izvesti predutrujanje, da/dN K1c in J1c tests - Programi za lomno mehaniko morajo obdelati rezultate po najnovešjih standardih: - K1c ASTM E399, ISO 12135, - J1c ASTM E1820 - zaželjeno B645, BS7448-1 prav tako - Programi za lomno mehaniko morajo imeti možnost posodobite v skladu s standardi, ne da bi bilo potrebno posodobiti glavni programski paket - Programi za lomno mehaniko morajo zajemati standardne preizkušance kot so: kompaktni natezni (CT), in enostransko zarezani upogibni (SENB), ter po možnosti diskasti kompaktni CT, ločno oblikovani natezni - Programi za lomno mehaniko morajo izvajati lastno analizo in poročanje - Programi za lomno mehaniko morajo izvajati preverjanja veljavnosti parametrov vzorca in rezultatov preizkusov glede na preizkusni standard - Programi za lomno mehaniko morajo imeti različne uporabniške ravni in možnost nastavitve vodenih testnih delovnih tokov - Merilni sistem za lomno mehaniko mora biti kompatibilen s sistemom za meritve rasti razpoke na osnovi padca potenciala z obstoječo Matelect DCM-2 napravo za J1C in da/dN teste - Video kamera za snemanje in ponovno prikazovanje I.i) Hidravlična enota in cevi - Hidravlična enota mora zagotoviti pretok olja vsaj 78 lpm pri 210 bar ali tudi lahko se dobavi tudi enotna hidravlična enota za celotno servohidravlično preizkuševaliče (za I. in II system) s kapaciteto vsaj 100 lpm pri 210 barih, vendar njena moč ne sme presegati vključno 45 kW. Napajanje z 400 V, 50Hz in 3 faze. - Enota mora biti zračno hlajena ali z zračnim predohlajanjem hladilne tekočine ali hladilnim stolpom z zaprto zanko. - Enota mora vsebovati digitalni zaslon s pritiskom in temperaturo olja ter vključevati enote za zaščito pregrevanja olja, pritisk olja, nivo olja ter stanje filtrov in temperature motorja. - Hrup mora biti manj kot 73 dB (A) - Hidravlična enota mora biti popolnoma zaprta v integriran protihrupni kontejner - Enota mora biti opremljena z največ 3-mikronskim (alil manjšim) tlačnim in povratnim filtrom - Priložene morajo biti vse hidravlične cevi za povezovanje naprave in hidravlične enote in sicer 6 m med hidravlično enoto in hidravličnim razdelilcem, 3 m med hidravličnim razdelilcem in preizkuševalnim strojem - Dobavljeno mora biti olje in vsi filtri za normalno delo hidravlične enote - Samostoječi hidravlični razdelilec z enim dovodom in tremi izhodi - Hidravlična enota ali aktuator mora biti opremljen z akumulatorom pritiska I.j) Ekstenziometri - Sistem mora biti opremljen s pritrdilnim eksenziometrom, pritrdilnimi merilniki COD (ang. Crack Opening Displacement) in brezkontaktnim ekstenziometrom - Pritrdilni ekstenziometer mora imeti 25 mm merilne razdalje s potjo od -2,5 mm do +12,5 mm - Pritrdilni ekstnziometer mora imeti nastavke za merilno razdalji za 25mm in 37,5 mm ali 40 mm - Pritrdilni ekstnziometer mora se lahko uporabi za meritve v temperaturnem območju od vsaj -80°C (ali manj) do vsaj +175°C (ali več). - Sistem mora vsebovati tudi merilec odpiranja razpoke (COD) z začetno 5mm merilno razdaljo in 2 mm hoda - Sistem mora vsebovati tudi merilec odpiranja razpoke (COD) z začeto 10 mm ali 12 mm merilno razdaljo in 4 mm hoda - COD merilca sta kompatibilna s standardom ASTM E399 in za meritve v temperaturnem območju od vsaj -80°C (ali manj) do vsaj +175°C, ali več. - Dobavljen mora biti brezkontaktni video ekstenziometer, ki deluje na principu Digital Image Correlation za meritve v vidnem polju do 240 mm. Druge izvedbe brezkontaktnih ekstenziometrov ne bodo upoštevane. - Za vidno polje vsaj 240 mm je treba uporabiti leče s goriščno razdaljo 16mm. - Montažni set za video ekstenziometer primeren za postavitev na vrata ambientne komore. - Pritrdilni nosilec dinamičnega video ekstenziometra za servohidravlični stroj s hidravličnimi čeljustmi - Brezkontaktni video ekstenziometer mora biti sposoben popolnoma integrirati v preizkusno programsko opremo in delovati samostojno - Brezkontaktni video ekstenzometer mora imeti možnost samodejnega merjenja dolžine merilnih vzorcev - Brezkontaktni video ekstenziometer mora biti sposoben izvajati meritev deformacije pri utrujenostih preizkusih do frekvence 20 Hz - Brezkontaktni videoekstenziometer mora imeti natančnost resolucije vsaj 2µm, ko se uporablja pri utrujenostnih preizkusih - Brezkontaktni videoekstenziometer mora imeti natančnost resolucije vsaj 0.5µm ko se uporablja pri kvazistatičnih testih - Vsi našteteti ekstenziometri morajo imeti sposobnost merjenja in krmiljenja, tako pri kvazistatičnih testih kakor tudi pri utrujenostnih preizusih. Pri tem noben od testov ne sme potekati z večjo frekvenco kot 5 Hz. I.k) Temperaturno ambijentna komora - Sistem mora biti opremljen s temperaturno ambijentno komoro za testiranje v temperaturnem območju od vsaj -80°C (ali nižjo) do vsaj +315°C (ali višjo) temperaturo. Torej, tako da celotni temperaturni razpon mora biti vsaj 430°C. - Komora se nadzira s temperaturnim regulatorjem in testno programsko opremo, vključno s spreminjanjem temperature, nastavitvijo časa doseganja temperature in časa ohranjanja na želeni temperaturi ter sprožitvijo testa - Ohlajanje se izvaja s tekočim dušikom do vsaj -80°C - Komora mora imeti notranjo luč, odstranljive klinaste odprtine in ogrevano sprednje okno, da omogoča enostavno spreminjanje nosilne linije in vidnost vzorca med preizkusom - Notranje dimenzije komore morajo biti vsaj 760x350x400mm (višina x širina x globina) - Zunanje dimenzije komore ne smejo presegati 970x530x900mm (višina x širina x globina) - Komora mora ležati na valjčnem nosilcu ali hladilna komora mora imeti lastni voziček, ki omogoča enostavno vstavljanje in odstranjevanje iz preizkusnega prostora - Okvir za nosilec in podporne noge za komoro ali komora mora imeti premični lastni voziček - Cevi za tekoči dušik vsaj 2m II. SISTEM ZA OBREMENJEVANJE v z in x-y SMERI z ±100 kN II.a) Nosilni okvir stroja - Nosilni okvir stroja mora biti izdelan kot dvoosna konstrukcija za dinamična testiranja, s katero je možno opraviti natezna in tlačna testiranja do sile vsaj ±100kN - Nosilni okvir stroja mora imeti togost vsaj 330kN/mm, ki se izmeri na višini jarma 1m od spodnje plošče, kar ustreza togosti 460 kN/mm pri 0,75 m višine zgornjega jarma od spodnje plošče. - Višinska razlika med zgornjim jarmom in osnovno pritrdilno ploščo mora biti vsaj 1630 mm - Razmik med stebroma mora biti vsaj 530 mm - Nosilni okvir stroja mora omogočiti hidravlično dvigovanje in spuščanje jarma ter zagotoviti hidravlično pritrditev za preprečitev pomika jarma. - Nosilni okvir stroja mora zaznati, da zgornji jarem ni pritrjen in preprečiti začetek izvajanja testa - Nosilni okvir stroja mora biti opremljen z vgrajenim stikalom za zasilno ustavitev v skladu z veljavnimi predpisi. Stikalo za ustavitev v sili ne sme biti na kablu, ki omogoča odmikanja od preizkusnega prostora. - Nosilni okvir stroja mora imeti na ogrodju stroja nameščeno prenosno ročico za fino in grobo nastavitev položaja aktuatorja. - Nosilni okvir stroja mora imeti možnost omejiti hitrosti bata aktuatorja na najmanj 10 mm/s - Sistem za obremenjevanje s silo ±100 kN mora izpolnjevati zahteve standarda EN ISO 13849-1 na ravni zmogljivosti D (PLd). Sistemi z nižjo ravnijo zmogljivosti ne bodo upoštevani II.b) Aktuator, hidravlični razdelilec in servo-ventili - Nosilni okvir stroja vključuje servohidravlični linearni pogon z aktuatojem, ki je nameščen pod spodnjo nosilno ploščo ogrodja stroja in površina bata aktuatorja mora biti plazemsko utrjena - Linearno-proporcionalni aktuator mora imeti statično in dinamično vrednost pri sili najmanj 100kN pri 210 barih, za izvajanje utrujanja, nateznih in tlačnih testiranj - Linearno-proporcionalni aktuator mora imeti celotni hod vsaj 150 mm (75mm) in hidrostatične ležeje za zaščito pred prečnimi obremenitvami - Linearno-proporcionalni aktuator mora biti opremljen z integriranim merilcem pomika bata aktuatorja, ki ima natančnost razreda 0.5 v skladu s ISO 9513 - Linearno-proporcionalni aktuator mora biti opremljen z razdelilnikom in servoventili, ki omogočajo pretok olja vsaj 55 l/min. - Hidravlični razdelilnik mora biti opremljen z vgrajenim hidravličnim izklopom, tako da je stroj mogožno izolirati in ustvariti hidravlični krog brez potrebe po dodatnih hidravlično krmilnih modulih ali na ekvivalenten način zagotoviti enako funkcionalnost. II.c) Sonda za silo do ±100kN - Servohidravlični stroj in aktuatorja morata biti integralno opremljen s sondo za linearno merjenje sile v statičnem in dinamičnem načinu obremenjevanja do ±100kN - Merilni sistem mora biti sposoben v realnem času zmanjšati vztrajnostne obremenitve iz signala sile in mora popolnoma avtomatično nastaviti dinamično kompezacijo sile ali nosilni okvir in oprema morajo biti namensko dizajnirani da zmanjšajo vpliv inercije pomikanja mas, tako da dosežejo dinamično togost. - Sistem za merjenje sile mora ustrezati standaru ISO 7500-1 del 1 razred 0.5 do 1/50 obsega merjene sile - Sonda za silo mora prenesti vsaj 150 % nominalno preobremenitev brez mehanskih poškodb II.d) Čeljusti - Sistem mora biti dobavljen s hidravličnimi koničastimi čeljustmi, ki so primerne za pritrjevanje okroglih (s premerom od 6.5 do vsaj 26 mm) in ploščatih vzorev (debeline od 1,1 do vsaj 25 mm) za statična in dinamična testiranja v nategu in tlaku do maksimalne sile ±100kN - Nadzorni sistem mora preprečiti odpiranje ali zapiranje čeljusti med testiranjem bodisi v kontroli sile ali pomika ali nadzorni sistet za ročno zapiranje čeljusti, ki je lociran na nosilnem okvirju je lahko le namenoma aktiviran ali dezaktiviran. II.e) Krmilno-kontrolni sistem - upravljalna konzola Upravljana kozola krmilno-kontrolnega sistema mora zaradi kompatibilnosti izpolnjevati popolnoma enake pogoje kot je navedneno po točkah od 37. do 51. v poglavju I.f) teh tehničnih zahtev in sistem lahko nadzira 100 kN servo-hidravlični stroj z dvema aktivatorjima s pomočjo enega računalnika in programa. II.f) Programska in računalniška oprema Programska in računalniška oprema mora biti avtonomna in zaradi kompatibilnosti izpolnjevati pogoje, ki so našteti v poglavju I.g) po točkah od 52. do 57. II.g) Programska oprema za utrujanje Z namenom zagotovitve kompatibilnosti s celotnim merilnim sistemom mora programska oprema za utrujanje izpolnjevati tehnične zahteve, ki so naštete v poglavlju I.h) po točkah od 58. do 60. G1) Program za dinamična testiranja Z namenom zagotovitve kompatibilnosti s celotnim merilnim sistemom mora programska oprema za utrujanje izpolnjevati tehnične zahteve, ki so naštete v poglavlju I.h) po točkah od 61. do 67. G2) Program za statična testiranja Z namenom zagotovitve kompatibilnosti s celotnim merilnim sistemom mora programska oprema za utrujanje izpolnjevati tehnične zahteve, ki so naštete v poglavlju I.h) po točkah od 68. do 72. G3) Programi za lomno mehaniko Z namenom zagotovitve kompatibilnosti s celotnim merilnim sistemom mora programska oprema za utrujanje izpolnjevati tehnične zahteve, ki so naštete v poglavlju I.h) po točkah od 73. do 80. II.h) Hidravlična enota, priključki & cevi - Hidravlična enota mora zagotoviti pretok olja vsaj 48 lpm pri 210 bar ali ena hidravlična enota z maksimalno močjo 45 kW in minimalnim pretokom olja vsaj 100 lpm pri 210 barih lahko napaj oba sistema ±250 kN in ±100 kN ter oba fliksibilna aktuatorja ±100 kN - V pimeru če je posebna hidravlična enota, ta mora biti zračno hlajena in vsebovati digitalni zaslon s pritiskom in temperaturo olja ter vključevati enote za zaščito pregrevanja olja, pritisk olja, nivo olja ter stanje filtrov in temperature motorja - Če je posebna hidravlična enota dobavljena, ta mora imeti hrup manjši kot 73 dB (A) - Če je posebna hidravlična enota dobavljena, ta mora biti opremljena z 3 mikronskim tlačnim in povratnim filtrom - Priložene morajo biti vse cevi v dolžini 6 m med razdelilcem in aktuatorjem stroja - Dobavljeno mora biti olje in vsi filtri za normalno delo hidravlične enote II.i) Dva fleksibilna aktuatorjas vsaj po ±100kN za 2D x-y testiranje - Vsak aktuator s pretokom vsaj 40 l/min s kapaciteto sile do ±100kN in hodom ±75mm ima lastni postavek za pritriditev in skupno krmilno-upravljalno konzolo, ki mora izpolnjevati zahteve navedene v točki II.e). - Skupni kontroler za oba aktuatorja in enoosni samostoječi stroj do ±100kN (sekcija II.a) mora imeti skupni osebni računalnik s predpisanimi zahtevami, kot je opisano v sekciji II.f) zgoraj. - Vsak aktuator mora biti priključen na hidravlični razdelilec, ki ima integriran izklop ter s tem niso potrebni dodatni hidravlični krmilni moduli ali v primeru samostoječega hidravličnega razdelilca mora biti sestavljen s primerno opremo. - Vsak aktuator mora biti opremljen z merilno celico za silo, ki izpolnjuje pogoje naštete v oddelku II.c), za kapaciteto vsaj do ±100kN - Vsak aktuator mora imeti hidravlične ali mehanske čeljusti za utrujanje vsaj do ±100 kN - Vsak aktuator mora imeti potisno ploščo za pritrditev - Samostoječi hidravlični razdelilec za povezavo z glavnim razdelilcem in fleksibilnimi aktuatorji mora biti dobavljen - Vsaki bat aktuatora rabi biti površinsko utrjen s plazmo namesto standardnega kromiranja - Set hidravličnih cevi v dolžini 2 x 6m med glavnim razdelicem in obema aktuatorjima morajo biti dostavljene in med montažo prilagojene vgradnim meram. lII.) Namestitev in usposabljanje Ponudnik mora izvesti inštalacijo in povezovanje opreme in sistemov. Demonstrirati uporabo programov. Dobavitelj mora organizirati vnos in postavitev opreme v laboratorijski prostor. Vse celice za silo je potrebno kalibrirati po ISO 7500 za nateg in tlak Potrditi je potrebno meritve, ki se nanašajo na jarem, pomik bata aktuatorja ali izmerjen pomik Kalibracija COD merilca Kalibracija video ekstenziometra Akreditacija za osnost po ASTM E 1012 Šolanje in usposabljanje za vsaj 4 uporabnike v trajanju 5 delovnih dni. lV.) Reference proizvajalca vsaj 3 (tri) istovrstne referenčne posle za dobavo primerljive opreme (servohidravličnih preizkuševalnih sistemov istih značilnosti ali kapacitet) v zadnjih 7 letih v skupni pogodbeni vrednosti dobave najmanj 700.000,00 EUR brez DDV. |  | Minimal technical requirements: - ±250kN Axial Dynamic Testing System I.a) Load Frame - The load frame shall be of a two-column construction and be fatigue rated, capable of testing in tensile and compressive modes with actuator with capacity of at least ±250kN and stroke of 150 mm (or ±75 mm) - The load frame must have a stiffness of at least 470 kN/mm (measured at 0.9 m crosshead height) - Maximum high of frame must be lower than 3700 mm - Top view of frame (basement size) must be smaller than 1200 x 1200 mm - The load frame must be selfstanding and fixed by antivibration screws in basement or has isolation pads - Vertical distance between the crosshead and platen to be at least 1970 mm - Distance between columns should be at least 635 mm - The load frame shall be equipped with a hydraulically operated system for raising, lowering, and clamping the moving crosshead - The load frame shall detect the upper crosshead being unclamped and prevent a test being started - The load frame must feature a frame-mounted, emergency stop in full compliance with current regulations. The emergency stop should not be on a cable that allows it to be moved away from the test-space - The load frame must feature a frame-mounted jog control handset with both coarse and fine actuator position adjustment - The load frame must feature a set up mode, restricting the actuators velocity to at least 10mm/s during test set-up - The ±250kN system must meet the requirements of EN ISO 13849-1 performance level D (PLd). Systems which meet a lower performance level will not be considered - Anti rotating device for actuator shoul be supplied - An alignment fixture with load capacity of 250 kN or higer must be included with the machine along with software to measure and adjust the alignment of the system I.b) Actuator, Manifold and Servo-valve - The load frame shall include a servohydraulic linear actuator, mounted in the test frame base - Load frame lower baseplate and actuator body has to be integrate in a single part (single block) - The linear actuator shall have a static & dynamic rating of at least 250kN at 210bar and be fatigue rated for testing in tensile, compressive, and through-zero modes where 75mm stroke displacement starts - The linear actuator shall have a minimum of 150 mm or total stroke (75mm) - The linear actuator must be fitted with an integrated LVDT for actuator position measurement with a measurement accuracy in class 0.5 according to ISO 9513 - The linear actuator should be fitted with a manifold and supplied with servo-valves rated for oil flows of at least 40 lpm and pressure 210 bars - The manifold shall be provided with an integrated hydraulic shut-off, so the machine can be isolated from the hydraulic circuit without the need for additional hydraulic control modules I.c) Load Cell for ±250kN - The system must be supplied with a linear fatigue rated load cell with a static and fatigue rating of ±250kN as integral part of system - Load cell should be possible to use as an integral part of the force measurement system and also in load mode control test of a testing machine. - The load cell must feature an accelerometer mounted directly on the linear load axis to measure inertial loads generated due to the moving fixture mass during a cyclic test - The control system must be capable of reducing inertia-induced loads from the load signal in real time and must feature a fully automatic dynamic load-compensation set-up routine or load frames and accessories are specifically designed to minimize inertial effect due to the moving mass, to achieve tough dynamic applications - The load measurement system shall meet ISO 7500-1 part 1 Class 0.5 down to 1/50th of loadcell full scale - The load cell must have a minimum 150 % overload capacity without mechanical damage I.d) Load Cell for ±5 kN 29. The system must be supplied with a linear fatigue rated load cell with a static and fatigue rating of ±5 kN 30. Load cell should be possible to use as an integral part of the force measurement system and also in load mode control test of a testing machine. Maximum load of load cell is 30% of overload (e.g. up to 6,5 kN) 31. Loadcell should be supplied with accessories for mounting in grips I.e) Grips - The system must be supplied with hydraulic wedge action grips suitable for gripping round (from dia 6.5mm to at least 30 mm) and flat specimens (from 1.1 mm to at least 35 mm), for static tensile and compression as well as full reverse-stress testing up to the full load capacity of the system ±250kN - The control system must prevent the operation of the grips (opening or closing) either whilst a test is running or whilst in load or strain control or the grips controls are manual and located on board the frame, so only an intentional user action could activate or disactivate them - The mechanical or hydraulic grips for fatigue rated testing up to ±50kN or 100kN must supplied. Grips must be suitable for use in a temperature chamber from -40°C to +175°C - The system should be supplied with mechanical wedges suitable for gripping round (from dia 6.5mm to at least 26 mm) and flat specimens (from 1.1 mm to at least 25.5 mm), for static tensile and compression as well as full reverse-stress testing up to ±50kN or 100kN - Two of water-cooled pullrods must be supplied for testing at high temperature up to +177°C? - Grips suitable for fracture mechanics testing of CT specimens up to a width of 25±0.5 mm must be supplied I.f) Control System – control unit - The system must feature a digital control system capable of controlling the actuator in position, load, and strain modes - The control loop must feature waveform generation with a minimum loop closure rate of 6 kHz - The controller must feature a facility for automatically updating the control loop terms to compensate for changes in specimen stiffness during a cycle. This facility must run at a minimum of 1kHz - The controller should have at least a 6 term control loop. (i.e. P, I, D, Lag, Feed Forward, Notch filter) - The controller must have channels for: position, axial load, and two optional channels (e.g. elongation or strain). These must be available for control and data acquisition - Data from all control and aquisition channels should be recorded synchronously and should be capable of at least the full 6 kHz loop update rate across all channels, irrespetive of how many channels are used. - The controller must have at least one analogue input for an external analogue waveform drive signal, and at least four assignable analogue outputs - All supplied transducers shall be automatically recognised and calibrated by the controller, if callibration is necessary - The controller must prevent any test proceeding with an uncalibrated transducer to prevent incorrect data being gathered. Therefore, without the calibration file the system cant work. - The controller must have a minimum of two limit detectors per connected transducer and must have a detection time of 1ms or better. The limit detectors must not be reliant in any way on the connection of a personal computer - In the event of any communications loss, the control cycles must be competed and than system must stop. - The control system must include both manual and automatic tuning functions - The system have to run with arbitary profiles to achieve the best tuning. - The automatic tuning function must tune all controllable channels (Position, Load, Strain) using only one parameter (e.g. specimen stiffness) - The control system must have a feature to ensure the load is kept within a predefined range during test set up and specimen loading. I.g) Software and PC Hardware - The system should have software supplied with the system to cover Dynamic Fatigue, Low Cycle Fatigue, Quasi-static, and Fracture mechanics testing - The PC user interface should clearly indicate to the user the current control mode, and clearly indicate from a distance the test status, e.g. running, finished, limit tripped or actuator power off - Factory integration of supplied Personal Computor with intel core i7 CPU, Keyboard and wide Screen Flat Panel Monitor at least 22'' size - Any PC offered with the system shall utilise the Windows 10 or newer Professional operating system and must feature a 22-inch LCD flat panel monitor or larger - In same case the software can to overlape the screeem if the windows size can be adjuseted to minimize the effect. - Preferibly the on screen control display must always be clearly visible and hence it should not be possible to minimise or obscure this area of the screen I.h) Fatigue Test Software - Systems should come with independent test software for different test types (e.g. quasi-static, fatigue, fracture mechanics etc.) - Test software should communicate with temperature chamber and be able to adjust and control the temperature via the test method where required - All software should have a context-specific help function G1) Dynamic test software - Shall be able to generate the following waveforms and provide looping capability: - Sine, haversine - Square, haversquare - Triangle, havertriangle - Ramp - Hold - User defined waveforms as either turning points or sampled data - Trapezoid waveforms with user defined ramp rates and hold times - Shall provide flexible data logging parameters which can be changed for each test step of a test - Incremental data logging - stores data if it changes by an amount greater than specified - Continuous data logging, linear or logarithmic cycle increments data logging - Incremental amplitude loop logging - logs data at onset of specimen failure as detected by an amplitude change - Dynamic software should have a quick test functionality, to automatically create a fatigue test method in minutes - Dynamic software should have the ability to perform closed-loop mixed-mode control of waveforms - Dynamic software should come with a library of example methods to aid user training and learning - For accurate data traceability, a read only copy of the test method should be preferibly stored in the results folder for every test set run - During a test the software shall have user customisable graphs of test data capable of real time auto-scaling and manual zoom in/out G2) Static test software - Static Software should have quick test functionality to allow running a testing in seconds with only one parameter. - Static software should contain a methods library of pre-built complying to international standards. Including ASTM E606, D3479, ASTM E466, ASTM E517 and ISO 12135 - Static software should come with the ability to view and record live video during a test and save it to the test results - Static software should have different user levels and the ability to set up guided or prompted test workflows for different operators - Static software should have its own analysis & reporting functionality G3) Fracture mechanics test software - Fracture software should have the capability to run pre-cracking, da/dN K1c and J1c tests - Fracture software must have methods to comply to the latest standards - K1c ASTM E399, ISO 12135, - J1c ASTM E1820 - Preferibly B645, BS7448-1 as well - Fracture software should have the ability to update the just the standards compliance without having to update the main software package - Fracture software must support the following specimen types: Compact Tension (CT) and Single Edge Notch Bend (SENB), and preferiby Disc Shape CT, Arc-Shape Tension - Fracture software should have its own analysis & reporting functionality - Fracture software should run validity checks of the specimen parameters and test results against the test standard - Fracture software should have different user levels and the ability to set up guided test workflows - Fracture softwere should be compatible with the Matelect 2DP potential-based availability growth measurement system by present Matelect DCM-2 device for J1C and da/dN tests - Test cam for video recording and playback module I.i) Hydraulic Supply & Hoses - A hydraulic power unit must be provided with a minimum flow rate of 78lpm at 210bar. It is acceptable also only one HPU for all servohydraulic system (I. and II. system), with minimum flow of oil 100 lpm at 210 bar, but maximum allowed power should not exeed 45kW. Supplied by 400 V, 50Hz and 3 phase operation. - The unit must be air cooled or with chiller or with cooling tower closed-loop. - The unit must be fitted with an interface that include a digital display of oil pressure and temperature and include protection devices for oil temperature, oil pressure, oil level, oil filter condition and motor temperature - Noise levels must not exceed 73 dB (A) - HPU should be supplied with fully integrated acoustic attenuation system - The unit or the load frame must be fitted with at least or lower 3 micron pressure filter at least (as close as better) - All hoses for interfacing between the machine and HPU must be supplied, as follows 6 m hoses from the HPU to the Distribution mainfold, 3 m hoses between distribution mainfold to testing system. - Oil and filters for hydraulic power unit must be sustained - Floor standing hydraulc distribution mainfold with a single inlet ad up to three outlets or alternateve solution, shared with the 100kN sysytem - Hydraulic power unit or the actuactor should be equiped with pressure accumulator I.j) Extensometry - System must come with a clip-on extensometer, clip-on COD gauges and a non-contacting extensometer - The clip-on extensometer should have a 25 mm gauge length with -2.5 mm to +12,5 travel - The clip-on extensometer should come with gauge length extenders for 25mm and 37.5 or 40 mm - The clip-on extensometer should be suitable for temperatures from at least -80°C (or lower) at least up to +175° (or higher) - There should be COD gauges provided with 5mm gauge length with 2mm travel - There should be COD gauges provided with 10 mm or 12mm gauge length and 4mm travel - The COD gauges must comply with ASTM E399 and be suitable for temperatures fromat least -80°C to at least +175°C or higher - The non-contacting extensometer should be supplied. A non-contacting video extensometer based on Digital Image Correlation (DIC) measurement system for Field of Viwe (FOV) 240 mm. Other forms of non-contacting extensometer will not be considered - Standard Field of View Lens with 16mm focal length should be used for FOV at least of 240 mm. - Mounting bracket to allow video extensometer to fitted to the front of a chamber door must be supplied - Dynamic video extensometer mounting bracket for servohydraulic machine with hydraulic grips have to be supplied - The non-contacting video extensometer must be able to fully integrate into the test software as well as run in a standalone mode - The non-contacting video extensometer should be able to automatically measure the gauge length of specimens - The non-contacting video extensometer should be able to perform strain control for fatigue tests, up to 20Hz - The non-contacting video extensometer must have an accuracy and resolution of at least 2µm when used for fatigue testing - The non-contacting video extensometer must have an accuracy and resolution of at least 0.5µm when used for quasi-static testing - All extensometers above should be capable of strain measurement and control for both quasi-static and fatigue testing. Please indicate if any are only suitable for quasi-static testing or testing below 5Hz I.k) Temperature Chamber - The system must be supplied with an environmental chamber capable of reaching temperatures from at least -80°C (or lower) to at least +315°C (or higher) temperature. Therefore, total temperature range should be at least 430°C. - The chamber must be controllable by temperature controller from the test software including changing temperature, setting dwell and soak times, and triggering events - LN2 cooling liquid for minimum temperature -80°C - The chamber must feature an internal light, removable wedge ports and a heated front window to allow easy changing of the loadstring and visibility of the specimen mid-test - The chamber internal dimensions should be at least 760x350x400mm (HxWxD) - The chamber external dimensions must not exceed 970x530x1200mm (HxWxD) - The chamber must sit on a roller carriage or chamber can have own a cart that can to allow easy insertion and removal from the test space - Frame Mounting Bracket and support legs or chamber should has own chart - Preferibly LN2 coolant Hose set with length minimum 2m II. SYSTEM WITH ±100 kN for z and x-y axial loading II.a) Load Frame - The load frame shall be of a two-column construction and be fatigue rated, capable of testing in tensile and compressive modes with a capacity of at least ±100kN and 150mm of usable stroke - The load frame must have a stiffness of at least 330 kN/mm (measured at 1m crosshead height) this crresponding to 460 kN/mm at 0.75 m crosshead height! - Vertical distance between the crosshead and platen to be at least 1630 mm - Distance between columns should be at least 530 mm - The load frame shall be equipped with a hydraulically operated system for raising, lowering, and clamping the moving crosshead - The load frame shall detect the upper crosshead being unclamped and prevent a test being started - The load frame must feature a frame-mounted, emergency stop in full compliance with current regulations. The emergency stop should not be on a cable that allows it to be moved away from the test-space - The load frame must feature a frame-mounted jog control handset with both coarse and fine actuator position adjustment - The load frame must feature a set up mode, restricting the actuators velocity to at least 10mm/s during test set-up - The ±100kN system must meet the requirements of EN ISO 13849-1 performance level D (PLd). Systems which meet a lower performance level will not be considered II.b) Actuator, Manifold and Servo-valve - The load frame shall include a servohydraulic linear actuator, mounted in the test frame base plate, and the actuator's rod finishing as to be done through plasma coating - The linear actuator shall have a static & dynamic rating of at least 100kN at 210bar and be fatigue rated for testing in tensile, compressive, and through-zero modes - The linear actuator shall have a minimum of 150 mm total stroke (75mm) and feature hydrostatic bearings to aid with protection from side loads - The linear actuator must be fitted with an integrated LVDT for actuator position measurement with a measurement accuracy in class of 0.5 according to ISO 9513 - The linear actuator should be fitted with a manifold and supplied with servo-valves rated for oil flows of at least 55lpm - The manifold shall be provided with an integrated hydraulic shut-off, so the machine can be isolated from the hydraulic circuit without the need for additional hydraulic control modules or similar/equivalent solution with the same functionality II.c) Load Cell with capacity ±100kN - The system and both actuators must be supplied with a linear fatigue rated load cell with a static and fatigue rating of ±100kN as integral part - The control system must be capable of reducing inertia-induced loads from the load signal in real time and must feature a fully automatic dynamic load-compensation set-up routine or load frames and accessories are specifically designed to minimize inertial effect due to the moving mass, to achieve tough dynamic applications - The load measurement system shall meet ISO 7500-1 part 1 Class 0.5 down to 1/50th of loadcell full scale - The load cell must have a minimum 150% overload capacity without mechanical damage II.d) Grips - The system must be supplied with hydraulic wedge action grips suitable for gripping round (dia 6 to at least 26 mm) and flat specimens (since 1,1 to at least 25 mm), for static tensile and compression as well as full reverse-stress testing up to the full load capacity ±100kN of the system - The control system must prevent the operation of the grips (opening or closing) either whilst a test is running or whilst in load or strain control or the grips controls are manual and located on board the frame, so only an intentional user action could activate or disactivate them II.e) Control System-control unit Control system designed in control unit must meet exactelly same conditions listed in points 37. to 51. of section I.f) this techincal requirements and system can control 100 kN servo-hydraulic testing machine and two actuators by one computer and software. II.f) Software and PC Hardware Systems should come with independent test software and PC Harware, but in order to be compatible with whole testing system, it should complete techical requirements given in section I.g) by points 52. to 57. II.g) Fatigue Test Software In order to be compatible with whole testing system, it should complete techical requirements given in section I.h) by points 58. to 60. G1) Dynamic test software In order to be compatible with whole testing system, it should complete techical requirements given in section I.h) by points 61. to 67. G2) Static test software In order to be compatible with whole testing system, it should complete techical requirements given in section I.h) by points 68. to 72. G3) Fracture mechanics test software In order to be compatible with whole testing system, it should complete techical requirements given in section I.h) by points 73. to 80. II.h) Hydraulic power unit, supply & Hoses - A hydraulic power unit must be provided with a minimum flow rate of 48lpm at 210bar or one HPU with maximum power 45 kW and minimum flaw rate of 100 lpm at 210 bar can supply both systems ±250 kN and ±100 kN with two flexible selfstanding actuators ±100 kN - If separate HPU is provided, the unit must be air cooled and fitted with an interface that include a digital display of oil pressure and temperature and include protection devices for oil temperature, oil pressure, oil level, oil filter condition and motor temperature - If separate HPU is provided the noise levels must not exceed 73 dB (A) - If separate HPU is provided the unit or the load frame must be fitted with a 3 micron pressure filter at least (as close as better) - All hoses for interfacing between the testing actuator and mainfold must be supplied in length of 6 m between - Oil and filters for HPU must be supplied II.i) Flexible 2 actuators with at least of ±100kN for 2D x-y testing system - Each actuator with at least 40 litre/min and equiped with servo valve capacity ±100kN and stroke ±75mm pedestal based actuator must share the same controller (the controller must conform to the requirements outlined in section II.e) above). - The integratl controller for both actuators and selfstanding testing machine ±100kN (secton II.a) must have its own dedicated PC system, with requirements outlined in section II.f) above. - Each actuator must be fitted with a manifold which has an integrated hydraulic shut-off, removing the need for additional hydraulic control modules or with a stand alone floor hydraulic service manifold with the proper fittings - Each actuator must be fitted with a load cell meeting the requirements of section II.c) above but at least for ±100kN capacity - Each actuator should be supplied with at least ±50 or 100 kN fatigue rated hydraulic or mechanical wedge grips - Each actuator has consists ped base actuator rig system - Floor standing distribution mainfold must be supplied or equivalent distribution solution - For each actuactor the rod finishing as to be done through plasma coating insted of standard chrome palating. - Hoses set of from 2 x 6 m length from mainfold to both standalone actuator have to be supplied and fit to distances among mainfold and actuators lII.) Installation and Training The tender should provide installation and integration of equipement and systems. Softwere demonstration basic system. Provider must organize entering equipment to laboratory space. All load cell calibration to ISO 7500 in tension and compresson Verification of displacement measuring system covering crosshead travel, actuator stroke or LVDT displacement COD Gauge calibration Video Extensiometer calibration Accredited Alignment Verification to ASTM E1012 – or supplied equipment must allow accreditated alignment verification according to that standard; Training for least 4 (four) persons within the scope of 5 (five) working days lV.) Company references at least 3 (three) reference transactions of the same type for the supply of comparable equipment (servo hydraulic testing systems) within the past 7 years with a total contract value of delivery at least EUR 700,000.00 excluding VAT. |
| 2.2 | Posebne zahteve za postavko Računalnik za nadzor mikroskopa (osebni oz. namizni računalnik) Special requirements for Microscope Control PC (presonal computer) |  |  |  |

Ponujeno blago mora v celoti izpolnjevati tudi naslednje tehnične specifikacije, s katerimi se zagotavlja izpolnjevanje ciljev v skladu z Uredbo o Zelenem javnem naročanju (Uradni list RS, št. 51/17 in 64/19):

The offered supplies must also fully meet the following technical specifications which ensure the fulfilment of the objectives in accordance with the Decree on green public procurement (Official Journal of the Republic of Slovenia RS, Nos. 51/17 and 64/19):

| Tehnične zahteve: OSEBNI RAČUNALNIKI | Technical specifications: PERSONAL COMPUTERS |
| --- | --- |
| ZAHTEVANO | REQUIRED |
| - Osebni računalniki morajo izpolnjevati najnovejše standarde Energy Star za energijsko učinkovitost, veljavne na dan objave obvestila o javnem naročilu ali na dan povabila k oddaji ponudbe, ali enakovredne standarde. Način dokazovanja: Ponudnik ob ponudbi priloži: - potrdilo Energy Star ali - potrdilo, da ima blago znak za okolje tipa I, ali - tehnično dokumentacijo proizvajalca, iz katere izhaja, da so zahteve izpolnjene. | - Personal computers shall comply with the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. Means of proof: The tenderer submits with the tender: - Energy Star certificate, or - certificate stating that the supply has Type I environmental label, or - manufacturer's technical documentation showing that this requirement is met. |
| - Osebni računalniki morajo biti zasnovani tako, da sta trdi disk in pomnilnik zlahka dostopna z uporabo univerzalnega razpoložljivega orodja (tj. z izvijačem, lopatico, kleščami ali pinceto) in zamenljiva. Način dokazovanja: Ponudnik ob ponudbi priloži: - jasna navodila za razstavljanje in popravilo, ki omogočajo, da se izdelki pri zamenjavi ključnih komponent ali drugih delov za nadgradnjo ali popravilo razstavijo brez škode (npr. sliko z eksplozijskim pogledom, na kateri so prikazani dostopni deli, ki se lahko zamenjajo, in potrebna orodja, videoposnetek), in - potrdilo, da ima blago znak za okolje tipa I, iz katerega izhaja, da blago izpolnjuje zahteve, ali - ustrezno dokazilo, iz katerega izhaja, da so zahteve izpolnjene. | - Personal computers must be designed in such a way that the hard disk and memory are easily accessible using universally available tools (i.e. with a screwdriver, spatula, pliers, or tweezers) and are replaceable. Means of proof: The tenderer submits with the tender: - clear disassembly and repair instructions that allow the products to be disassembled without damage when replacing key components or other parts for upgrading or repair (e.g. exploded view image, showing accessible parts that can be replaced and the necessary tools, video), and - certificate stating that the supply has Type I environmental label, or - suitable proof showing the fulfilments of the demand. |
| - Glasnost osebnega računalnika mora biti izmerjena v skladu s standardom SIST EN ISO 7779 in v času zapisovanja na trdi disk ne sme presegati 26 dB. Podatek mora biti označen z oznako LpAm v dB. Način dokazovanja: Ponudnik ob ponudbi priloži: - potrdilo, da ima blago znak za okolje tipa I, iz katerega izhaja, da blago izpolnjuje zahteve, ali - ustrezno dokazilo, iz katerega izhaja, da so zahteve izpolnjene. | - The volume of the personal computer must be measured in accordance with the SIST EN ISO 7779 standard and must not exceed 26 dB at the time of writing to the hard disk. The data must be labelled with the designation LpAm in dB. Means of proof: The tenderer submits with the tender: - certificate stating that the supply has Type I environmental label, or - suitable proof showing the fulfilments of the demand. |
| TEHNIČNE ZAHTEVE: ZASLON | TECHNICAL SPECIFICATION: SCREEN |
| - Zasloni morajo izpolnjevati najnovejše standarde Energy Star za energijsko učinkovitost, veljavne na dan objave obvestila o javnem naročilu ali na dan povabila k oddaji ponudbe. Način dokazovanja: Ponudnik ob ponudbi priloži: - potrdilo Energy Star ali - potrdilo, da ima blago znak za okolje tipa I, ali - tehnično dokumentacijo proizvajalca, iz katere izhaja, da so zahteve izpolnjene. | - Screens shall meet the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. Means of proof: The tenderer submits with the tender: - Energy Star certificate, or - certificate stating that the supply has Type I environmental label, or - manufacturer's technical documentation showing that the requirements are met. |

Operacijo sofinancirata Republika Slovenija, Ministrstvo za izobraževanje, znanost in šport in Evropska unija iz Evropskega sklada za regionalni razvoj. Operacija se izvaja v okviru Operativnega programa za izvajanje evropske kohezijske politike 2014-2020 v okviru prednostne osi 1: »Mednarodna konkurenčnost raziskav, inovacij in tehnološkega razvoja v skladu s pametno specializacijo za večjo konkurenčnost in ozelenitev gospodarstva«, prednostne naložbe 1.1 »Krepitev infrastrukture za raziskave in inovacije ter zmogljivosti za razvoj odličnosti na tem področju, pa tudi spodbujanje pristojnih centrov, zlasti takšnih, ki so evropskega pomena«, specifičnega cilja 1.1.1 »Učinkovita uporaba raziskovalne infrastrukture ter razvoj znanja/kompetenc za boljše nacionalno in mednarodno sodelovanje v trikotniku znanja«.

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

Standardi ENERGY STAR: http://www.eu-energystar.org/. Različica, veljavna ob izdaji teh okoljskih primerov in meril, je 6.1, posodobitve pa so dostopne na spletni strani http://www.eu-energystar.org/specifications.htm.

ENERGY STAR standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 6.1, and updates are available at http://www.eu-energystar.org/specifications.htm.

Hard Drive Operating.

LpAm (Low power AM) je oznaka za tlak zvočnih emisij.

Hard Drive Operating.

LpAm (Low power AM) is the designation for the sound emission pressure.

Standardi Energy Star: http://www.eu-energystar.org/. Različica, veljavna ob izdaji teh okoljskih primerov in meril, je 7.0, posodobitve pa so dostopne na spletni strani http://www.eu-energystar.org/specifications.htm.

Energy Star Standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 7.0, and updates are available at http://www.eu-energystar.org/specifications.htm.