Project title: Upgrading National Research Infrastructures – RIUM

COMPLIANCE DECLARATION WITH MINIMUM TECHNICAL REQUIREMENTS OF THE OFFERED EQUIPMENT – LOT 5

Name of the tenderer:

Address of the tenderer:

Subject of the public procurement: Purchase of Advanced Equipment for Development of Special Optical Fibers

Lot 5: The preform analyzer

Designation of the public procurement: 302/41 - RIUM56-FERI 03.1/2021

INSTRUCTIONS:

- This Annex No. 2 must be completed and signed. The tenderer must indicate the actual technical specifications of the equipment offered in line with all requirements.

- The tenderer shall upload this Annex No. 2 to the e-JN information system in the section “Other Annexes”.

We declare that the offered equipment fully meets the minimum technical requirements of the client for the public procurement titled “Purchase of Advanced Equipment for Development of Special Optical Fibers” – Lot 5, and that we are ready to realize the order according to the following specification:

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 1 | System hardware: |  |  |  |
|  | - Z-axis longitudinal measurement range: minimal 800 mm. | Z-axis longitudinal measurement range: _______________ mm |  |  |
|  | - Automated rotary head with manual preform clamping mechanism. | - YES | - NO |  |
|  | - Measurement tank transverse measurement range: up to 90 mm. | Measurement tank transverse measurement range: _____________ mm |  |  |
|  | - Built-in silica reference calibration rod. | - YES | - NO |  |
|  | - Suitable length oil tank to accommodate applicable preform measurement range. | - YES | - NO |  |
|  | - Automated preform tip detection system. | - YES | - NO |  |
|  | - Controller with system management over HMI. | - YES | - NO |  |
|  | - Index matched oil to silica at 30 °C. | - YES | - NO |  |
|  | System software supporting the following features: |  |  |  |
|  | - Measurement of multimode & single-mode preforms. | - YES | - NO |  |
|  | - Automated preform rotation & longitudinal translation. | - YES | - NO |  |
|  | - Preform geometry analysis. | - YES | - NO |  |
|  | - Equivalent step index (ESI) single-mode core profile analysis. | - YES | - NO |  |
|  | - Prediction of mode field diameter and cut-off wavelength from single-mode preform profiles. | - YES | - NO |  |
|  | - Alpha fit to multimode core profile geometry. | - YES | - NO |  |
|  | - Measurement analysis report preparation tools. | - YES | - NO |  |
|  | - Configuration for generating automated measurements with scripting support. | - YES | - NO |  |
|  | - Ability to run routines for machine calibration. | - YES | - NO |  |
|  | Performance requirements: |  |  |  |
|  | - Refractive index repeatability: better than 0.00005. (Excludes regions with rapidly changing index of refraction.) | Refractive index repeatability: ____________ |  |  |
|  | - Diameter repeatability: better than 0.05 mm. | Diameter repeatability: ________________ mm |  |  |
|  | - Concentricity repeatability: better than 0.05 mm. (Repeatability is defined as 1 standard deviation on a set of 10 measurements, preform unmoved) | Concentricity repeatability: ________________ mm |  |  |
|  | - Preform diameter range: between 5 mm and 90 mm. | Preform diameter range: between __________ mm and __________ mm |  |  |
|  | - Scan time ( ≥ 4000 data points, ≥ 5 um step): less than 60 seconds. | Scan time: ___________ s with ________ data points at _______ um step. |  |  |

| Place and date: | Seal: | Authorized signature: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.