Naziv projekta: Nadgradnja nacionalnih raziskovalnih infrastruktur – RIUM

Project title: Upgrading National Research Infrastructures – RIUM

Tehnične specifikacije za sistem za lasersko litografijo

Technical specifications for a laser lithography system

- PREDMET JAVNEGA NAROČILA 						1. SUBJECT OF THE PUBLIC TENDER

| Predmet javnega naročila je sistem za lasersko litografijo, vključno z dostavo, vnosom, namestitvijo, montažo, konfiguracijo, inštalacijo in zagonom naprave ter usposabljanjem, skladno s tehničnimi specifikacijami. Lokacija dobave in vnos opreme: Univerza v Mariboru, Fakulteta za elektrotehniko, računalništvo in informatiko, Koroška cesta 46, 2000 Maribor, Slovenija, Laboratorij za elektro-optične in senzorske sisteme objekt G, čisti prostor G-105. Izbrani ponudnik bo moral opremo vnesti in namestiti. Naročnik bo ponudnikom omogočil ogled prostora na lokaciji vnosa in namestitve opreme, skladno s točko 11 Navodil ponudnikom. Rok dobave in namestitve: najpozneje 120 (sto dvajset) dni od dneva pričetka veljavnosti pogodbe. Garancijski rok: najmanj 1 (eno) leto od dneva podpisa prevzemnega zapisnika. Celotna 1-letna garancija vključuje letni servis. Za vso opremo velja, da mora zanjo v času garancijskega obdobja proizvajalec opreme zagotavljati dostop do vseh posodobitev programske opreme brez dodatnih stroškov. Vsi popravki, servisni paketi ter nadgradnje in podpora morajo biti od proizvajalca strojne oziroma programske opreme. Ponudnik mora za opremo zagotoviti: - da so vsi ponujeni moduli/komponente združljivi in medsebojno povezljivi, - da je mogoče vse zahtevane module/komponente povezati v enoten in brezhibno delujoč sistem v skladu s tehničnimi specifikacijami, - odpraviti vse napake, ki bi se pojavile zaradi morebitne medsebojne nezdružljivosti ali nepovezljivosti oz. zaradi neustrezne povezave v enoten, brezhibno delujoč sistem. Usposabljanje: Dobavitelj mora izvesti osnovno začetno usposabljanje, ki vključuje strojne in programske module na dobavljenem predmetu pogodbe za najmanj 1 (eno) osebo uporabnika v obsegu 1 (enega) delovnega dneva. V ponudbi mora biti zajeti ves droben material, ki je potreben za priključitev naprav, tudi če ni izrecno naveden. | Subject of the public tender is a laser lithography system, placement, assembly, configuration, installation, starting of the device and training in line with technical specifications. Equipment delivery and placement: University of Maribor, Faculty of Electrical Engineering and Computer Science, Laboratory for Electro-Optics and Sensor Systems, Koroška cesta 46, 2000 Maribor, Slovenia, Building G, clean room G-105. The equipment shall be delivered and placed by the successful tenderer. The contracting authority will enable the tenderers to inspect the premises on site in line with Point 11 of the Instructions to tenderers. Delivery time: at the latest 120 (one hundred and twenty) days from the day of entry into force of the contract. The warranty period: At least 1 (one) year from the day of signature of the take-over protocol, including annual service. The equipment manufacturer is obliged to provide access to all software updates at no additional cost during warranty period. All patches, service packs, upgrades and support must be provided by the hardware or software manufacturer. The provider of the equipment must guarantee: - that all modules/components offered are compatible and interconnected, - that all required modules/components can be integrated into a single and seamlessly functioning system in line with technical specifications, - to eliminate all errors resulting from possible incompatibility or mismatch, or from improper connection to a single, seamlessly functioning system. Training: The supplier must carry out basic initial training, which includes hardware and software modules on the delivered subject of the contract for at least 1 (one) person within the scope of 1 (one) workday. The offer must include all the fine material, necessary for connecting the devices, even if it is not explicitly stated. |
| --- | --- |

- TEHNIČNE ZAHTEVE							2. TECHNICAL SPECIFICATIONS

| Sistem za lasersko litografijo Univerza v Mariboru naroča vsestranski laserski litografski sistem z visoko ločljivostjo brez maske za širok nabor aplikacij, kot so mikrofluidika, mikromehanika ali fotonika. Sistem mora biti primeren za mikroizdelave, izdelavo fotomask in hitro izdelavo prototipov neposredno na tankih ali debelih substratih. Sistem za lasersko litografijo mora omogočiti visoko stopnjo prilagodljivosti za kreiranje mikrostruktur v/na foto občutljivih plasteh. Sistem mora podpirati rastrski in vektorski način osvetlitve/pisanja. Sistem za lasersko litografijo mora biti opremljen z motorizirano mizo za vzorce z visoko ločljivostjo pri pomikih v X in Y smereh, z nizkimi pogreški pri premikih in hitrimi premiki pri skeniranju. Sistem za poravnavo mora biti opremljen s kamero visoke ločljivosti in z napredno programsko opremo/sistemom za zajem slike ter podpirati aplikacije, ki zahtevajo več plasti. Laserski vir mora podpirati osvetlitev širokopasovnih in posebnih i-line fotorezistov, vključno s SU-8, AZ-nLOF, TOK IP in drugimi. Najmanjša velikost značilnosti/strukture, ki jo lahko zapišemo s sistemom laserske litografije, je podana v minimalnih tehničnih zahtevah. Sistem mora biti primeren za naslednje raziskovalne aplikacije: - Polprevodnik - Fotonske naprave - Izdelava mask - 3D litografija - Difrakcijski optični elementi - Mikrofluidika V nadaljevanju so podane minimalne tehnične zahteve. Ponudba mora izpolnjevati ali presegati vse minimalne tehnične zahteve. | The laser lithography system The University of Maribor is purchasing a versatile high resolution maskless laser lithography system for a wide range of new applications such as microfluidics, micromechanics, or photonics. The system must be suited for microfabrication, photomask fabrication and fast prototyping directly on thin or thick film substrates. The laser lithography system must give the highest degree of freedom to create microstructures in photo sensitive layers. The system must support raster and vector exposure/write modes. The laser lithography system must be equipped with high precision stages for X and Y motion, thus allowing extreme low mechanical errors and fast scan movements. The alignment system must be equipped with a high-resolution camera and processed by advanced imaging software/system, thus supporting applications which require multiple layers. Laser source must support for exposure of broadband and special i-line photoresists, including SU-8, AZ-nLOF, TOK IP and others. The minimum feature/structure size, that can be written with the laser lithography system, is specified in the minimum technical requirements. The system must be suitable for the following research applications: - Semiconductor - Photonic devices - Mask making - 3D Lithography - Diffractive optical elements - Microfluidics Minimum technical requirements are given below. The offer must match or exceed all minimum technical requirements. |
| --- | --- |

| Številka postavke Item No. | Število kosov No. of pieces | ZAHTEVANO | REQUIRED |
| --- | --- | --- | --- |
| 1 | 1 | Minimalne tehnične zahteve: Laserski vir in video sistem: - Valovna dolžina laserskega vira: 375 nm - Velikost strukture: ≤ 2 μm - Življenjska doba laserskega vira: ≥ 10000 ur - Moč laserskega vira: ≥ 60 mW - Avtomatski sistem fokusiranja laserskega vira ali motorizirana mizica s fokusirnim sistemom - Kamera z visoko ločljivostjo za poravnavo z vzorcem: ločljivost vsaj 1600 x 1200 pikslov - Točnost poravnave: ≤ 1 um Miza za vzorce in lastnosti substrata/vzorca: - Motorizirana miza z linearnimi motorji za pomike po X in Y-osi - Delovno področje: ≥ 100 x 100 mm - Ločljivost motorjev mize: ≤ 100 nm - Velikost substrata: od 3 x 3 mm^2; do 125 x 125 mm^2 - Debelina substrata: ≥ 0.1 mm in ≤ 10 mm - Vkauumsko prijemalo za pritditev substrata do dimenzij 125 x 125 mm^2 Zmogljivosti zapisovanja: - Način zapisovanja: vektorski, rasterski ali skanirni in kombinacija obeh - Hitrost zapisovanja v vektorskem načinu: ≥ 100 mm/s - Hitrost zapisovanja v rasterskem ali skanirnem načinu: ≥ 4 mm^2/min - Število nivojev sivin v rasterskem režimu delovanja: ≥ 500 nivojev - Uniformnost dimenzije pri rasterskem ali skanirnem načinu zapisovanja: ≤ 200 nm - Uniformnost linije pri vektorskem načinu zapisovanja: ≤ 50 nm Zmogljivosti sistema za upravljanje za lasersko litografijo: - Programska oprema za nadzor sistema za lasersko litografijo - Podpora za sprotno obdelovanje prednastavljenih receptov - Podpora za pretvorbo formatov: Rastrska slika, TIFF, GDSII, DXF, CIF, GDSII, in Gerber datoteke - Podpora za generiranje 3D datotek iz rastrskih, STEP, STL, SVG, DXF in OBJ datotek - Podpora za risanje osnovnih oblik brez nalaganja izvornih datotek - Podpora za oddaljen dostop in pomoč | Minimum technical requirements: Laser source and video system: - Wavelength: 375 nm - Feature size: ≤ 2 μm - Laser lifetime: ≥ 10000 hours - Laser power/intensity: ≥ 60 mW - Laser auto focus system or a motorized focus stage - High resolution camera system for alignment: ≥ resolution 1600 x 1200 pixels - Alignment accuracy precision: ≤ 1 um Stage system and substrate/sample properties: - X and Y stage system with linear motors - Working area: ≥ 100 x 100 mm - Stage resolution: ≤ 100 nm - Substrate size: from 3 x 3 mm^2; to 125 x 125 mm^2 - Substrate thickness: ≥ 0.1 mm and ≤ 10 mm - Vacuum chunk for clamping substrate with dimensions up to 125 x 125 mm^2 Performance requirements: - Write mode: vector and raster or scanning and combination of both - Write speed in vector exposure mode: ≥ 100 mm/s - Write speed in raster exposure or scanning mode: ≥ 4 mm^2/min - Number of grey levels in raster exposure mode: ≥ 500 levels - Critical dimension uniformity with raster or scanning exposure mode: ≤ 200 nm - Line uniformity with vector exposure mode: ≤ 50 nm The laser lithography system control/software: - Software for controlling the machine - Software for on-the-fly processing of the jobs/recipes - Software for conversion of the following file formats: Bitmaps, TIFF, GDSII, DXF, CIF, GDSII, and Gerber files - Software must support conversion for 3D source files from Grayscale bitmaps and parametric, such as STEP, STL, SVG, DXF and OBJ - Software must support drawing basic shapes without source files - Software must support remote support |

Operacijo sofinancirata Republika Slovenija, Ministrstvo za izobraževanje, znanost in šport in Evropska unija iz Evropskega sklada za regionalni razvoj. Operacija se izvaja v okviru Operativnega programa za izvajanje evropske kohezijske politike 2014-2020 v okviru prednostne osi 1: »Mednarodna konkurenčnost raziskav, inovacij in tehnološkega razvoja v skladu s pametno specializacijo za večjo konkurenčnost in ozelenitev gospodarstva«, prednostne naložbe 1.1 »Krepitev infrastrukture za raziskave in inovacije ter zmogljivosti za razvoj odličnosti na tem področju, pa tudi spodbujanje pristojnih centrov, zlasti takšnih, ki so evropskega pomena«, specifičnega cilja 1.1.1 »Učinkovita uporaba raziskovalne infrastrukture ter razvoj znanja/kompetenc za boljše nacionalno in mednarodno sodelovanje v trikotniku znanja«.

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.