Project name: Upgrading national research infrastructures–RIUM

Technical specifications for equipment for determining the thermodynamic and spectroscopic properties of materials

SUBJECT-MATTER OF THE PUBLIC PROCUREMENT

The subject-matter of the public contract is the supply of equipment for determining the thermodynamic and spectroscopic properties of materials, including placement, assembly, configuration, installation, start-up of the equipment as well as user training, in line with the Technical Specifications.

The public contract is divided into four (4) lots, namely:

- 	Lot 1: CD spectrophotometer capable of measuring absorbance and fluorescence

-	Lot 2: Fluorescence spectrophotometer for measuring fluorescence, phosphorescence, and chemiluminescence

-	Lot 3: UV-Vis spectrophotometer with integrating sphere

-	Lot 4: Equipment for the determination of thermodynamic properties of biological systems (ITC and DSC)

Delivery deadline: no more than 90 days from the date of entry into force of the contract for all lots.

Location for supply, delivery and installation of equipment: University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, room D2-203 for lot 1, lot 2 and lot 3 and room D1-212 for lot 4.

Warranty period: at least two (2) years from the signing of the handover record for all lots.

During the warranty period, the equipment provider has to ensure access to all software updates (firmware, drivers, etc.) for all equipment from a particular lot at no additional cost. All patches, service packs, upgrades and support have to come from the manufacturer of the hardware or software.

Minimum technical requirements are provided below. The tender must meet or exceed all minimum technical requirements for an individual lot.

TECHNICAL REQUIREMENTS

Lot 1: CD spectrophotometer capable of measuring absorbance and fluorescence

The possibility of measuring polarised light in transparent samples or solutions is to investigate the effect of circular dichroism or linear dichroism. Measurement of polarised light and kinetics studies on various chemical samples: determination of protein structure, twisting of macromolecules, DNA, RNA and other polymers, and chiral sites. The flow cell option on the UV-Vis CD allows the HPLC instrument to be assembled to determine complex chemical and biological patterns.

| Item number | Number of pieces | REQUIRED |
| --- | --- | --- |
| 1.1 | 1 | UV-Vis spectrophotometer for measurement of CD and LD with air-cooled light source: - Xenon lamp of minimal strength of 150 W - Wavelength range of 165 nm or lower—950 nm or higher - Resolution: ≤ 0,025 nm - Wavelength accuracy ≤ 0.1 nm at 200 nm in ≤ 0.5 nm at 750 nm - Wavelength reproducibility: ≤ 0.1 nm - CD resolution: 0.00005 mdeg or better - Stray light: ≤ 0.0003 % at 200 nm - Possibility of step scanning - Possibility of various spectral measurements: CD, LD, fluorescence or absorbance (spectral measurements: measurement of time dependence – kinetic measurements, the possibility of constant temperature measurements or measurements at temperature gradient.) - Built-in system for validation of wavelength correctness. |
| 1.2 | 1 | Automated, software-controlled nitrogen gas flow meter |
| 1.3 | 1 | Peltier thermostat cell holder with built-in stirrer for minimally 6 cuvettes and a temperature range from -20 °C or less to +105 °C or more. Temperature measurement should be possible directly in the sample, light output for both transmissions (CD, LD) measurements as well as for fluorescence measurements. If Peltier elements need a circulator for continuity function, the cooling-heating circulator, controlled from the instrument control software should be provided with it. |
| 1.4 | 1 | Set for fluorescence measurements that enables: - Measurement of FDCD - Total fluorescence measurement (detection between 200 nm or lower—800 nm or higher) - Measurement of emitted light spectre (measurement of spectra in wavelength range between 200 nm or lower—800 nm or higher) |
| 1.5. | 1 | Accessories for coupling the CD instrument as a detector to an existing HPLC unit |
| 1.6. | 1 | Automatic titration system: - Titration in the cuvette during spectroscopic measurements - Control with the provided software - System with two syringes |
| 1.7 | 1 | Personal computer: - MS Windows 10 Professional 64 bit operating system - at least 3.6 GHz Intel® Core® i7 processor or equivalent - 16 GB RAM, DVD recorder, 1 TB hard disk and 256 GB SSD or more - 24-inch monitor with Full HD or higher - wired keyboard and wired computer mouse |
| 1.8 | 1 | Software for instrument management, data acquisition and data processing. |
| 1.9 | 6 | Quartz glass cuvettes for CD measurements with 0,2 mm optical path length with an associated spacer for each cuvette. |
| 1.10 | 6 | Quartz glass cuvettes for CD measurements with 1 mm optical path length PTFE plug and with an associated spacer for each cuvette. |
| 1.11 | 6 | Quartz glass cuvettes for CD measurements with 2 mm optical path length PTFE plug and with an associated spacer for each cuvette. |
| 1.12 | 6 | Quartz glass cuvettes for CD measurements with 10 mm optical path length PTFE plug and with an associated spacer for each cuvette. |
| 1.13 | 1 | Additional xenon lamp as a replacement part |
| 1.14 | 1 | Supply, installation and commissioning of equipment in the customer's premises D2-203, UM FKKT |
| 1.15 | 1 | Training for at least five (5) users lasting 2 (two) working days of 8 hours at the location of equipment installation |
| 1.16 | 1 | Warranty of at least two years from the signing of the acceptance report |

Lot 2: Fluorescence spectrophotometer for measuring fluorescence, phosphorescence, and chemiluminescence

Possibility of fluorescence measurement to identify fluorophore - containing compounds. Measurement of fluorescent materials for use in optoelectronics, measurement of excitation barriers (bandgap) in semiconductors, measurement of the concentration of chemical and biological samples with fluorescent markers and observation of cellular and enzymatic samples using fluorescent markers.

| Item nuber | Number of pieces | REQUIRED |
| --- | --- | --- |
| 2.1 | 1 | Fluorescence spectrophotometer: - Excitation and emission wavelength range 200 nm or lower - 850 nm or higher - Light source: xenon lamp - Wavelength accuracy ≤ 1.5 nm - Wavelength repeatability ≤ 0.3 nm - Spectral width of 1.5 nm or less to 20 nm or more - Sensitivity: signal to noise ratio (RMS) ≥ 750: 1 |
| 2.2 | 1 | Polarizer for measuring anisotropy in the UV-Vis range |
| 2.3 | 1 | With Peltier elements thermostat holder for at least 4 cuvettes with a temperature gradient control between 0 ° C or less and 90 ° C or more, with the possibility of mixing the sample in the cuvettes and a sensor for measuring the temperature in the cuvettes. If the system needs heat dissipation for its operation, the supplier must also provide a heat-cooling circulator, which takes care of heat dissipation from Peltier elements, controlled by software and must ensure long-term operation of Peltier elements in the entire temperature range. |
| 2.4 | 4 | Cuvettes and cuvette attachments size 3 x 3 mm for fluorescence measurement with associated spacer for each cuvette |
| 2.5 | 4 | Cuvettes and cuvette attachments of size 5 x 5 mm for fluorescence measurement with associated spacer for each cuvette |
| 2.6 | 4 | Cuvettes and cuvette attachments of size 10 x 10 mm for fluorescence measurement |
| 2.7 | 1 | Personal computer: - MS Windows 10 Professional 64 bit operating system - at least 3.6 GHz Intel® Core® i7 processor or equivalent - 16 GB RAM, DVD recorder, 1 TB hard disk and 256 GB SSD or more - 24-inch monitor with Full HD or higher - wired keyboard and wired computer mouse |
| 2.8 | 1 | Software for instrument management, data acquisition and data processing. |
| 2.9 | 1 | Replacement fluorescence spectrometer lamp |
| 2.10 | 1 | Supply, installation and commissioning of equipment in the customer's premises D2-203, UM FKKT |
| 2.11 | 1 | Training for at least five (5) users lasting 2 (two) working days of 8 hours at the location of equipment installation |
| 2.12 | 1 | Warranty of at least two years from the signing of the acceptance report |

Lot 3: UV-Vis spectrophotometer with integrating sphere

The UV-Vis spectrophotometer will be used for spectrophotometric analysis of solid, powder and liquid samples in the UV and visible light range. Measurement of qualitative and quantitative analyses of organic and inorganic materials.

| Item nuber | Number of pieces | REQUIRED |
| --- | --- | --- |
| 3.1 | 1 | UV-Vis double beam spectrophotometer: - Wavelength range 190 nm or lower - 900 nm or higher - Double monochromator - Wavelength accuracy ≤ 0.08 nm - Wavelength repeatability ≤ 0.025 - Linear range of spectrophotometer from 0 Abs or less to 6 Abs or more - Stray light at 220 nm ≤ 0.00007% - Photometric accuracy ≤ 0.0003 A or better using double aperture method - Photometric repeatability ≤ 0.00008 Abs or better at 0.5 A - Photometric stability ≤ 0.0002 A/h - Photometric noise (RMS) ≤ 0.00005 A at 500 nm, 0 Abs |
| 3.2 | 1 | 150 mm integration sphere for measuring solid and liquid samples with a built - in detector for measuring wavelengths between 200 nm or lower - 900 nm or higher. |
| 3.3 | 1 | With Peltier elements thermostat holder for at least 6 cuvettes with an optical path of 10 mm, with a temperature gradient control between 0 ° C or less and 100 ° C or more, with the possibility of mixing samples in cuvettes and a sensor for measuring cuvette temperature. If the system needs heat dissipation for its operation, the supplier must also provide a heat-cooling circulator, which takes care of heat dissipation from Peltier elements, controlled by software and must ensure long-term operation of Peltier elements in the entire temperature range. |
| 3.4 | 6 | Quartz glass cuvettes (V = 1400 L) with black sides, PTFE plug and 10 mm optical path, which allow measurements in the wavelength range of the UV-Vis spectrophotometer. |
| 3.5 | 6 | Quartz glass cuvettes (V from 1700 to 1750 L) with PTFE plug and 5 mm optical path, which allow measurements in the wavelength range of the UV-Vis spectrophotometer with the corresponding spacer for each cuvette. |
| 3.6 | 6 | Quartz glass cuvettes (V = 700 L) with PTFE plug and 2 mm optical path, enabling measurements in the UV-Vis wavelength range of the spectrophotometer with an associated spacer for each cuvette. |
| 3.7 | 6 | Quartz glass cuvettes (V from 350 to 400 L) with PTFE plug and 1 mm optical path, which allow measurements in the wavelength range of the UV-Vis spectrophotometer with the corresponding spacer for each cuvette. |
| 3.8 | 1 | Personal computer: - MS Windows 10 Professional 64 bit operating system - at least 3.6 GHz Intel® Core® i7 processor or equivalent - 16 GB RAM, DVD recorder, 1 TB hard disk and 256 GB SSD or more - 24-inch monitor with Full HD or higher - wired keyboard and wired computer mouse |
| 3.9 | 1 | Software for instrument management, data acquisition and data processing. |
| 3.10 | 1 | Supply, installation and commissioning of equipment in the customer's premises D2-203, UM FKKT |
| 3.11 | 1 | Training for at least five (5) users lasting 2 (two) working days of 8 hours at the location of equipment installation |
| 3.12 | 1 | Warranty of at least two years from the signing of the acceptance report |

Lot 4: Equipment for the determination of thermodynamic properties of biological systems (ITC and DSC)

The isothermal titration microcalorimeter (ITC) will be used to titrate microlitre amounts of biological protein solutions to determine the affinity of ligands to proteins and the heat released by different interactions in the solution (e.g., micellization). The DSC calorimeter will allow measurements to be carried out on small (microliter) volumes of biological samples (protein solutions and DNA fragments). Determination of heat released/consuming at phase transitions, determination of the melting temperature of proteins in different buffers, and thermal capacity.

| Item number | Number of pieces | REQUIRED |
| --- | --- | --- |
| 4.1 | 1 | Isothermal titration calorimeter (ITC): - The cell is made from chemically inert corrosion-resistant material (e.g., Hastelloy or gold). - Cell volume: 0.2 ml or less. - Temperature range: 5 °C or less to 80 °C or more - Temperature stability at 25 °C: better than 0.15 mK - Calorimeter response time (time constant):10 s or less - Noise level: 2 nW or less. - Titrant syringe should be made of glass or other chemically inert, corrosion-resistant material. The material should be stable in acidic and alkaline solutions and organic solvents. The material should be suitable for work with biomacromolecules. - Titrant syringe volume: 0.04 ml or more. - Injection volume precision: better than 1 % at 2 ml addition - Adjustable stirring speed - The instrument has to be equipped with a cleaning unit that enables automatic cleaning of the instrument (syringe and cell), - Instrument should enable upgrade with auto-sampler, compatible with cleaning unit. The autosampler must enable thermostating of samples at 4 °C. |
| 4.2 | 1 | Differential scanning calorimeter (DSC) - Capillary shaped cell made from chemically inert, corrosion-resistant material. It should enable usage in a alkaline and acidic solutions and resistance to organic solvents. - Cell volume: 0.1—0.4 ml - Temperature range: −10 °C or less to 130 °C or more. - Short term noise 0.02 μW or lower. - Baseline repeatability: 0.1 μW or better. - Calorimeter response time (time constant): 5 s or less - The instrument has to be equipped with a cleaning unit that enables automatic cleaning of the instrument - It should be possible to upgrade the instrument with autosampler. If the producer demands degassing of samples before experiment, degassing station should be included. Detailed description of degassing station is given at 4.4. |
| 4.3 | 1 | Cell cleaning unit - Enables automatic cleaning of the cell and is controlled by instrument software. |
| 4.4 | 1 | Degassing station: - Degassing temperature should be in range between 10 °C or less and 75 °C or more. |
| 4.5 | 1 | ITC test kit |
| 4.6 | 1 | DSC test kit |
| 4.7 | 1 | Additional set of syringes for ITC. |
| 4.8 | 2 | Personal computer: - MS Windows 10 Professional 64 bit operating system - at least 3.6 GHz Intel® Core® i7 processor or equivalent - 16 GB RAM, DVD recorder, 1 TB hard disk and 256 GB SSD or more - 24-inch monitor with Full HD or higher e) wired keyboard and wired computer mouse |
| 4.9 | 1 | Software: - ITC software that enables control of the instrument and treatment of obtained data. - At least two licenses for research and study proposes. |
| 4.10 | 1 | Software: - DSC software that enables control of the instrument and treatment of obtained data. - At least two licenses for research and study proposes. |
| 4.11 | 1 | Supply, installation and commissioning of equipment in the customer's premises D2-203, UM FKKT |
| 4.12 | 1 | Training for at least five (5) users lasting 2 (two) working days of 8 hours at the location of equipment installation |
| 4.13 | 1 | Warranty of at least two years from the signing of the handover record |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. Direct confirmation of the operation is carried out under priority axis 1 of the Operational Program for Implementation of the European Cohesion Policy 2014- 2020: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centres, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle''.

For item 1.7 (personal computer and display or monitor), the Contracting Authority requires the fulfilment of the objective set out in Point 7 of the second paragraph of Article 6 of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17 and 64/19, 49/20 — ZIUZEOP, 152/2020– ZZUOOP, 121/2021).

For item 2.7 (personal computer and display or monitor), the Contracting Authority requires the fulfilment of the objective set out in Point 7 of the second paragraph of Article 6 of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17 and 64/19, 49/20 — ZIUZEOP, 152/2020– ZZUOOP, 121/2021).

For item 3.8 (personal computer and display or monitor), the Contracting Authority requires the fulfilment of the objective set out in Point 7 of the second paragraph of Article 6 of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17 and 64/19, 49/20 — ZIUZEOP, 152/2020– ZZUOOP, 121/2021).

For item 4.8 (personal computer and display or monitor), the Contracting Authority requires the fulfilment of the objective set out in Point 7 of the second paragraph of Article 6 of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17 and 64/19, 49/20 — ZIUZEOP, 152/2020– ZZUOOP, 121/2021).