## technical specifications
Technical specifications: biological x-ray unit for cell and tissue irradiation | provided by the tenderer (YES/NO) | Title of the tehnical documentation | Page number in the technical documentation
The X-ray unit should be suitable for standard cell laboratories, where it will be used for cell cultures that are grown in different size of culture flasks (petri dishes, T25-T75). Therefore it should allow larger filed coverage for lager samples that allows less IR time.
To use it in a standard laboratory, X-Ray unit should be table top designed with dimensions not more than 80cm x 60cm x 65cm and plugs for standard A/C outlet.
The energy for X-ray unit should be in range 10- 150V, with tube current between 0.1 - 6.25 mA and switching off the current should not affect the radiation tube. Maximum tube power should be 940W.
The dosing rate of x-ray tube should offer up to 50 Gy / min (unfiltered dose) and up to 13Gy / min (0.5mm Al). Maximum coverage should be 30 cm. The angle of the X-ray beam should have a maximum divergence of 40 °. The exposure parameters should maintain the reproducibility and accuracy of the experiments in order to allow the comparison between the experiments.
It should offer automatic warm-up and continuous use with no downtime required between the experiments.
It has to be safe before, during and after the dosing with no require additional radiation protection. The dose needs to be exact and monitored with integrated dosimeter with Automatic Dose Control setting.
Even scientist that do not have specialized x-ray knowledge should have to possibility to operate it, therefore it should have simple control panel, preferably touchscreen and a secure system requiring a password and software to monitor all the experiments that are done on X-ray tube.
It should also contain straight forward manuals in English language.
One should be able to download all the experiments done with x-ray unit (history of all the experiments) in office file,like Word or Excel,or equivalent, to a separate computer for the track of the experiments and for further analysis.
All the consumables needed to operate the X-ray machine should be included.

## general conditions
General conditions: | provided by the tenderer (YES/NO)
The life time of the equipment: minimum 10 years.
Warranty period: 1 year
Installation location: NIB, Večna pot 111, Lubljana.
Transportation of equipment to the installation location should be included.
Installation in the foreseen space should be included.
Basic consumables for installation and commissioning of the equipment should be provided.
Packing of the equipment should be included.
Commissioning and hand over of the equipment after onsite training. Site acceptance test (SAT) should consist of declared parameter verification.
Instructions for use should be delivered upon delivery of the equipment.
On-site training should be provided for unlimited number of attendees.
Remote support for experimental setup and troubleshooting during warranty period free of charge.
Validation and calibration or other appropriate inspection of the operation of the equipment before the expiration of the warranty.
Service must be provided by the provider during warranty period and life time of the equipment.
During warranty period the service response time online-remote within 72h, the instrument repair service on-site with new instrument parts within 45 days, or in agreement with the customer. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
When it is necessary to transport the equipment to the service, the transport costs are covered by the supplier within the warranty period.
Replacement with new equipment: if the repair within the warranty period lasts more than 45 days from the reporting of the error, the provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Free software and updates should be provided for the whole life time of the equipment.
Period of availability of spare parts and servicing after warranty period: 8 years.