## technical specifications
Technical specifications: | provided by the tenderer (YES/NO) | Title of the tehnical documentation | Page number in the technical documentation
System for scanning, reading and classification of preparations for Toxicology and Mutagenesis analyses set up for investigations: • in vitro micronucleus test (stained with Giemsa); • in vitro micronucleus test (stained with DAPI); • Comet Assay; • Dicentric chromosome score; • AMES II / AMES MPF analysis.
Consisting of a bright field and fluorescence microscope unit and storage of the results which, based on specific algorithms, must automatically provide the result of the investigation and also allow the reading of the samples manually.
Ability to search and automatically analyze micronuclei, identifiable on mono / binucleated cells automatically recognized by the system, with the possibility of automatic calculation of the CBPI coefficient, in accordance with international guidelines.
Read and quantify micronuclei on both brightfield and fluorescence prepared slides quickly and accurately.
Ability to scan and analyze slides for the comet assay on the same platform.
Ability to scan and analyze slides for the Foci (γH2AX) Test on the same platform.
Ability to perform Ames II / Ames MPF Test on the same platform (possibility to scan multiwell plate).
Possess analysis algorithms compliant with OECD guidelines.
Possibility to scan, acquire and identify dicentric chromosomes on metaphases analyzed in bright field (Giemsa stained) in total automation, with the possibility of relocating the metaphases in live.
Possibility of performing multiple analyzes on the same slide or on a battery of slides in totally automatic way, up to a maximum of 8 positions.
Possess the ability to scan slides in "multi-position" mode, i.e. define multiple areas on the slide to be evaluated with the same algorithms or differentiate according to the application.
Be equipped with a fully motorized microscope on the x, y, z axes (z stacking), with 7-position objective revolvers (5x, 10x, 20x, 63x all fluorite), six-position filter revolver (with Dapi, Green filters , Orange and Red narrow band), including RGB LED system and fluorescent lighting system directly coupled to the stand with long-lasting metal halide technology (2000 hours).
Owning an integrated management system that fully complies with the GDPR requirements, carrying by default and by design all the technological protections designed to satisfy the protection of personal, particular data, also with the use of the solution.
Be upgradable for analysis of additional endpoints and automatic sample feeding.

## general conditions
General conditions: | provided by the tenderer (YES/NO)
The life time of the equipment: minimum 10 years.
Warranty period: 1 year, including all required services, spare parts, labour costs and travelling costs.
Installation location: NIB, Večna pot 111, Ljubljana.
Transportation of equipment to the installation location should be included.
Installation in the foreseen space should be included.
Packing of the equipment should be included.
Basic consumables for installation and commissioning of the equipment should be provided.
Commissioning and hand over of the equipment after onsite training. Site acceptance test (SAT) should consist of declared parameter verification.
Instructions for use should be delivered upon delivery of the equipment.
On-site training should be provided for at least 5 attendees: - at least 1 day for machine operation
On-line or on-site (depending on epidemic situation) training should be provided for unlimited number of attendees: - at least 3 days for software
Remote support for experimental setup and troubleshooting during warranty period free of charge.
Validation and calibration or other appropriate inspection of the operation of the equipment before the expiration of the warranty.
Service must be provided by the provider during warranty period and life time of the equipment.
During warranty period the service response time must be 2 working days and the instrument repair time 7 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the customer. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
When it is necessary to transport the equipment to the service, the transport costs are covered by the supplier within the warranty period.
Replacement with new equipment: if the repair within the warranty period lasts more than 45 days from the reporting of the error, the provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Period of availability of spare parts and servicing after warranty period: 8 years.
Free software updates should be provided for the whole life time of the equipment.