Form No. 2

(Form must be loaded in application e-JN into the section »Pro-forma invoice«)

Price offer (Pro-forma invoice)

On the basis of the invitation to tender »Mechanical material testing system (2 Lots)« we make an offer, as follows:

I. Offer number: _______________

| NAME OF THE TENDERER: |  |
| --- | --- |
| ADDRESS OF THE TENDERER: |  |

II. The Tenderer shall fill in the information only for the lot for which he submits the tender:

Price offer for Lot No. 1:

| Item | Quantity | Value excluding VAT in EUR | VAT value (in % and in EUR) | Value with VAT in EUR |
| --- | --- | --- | --- | --- |
| Hardness testing machine | set |  |  |  |

We declare that the total offer value presented above covers all costs that will incur in connection with the realization of the contract, in the event that our tender is selected.

Price offer for Lot No. 2:

| Item | Quantity | Value excluding VAT in EUR | VAT value (in % and in EUR) | Value with VAT in EUR |
| --- | --- | --- | --- | --- |
| Servo hydraulic fatique testing machine - 1000 kN load frame (for static and dynamic testing) | set |  |  |  |

We declare that the total offer value presented above covers all costs that will incur in connection with the realization of the contract, in the event that our tender is selected.

III. Offer under additional criteria (applies to both lots):

LOT 1:

| Equipment | Criterion | Offered characteristic (to be filled in by the tenderer) |
| --- | --- | --- |
| Research and measurement data acquisition system for capturing electrical quantities | Technical advantage: Samples holder enabling simultaneous assembly of multiple points The technical advantage under evaluation: Higher number of samples which can be fixed into holder (holder positions number) one time for the test. Points are awarded according to the following criterion: - 1 sample position on holder (0 points) - 2-4 samples positions on holder(5 points) - 5 or more samples positions on holder (10 points) | Holder positions number: ___________ |
|  | Delivery time shorter than required Points are awarded according to the following criterion: - 15 weeks or more (0 points) - 10 – 14 weeks (5 points) - Less than 10 weeks (10 points) | Delivery time offered: _________ weeks |
|  | On-site presence of service technician from the error reporting Points are awarded according to the following criterion (in what time the Tenderer can provide a service technician from the day of reporting the error by the Contracting Authority): - 5 working days or more (0 points) - 4 working days (2 points) - 3 working days (5 points) - 2 working days or less (10 points) | On-site presence of service technician from the error reporting offered: __________ working days |

LOT 2:

| Equipment | Criterion | Offered characteristic (to be filled in by the tenderer) |
| --- | --- | --- |
| Servo hydraulic fatique testing machine - 1000 kN load frame (for static and dynamic testing) | Depth of T-slotted platform The technical advantage under evaluation: depth of T-slotted platform Points are awarded according to the following formula: - less than 1650 mm (0 points) - 1650 mm to 1800 mm (10 points) - 1801 mm to 1950 mm (20 points) - more than 1950 mm (30 points) | depth of T-slotted platform offered: ________________ mm |
|  | Column spacing The technical advantage under evaluation: column spacing Points are awarded according to the following formula: - less than 740 mm (0 points) - 740 mm to 790 mm (20 points) - 791 mm to 840 mm (40 points) - more than 840 mm (60 points) | column spacing offered: ________________ mm |
|  | Additional video extensometer based on DIC with FOV (field of view) 150 mm and resolution less than 2,5 µm The technical advantage under evaluation: additional video extensometer based on DIC with FOV (field of view) 150 mm and resolution less than 2,5 µm The Tenderer who offers an additional video extensometer based on DIC with FOV (field of view) 150 mm and resolution less than 2,5 µm receives 120 points. | Additional video extensometer based on DIC with FOV (field of view) 150 mm and resolution less than 2,5 µm offered: YES DIC with FOV (field of view) ____mm and resolution ___ µm NO |
|  | Additional external amplifier with EtherCAT (or comparable) connection with at least 8 input channels (at 24 bits resolution; sample rate more than 5 kHz) The technical advantage under evaluation: additional external amplifier with EtherCAT (or comparable) connection with at least 8 input channels (at 24 bits resolution; sample rate more than 5 kHz) The Tenderer who offers an additional external amplifier with EtherCAT (or comparable) connection with at least 8 input channels (at 24 bits resolution; sample rate more than 5 kHz) receives 30 points. | Ponujen dodatni zunanji ojačevalnik: YES Name of additional external amplifier __________ No. of input channels ____ resolution ___ sample rate ____kHz NO |
|  | Ability to interface with existing HBM MGC Plus external amplifier(s) The technical advantage under evaluation: Ability to interface with existing HBM MGC Plus external amplifier(s) The Tenderer who offers an ability to interface with existing HBM MGC Plus external amplifier(s) receives 5 points. | Ability to interface with existing HBM MGC Plus external amplifier(s): YES NO |
|  | Additional signal conditioners as specified The technical advantage under evaluation: Additional signal conditioners as specified The Tenderer who offers an additional signal conditioners as specified receives 5 points for each, but no more than 40 points. | Additional signal conditioners as specified offered: YES No. of offered: ___________ NO |
|  | Additional analogue IO channels (input) The technical advantage under evaluation: Additional analogue IO channels (input) The Tenderer who offers an additional analogue IO channels (input) receives 2 points for each, but no more than 20 points. | Additional analogue IO channels (input) offered: YES No. of offered: ___________ NO |
|  | Additional servo valve(s) with the same characteristics as the built-in servo valve(s) has - as spare part(s) The technical advantage under evaluation: Additional servo valve(s) with the same characteristics as the built-in servo valve(s) The Tenderer who offers an additional servo valve(s) with the same characteristics as the built-in servo valve(s) receives 15 points for each, but no more than 45 points. | Additional servo valve(s) with the same characteristics as the built-in servo valve(s) has - as spare part(s) offered: YES No. of offered: ___________ NO |
|  | General warranty The minimum requirement for a general warranty is 2 years (0 points). The general warranty offered for the equipment is assessed. | General warranty offered: ____ years |
|  | Delivery time Minimum requirement for the delivery time is not more than 40 weeks (0 points). The delivery time offered is assessed. | Delivery time offered: ____ weeks |
|  | Service response time Minimum requirement for the service response time is not more than 5 working days (0 points). The service response time offered is assessed. | Service response time offered: ____ working days |
|  | Training Minimum requirement for the training included at the site of the final user is at least 5 days (0 points). The training in days offered is assessed. | Training offered: ____ days |
|  | Guaranteed serviceability time Minimum requirement for the guaranteed serviceability time is at least at least 10 years (0 points). The guaranteed serviceability time in years offered is assessed. | Guaranteed serviceability time offered: ____ years |

* If the Tenderer offers equipment that has an individual technical advantage, as stated above, in order to award points within the criteria, it is obligatory to attach a proof (prospectus material, technical documentation,…) to the tender, which will show that the offered equipment really has an individual technical advantage.

IV. Validity period of the offer

The offer is valid for at least 90 days after the deadline for submission of tenders.

Should the validity of the tender be shorter, the tender shall be rejected/excluded from the procedure to award the public contract.

| Place and date: | Name and surname: _____________________ |
| --- | --- |
|  | (stamp and signature) |