Project name: Upgrading national research infrastructures - RIUM

- CORRECTION

DECLARATION ON MEETING THE MINIMUM TECHNICAL REQUIREMENTS OF THE CONTRACTING AUTHORITY AND SPECIFICATIONS OF THE EQUIPMENT OFFERED – LOT 2

Tenderer’s name:

Tenderer’s address:

Subject-matter of the public contract: Equipment for determining the thermodynamic and spectroscopic properties of materials

Procurement code: 302/38 - RIUM51-FKKT 03 - 1/2021

INSTRUCTION:

- This Attachment No. 2 shall be completed and signed. The tenderer must indicate whether the offered equipment meets (or exceeds) the required technical specifications of the offered equipment according to all requirements of the Contracting Authority.

- The tenderer shall load this Attachment No. 2 in Word and PDF format (including attached evidence of compliance with technical requirements) in the e-JN information system in the section "Other Attachments".

We declare that the offered equipment fully meets the minimum technical requirements of the Contracting Authority for the subject-matter of the public procurement entitled Equipment for determining the thermodynamic and spectroscopic properties of materials, and that we are ready to realize the public procurement according to the following specification:

| EQUIPMENT |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Item number | Number of pieces |  | REQUIRED | OFFERED |  | PROOF |  |
| 2.1 | 1 |  | Fluorescence spectrophotometer: |  |  |  |  |
|  |  |  | - Excitation and emission wavelength range 200 nm or less - 850 nm or higher | YES NO |  |  |  |
|  |  |  | - Light source: xenon lamp | YES NO |  |  |  |
|  |  |  | - Wavelength accuracy ≤ 1.5 nm | YES NO |  |  |  |
|  |  |  | - Wavelength repeatability ≤ 0.3 nm | YES NO |  |  |  |
|  |  |  | - Spectral width of 1.5 nm or less to 20 nm or more | YES NO |  |  |  |
|  |  |  | - Sensitivity: signal to noise ratio (RMS) ≥ 750: 1 | YES NO |  |  |  |
| 2.2 | 1 |  | Automatic polarizer for measuring anisotropy in the UV-Vis range | YES NO |  |  |  |
| 2.3 | 1 |  | With Peltier elements thermostat holder for at least 4 cuvettes with a temperature gradient control between 0 ° C or less and 90 ° C or more, with the possibility of mixing the sample in the cuvettes and a sensor for measuring the temperature in the cuvettes. If the system needs heat dissipation for its operation, the supplier must also provide a heat-cooling circulator, which takes care of heat dissipation from Peltier elements, controlled by software and must ensure long-term operation of Peltier elements in the entire temperature range. | YES NO |  |  |  |
| 2.4 | 4 |  | Cuvettes and cuvette attachments size 3 x 3 mm for fluorescence measurement with associated spacer for each cuvette | YES NO |  |  |  |
| 2.5 | 4 |  | Cuvettes and cuvette attachments of size 5 x 5 mm for fluorescence measurement with associated spacer for each cuvette | YES NO |  |  |  |
| 2.6 | 4 |  | Cuvettes and cuvette attachments of size 10 x 10 mm for fluorescence measurement | YES NO |  |  |  |
| 2.7 | 1 |  | Personal computer: |  |  |  |  |
|  |  |  | - MS Windows 10 Professional 64 bit operating system | YES NO |  |  |  |
|  |  |  | - At least 3.6 GHz Intel® Core® i7 processor or equivalent | YES NO |  |  |  |
|  |  |  | - 16 GB RAM, DVD recorder, 1 TB hard disk and 256 GB SSD or more | YES NO |  |  |  |
|  |  |  | - 24-inch monitor with Full HD or higher | YES NO |  |  |  |
|  |  |  | - Wired keyboard and wired computer mouse | YES NO |  |  |  |
| 2.8 | 1 |  | Software for instrument management, data acquisition and data processing. | YES NO |  |  |  |
| 2.9 | 1 |  | Supply, installation and commissioning of equipment in the customer's premises D2-203, UM FKKT | YES NO |  |  |  |
| 2.10 | 1 |  | Training for at least five (5) users lasting 2 (two) working days of 8 hours at the location of equipment installation | YES NO |  |  |  |
| 2.11 | 1 |  | Warranty of at least two years from the signing of the acceptance report | YES NO |  |  |  |
| Requirements ensuring that objectives are met in accordance with the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17, 64/19 and 121/21) for item 2.7 Personal computer: personal computer and display or monitor |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  | MEANS OF PROOF |
|  |  | - The personal computer meets the latest Energy Star energy efficiency standards applicable on the date of publication of the public procurement notice or on the date of the invitation to tender, or equivalent standards. The personal computer falls within the highest energy class available on the market. |  |  | YES NO |  |  |
|  |  | - Personal computer is designed in such a way that the hard drive and the memory are easily accessible using available universal tools (i.e., a screwdriver, spatula, pliers or tweezers) and replaceable. |  |  | YES NO |  |  |
|  |  | - The volume of the personal computer must be measured in accordance with the SIST EN ISO 7779 standard and must not exceed 26 dB at the time of writing to the hard disk. The data must be labelled with the designation LpAm in dB. |  |  | YES NO |  |  |
|  |  | - The display or monitor meets the latest Energy Star energy efficiency standards applicable on the date of publication of the public procurement notice or on the date of the invitation to tender. The monitor falls within the highest energy class available on the market. |  |  | YES NO |  |  |

|  | Stamp: |
| --- | --- |
| Place and date: |  | Signature of the responsible person: |

The operation is co-financed by the Republic of Slovenia, the Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. Direct validation of the operation is carried out under the Operational Programme for the Implementation of the European Cohesion Policy 2014-2020 under Priority Axis 1: "International competitiveness of research, innovation and technological development in line with smart specialisation for greater competitiveness and greening of the economy", investment priority 1.1. infrastructure for research and innovation and capacity to develop excellence in this field, as well as the promotion of competent centres, especially those of European importance ", specific objective 1.1.1 Effective use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle”.

The tenderer must indicate whether the offered equipment meets (or exceeds) the required technical specifications of the offered equipment according to all requirements of the Contracting Authority.

The tenderer must state the name of the attached catalogue or prospectus of the equipment manufacturer and the page number indicating the relevant technical specification. Exceptionally, the tenderer may also provide a link to the equipment manufacturer's website stating the relevant technical specifications. In the case of a link to an online document, the tenderer must provide a page number indicating the relevant technical specification. The tenderer shall enclose to this Attachment No. 2, along with the tender, evidence of compliance with the minimum technical requirements and, in the case of a link to the website, ensure their adequacy and activity in word form.

For item 2.7 (personal computer and display or monitor), the Contracting Authority requires the fulfilment of the objective set out in Point 7 of the second paragraph of Article 6 of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17 and 64/19, 49/20 — ZIUZEOP, 152/2020– ZZUOOP, 121/2021).

For item No. 2.7 Personal computer (personal computer and display or monitor), the tenderer shall indicate the certificate or proof of compliance with the requirements of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17, 64/19 and 121/21) as specified in the Technical Specifications that are part of these procurement documents and submits it with the tender, together with the rest of of the tenderer’s technical documents.