Technical specifications for the Contract: "Procurement of a fully automated device for coating flat substrates with different technological systems and drying methods in the same unit"

Project title: Upgrading of national research infrastructures - RIUM

Public Procurement Code: 302/14-RIUM72-FS12-P18/2022 "Procurement of a fully automated device for coating flat substrates with different technological systems and drying methods in the same unit"

SUBJECT OF THE PUBLIC PROCUREMENT

Subject of the public contract: The supply of Laboratory devices for fully automated surface coating of flat substrates with different technological systems and drying methods in the same unit, including delivery, installation, commissioning of equipment and user training in accordance with the technical specifications.

The device must allow the application of coatings on flat surfaces of various thicknesses and hardnesses, such as textiles, paper, films and foils, membranes, glass, metal, ceramics, etc., and sizes of at least DIN A4.

The device must perform at least the following technological methods of coatings:

- Screen printing, and

- Slot-die coating,

The drying of the coated surfaces must take place in the same unit, and the transition from application to drying must take place automatically, without intermediate intervention by the operator, both in the work area and in the machine control console.

Drying operations must include at least:

- drying with NIR irradiation.

The device must be enclosed in a housing with a built-in LEL detector, which allows safe work even with solvent-based media.

Delivery time: maximum 210 days from the signing of the Contract.

Warranty and maintenance: At least 2 years from the signing of the acceptance report, which must include at least repair or replacement of parts in the case of failure within a maximum of ten (10) working days after the submission of the application, and the rest as follows from the content of the tender documentation. For all equipment, the equipment provider / manufacturer must provide access to all technical and software updates at no additional cost during the warranty period. All repairs, service packs and upgrades and support (hardware and software) must be from the equipment manufacturer. The bidder must provide spare parts and support for at least 5 years after the end of the warranty period.

The provider must provide for the equipment (without additional costs for the client):

- Equipment, as derived from the technical specifications, must be new, packaged and delivered to the required user space, at the same time and complete;

- Appropriate transport of equipment from the manufacturer/supplier to the place of installation, transport insurance, unloading, installation in Laboratory S18/18 (basement, Smetanova 17, 2000 Maribor) internal transport route of approx. 50 m, including staircase of approx. 1.3 m and length approx. 3.15 m (attachment: Photos of the location of the accommodation and transport routes);

- Installation and connection of all elements and equipment units at the University of Maribor, Faculty of Mechanical Engineering, Smetanova ulica 18, 2000 Maribor in Laboratory S18/18;

- Connection of the necessary elements / units for operation and connection to existing installations / systems;

- Compatibility and interoperability of all offered equipment units;

- elimination of all errors that could occur due to possible incompatibility, incompatibility or inadequate connection to a single, flawlessly functioning system;

- Test run in accordance with the technical specifications;

- Demonstration with training of at least 3 users for a period of at least 3 days of 8 hours daily, to work with individual methods of coating and drying, including all hardware and software units, on the delivered subject of the Contract;

- Submit the manufacturer's technical documentation and instructions for handling the equipment in electronic and physical form for the hardware and software and the application programme, all at least in English.

The tender must include all small / supporting material required for installation and commissioning, even if it is not stated explicitly in the technical requirements.

The price of the equipment must be given in Euros and include all the elements of which it consists (equipment units by items and accompanying services, warranty period, any discounts, transport costs, labour costs, delivery and installation, commissioning and demonstration costs, insurance, training costs, instructions, etc.), or all costs related to the implementation of the public procurement in question. The bid value is given by the client's DDP (Incoterms 2010 - Delivery Duty Paid).

The providers must submit a bid for all equipment items. The Contracting Authority will not consider bids that relate only to part of the equipment, and will reject such offers as inadmissible.

TECHNICAL REQUIREMENTS

The equipment consists of a Laboratory unit for coating flat substrates from DIN A10 to at least DIN A4, built-in modules for technologically different coatings (at least screen printing and slot-die coating), and at least an NIR drying irradiation unit, operating continuously in-line in S2S (Sheet-to-Sheet) mode in multiple successive applications, housings for the said unit with integrated LEL detector, and connected to all necessary units for automated operation and monitoring of the coating process (control panel with the appropriate software program), including a system for dosing of coating fluid/ink.

The single offer must include all components that will enable the realisation of all customer requirements.

The equipment will be installed in Laboratory S18/18 of the member of the University of Maribor, Faculty of Mechanical Engineering; in its vicinity are / will be a 230 V socket, a 380 V socket (5-pin, 16-ampere), an Ethernet network connection; a hood (size 3m x 2m) for air extraction, and a compressor for compressed air (6 bar).

The minimum technical requirements are given below. The tender must meet or exceed all minimum technical requirements.

| Item number | Number of pieces | REQUIRED |
| --- | --- | --- |
| 1.1 | 1 | Laboratory unit - System for controlled clamping (alignment), fixing and linear movement of a flat substrate of sizes from DIN A10 to at least DIN A4 and thicknesses up to at least 3 mm, - System for vacuum-controlled fixing and transport of the substrate on the guide with connection to compressed air at a pressure of up to 6 bar, including a control valve and vacuum gauge. - a system using different coating modules, ie. At least: - Screen printing, and - Slot-die coating, - and using at least one drying method, ie.: - NIR irradiation Coating units must be able to coat surfaces of sizes from DIN A10 to at least DIN A4. - A computer-aided automated system that enables: - Stencil / deadlock positioning alignment and real-time monitoring - Dosing the coating liquid / ink according to the module used, the method and speed of coating applied. - A controlled and repeatable sequential method of coating application with selected module and drying with the selected method, in the same unit, without the intermediate intervention of the operator in the process. - Adjustment of the screen / slot offset from the substrate surface in the tolerance range +/- 4 m - Setting the accuracy of repeated / consecutive coatings in the tolerance range +/- 20 m - Setting the coating speed according to the selected module: At least between 0.4 - 4 m / min for the slot-die application mode, and at least between 0.4 - 25 m / min for the print application mode. Software to control and monitor the device in English, that allows to export process data in Excel (at least in CSV format), with an operation display on the operating screen (at least an 8-inch colour touch screen with English navigation) Housing for the Laboratory unit: - Must be made of solvent-resistant stainless material. - Must have lighting. - Must be equipped with an explosive gas and steam detector (Lower Explosive Limit - LEL detector). |
| 1.2 | 1 | Printing module containing: - At least one screen with a 60 mesh size for printing a DIN A4 large surface sample, and at least one pneumatic feedable squeegee / blade. |
| 1.3 | 1 | Slot-die coating module: - Made of stainless steel for applying a layer on a surface of at least 180 mm wide, and with the possibility of using washers (shim masks) for applying at least 50 m thick coat, with a gap clearance accuracy in the tolerance range between +/- 2.5 m. - With a system that allows the application of liquid / ink with a viscosity of at least 1-10,000 mPas and a thickness of at least 0.05-80 m - With a system that enables controlled dosing of liquid / ink into the slot, and, thus, uniformity of application in the tolerance range +/- 5%. |
| 1.4 | 1 | Unit for drying of coatings with NIR irradiation: - With a heating power of 3 transmitters of at least 2.5 kW - With system for controlling the radiation power and time - With glare protection |
| / | / | All devices must be up to 3300 cm (length) x 1500 cm (width) x 2100 cm (height) |
| / | / | The device must enable subsequent / later upgrades (mechanical and software) - By other methods of application, at least by: - Flexo printing mode, - Gravure printing, - Gravure-ofset printing, and - Lamination. - By other means of drying, at least by: - UV radiation based on an LED or mercury light source. without further interference with the size of the device. |
| 1.5 | 1 | Packaging, Transport and Installing in the Laboratory - Machinery mounted on foil-wrapped wooden boards, - Transport of equipment from the manufacturer / supplier to the place of installation with transport insurance, including unloading and installation in the Laboratory |
| 1.6 | 1 | Commissioning and Training Connection of all the necessary elements for operation (connection to the compressed air supply, dosing unit with container for storing the coating liquid / ink and supply to the module system) by checking the operation or test run. Demonstration and training of at least 3 users for a period of at least 3 days of 8 hours daily, to work with individual methods of application and drying, which includes all hardware and software units. |
| 1.7 |  | Warranty period: at least 2 years. |
| 1.8 |  | Provision of spare parts: at least 5 years after the end of the warranty period. |
| 1.9 | 1 | Documentation (in electronic or physical form) for the use and maintenance of the device and its components, at least in English. |
| 1.10 | 1 | Delivery time: maximum 210 days from the entry into force of the contract. |

The operation is co-financed by the Republic of Slovenia, the Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. Direct validation of the operation is carried out under the Operational Programme for the Implementation of the European Cohesion Policy 2014-2020 under Priority Axis 1: "International competitiveness of research, innovation and technological development in line with smart specialisation for greater competitiveness and greening of the economy", investment priority 1.1. infrastructure for research and innovation and capacity to develop excellence in this field, as well as the promotion of competent centres, especially those of European importance ", specific objective 1.1.1 Effective use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle ”.