## Technical Specifications
3D-scanning laser measurement device for contactless vibration measurement of surfaces in three dimensions
|  |  | The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
Sensor specifications
Ability to measure the real 3D vibration on each measurement point in all three spatial dimensions (x-, y-, and z-component) without physical contact with the measured object
Peak velocity measurement range at least 2.5 mm/s–30 m/s (full scale) depending on selected range
Useable frequency bandwidth at least 100 kHz
Laser spot diameter max. 66 µm at 150 mm standoff distance
Laser spot diameter max. 260 µm at 1000 mm standoff distance
Minimum standoff distance for vibration measurement max 125 mm
Weight of one scanning sensor head not exceeding 12 kg
Laser safety class 1 or laser safety class 2 (EN 60825-1)
Improved signal-to-noise ratio on larger working distances and surfaces with poor reflectivity using multi-path interferometer with reception diversity in the optical setup or at least equivalent technology with an Optical Signal Robustness (OSR) > 10000 mm/dropout
Integrated geometry scan unit
Integrated video camera in one of the scanning heads for setting measurement points on the measured objects via the supplied software
Possibility to setup a measurement point grid with at least 500x500 points in one scan cycle
Angular resolution of integrated scanning mirrors less than 0.001°
Autofocus for the laser beams via the data acquisition software
Fully digital signal acquisition without additional D/A and A/D conversion for noise reduction
Data acquisition frontend with at least 8 analog voltage inputs with a useable frequency range of at least 100 kHz
Data acquisition frontend with integrated signal generator with at least 4 outputs
Included tripods and/or frame with sufficient load capacity for fixing the scanning heads
Included computer with data acquisition software
Software specifications
Flexible measurement grid definition
Remote control of at least one external single-point vibrometer used for phase reference
3D geometry scan function
Internal arbitrary signal generator providing standard and user-defined waveforms for excitation
Integrated control of an external reference vibrometer with automatic correction of the phase relationship resulting from the different signal delays of the two vibrometers
High-pass, low-pass, band-pass and notch digital filters for data acquisition
Visualization of measured vibration in the frequency domain
Frequency or band selective 1D and 3D animation for x, y and z vibr. vector separately or combined
User-definable clipping planes and profile cuts
Display and animation in pseudo colors, video image “skin” or imported texture
Editable point selection for display
Analysis and visualization of live or stored data
Import/export formats:
Vibration data: UFF, ASCII, WAV or equivalent
Geometry: UFF, STL or equivalent
ODS graphics and animation: AVI, GIF, JPG, BMP, TIFF, PNG or equivalent
Universal file format data conversion for exchange with modal analysis and finite element software packages
Visual Basic for Applications (VBA) compatibility for automation
API for extraction of raw data via third-party applications
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no

## Splošni pogoji
GENERAL CONDITIONS / SPLOŠNI POGOJI
The life time of the equipment / Življenjska doba opreme:
minimum 10 years / najmanj 10 let
Warranty period / Garancijski rok:
2 years, including all required services, spare parts, labour costs and travelling costs /
2 leti, v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in potni stroški serviserja
Installation location / Lokacija inštalacije:
NIB, Večna pot 111, Ljubljana, base floor (1st level) / NIB, Večna pot 111, Ljubljana, pritličje (1. etaža)
Transportation of equipment to the installation location / Prevoz opreme do lokacije inštalacije:
| included | / vključeno
Installation in the foreseen space / Inštalacija v predviden prostor:
| included | / vključeno
Packing of the equipment / Pakiranje opreme:
| included | / vključeno
Condition of the equipment / Stanje opreme:
the supplied equipment must be new and unused / dobavljena oprema mora biti nova in neuporabljena
Commissioning and hand over of the equipment / Primopredaja in prevzem opreme:
after training, performance check of the equipment by a certified representative of the provider /
po izvedenem usposabljanju se preveri delovanje opreme s strani certificiranega predstavnika izvajalca
Instructions for use / Navodila za uporabo:
handed upon delivery of the equipment / predana ob dobavi opreme
Training / Usposabljanje:
2-day on-site training at delivery / 2-dnevno usposabljanje na lokaciji naročnika ob dobavi
on-site training should be provided for at least 2 attendees / usposabljanje na lokaciji naročnika za vsaj 2 udeleženca
Service / Servis:
service must be provided by the provider during warranty period and life time of the equipment /
servis mora biti zagotovljen s strani izvajalca v času garancijske dobe ter v celotni življenjski dobi opreme
any warranty request must be resolved within 60 working days, or the provider provides an equivalent temporary replacement until repair or replacement / vsaka garancijska zahteva mora biti rešena v 60 delovnih dneh, sicer mora izvajalec zagotoviti enakovredno začasno zamenjavo do popravila opreme oziroma zamenjave
when it is necessary to transport the equipment to the service, the transport costs are covered by the provider within the warranty period / v kolikor je potreben transport opreme na servis, stroški transporta do servisa bremenijo izvajalca v času garancijske dobe
Certificates / Certifikati:
the EU Declaration of conformity (CE) for the measurement system is required / zahteva se izjava EU o skladnosti (CE) za merilni sistem
Period of availability of calibration, spare parts and service for the measurement system /
Doba zagotavljanja kalibracije, rezervnih delov in servisiranja:
at least for the life time of the equipment / vsaj za celotno življenjsko dobo opreme
Free software and updates / Brezplačna programska oprema in njene posodobitve:
should be provided for the whole life time of the equipment /
morajo biti zagotovljene tekom celotne življenjske dobe opreme
yes/da
no/ne
no/ne
no/ne
no/ne
no/ne
no/ne
no/ne
no/ne
no/ne
no/ne
yes/da
yes/da
yes/da
yes/da
yes/da
yes/da
yes/da
yes/da
yes/da
yes/da
no/ne
no/ne
yes/da
no/ne
yes/da
no/ne
yes/da
no/ne
yes/da
no/ne
yes/da
yes/da
no/ne