Project title: Upgrading National Research Infrastructures – RIUM

Declaration of Compliance with the ContRacting Authority’s Minimum Technical Requirements and Specifications of the Offered Equipment

Name of the tenderer:

Address of the tenderer:

Subject of the public procurement: Purchase of a drive test bench system for hardware simulation of interaction between a drive/source and the power grid

Designation of the public procurement: 302/9 – RIUM19P-67-FERI01/2022

INSTRUCTIONS:

- This Attachment No. 2 has to be completed and signed. The tenderer must mark whether the offered equipment is meeting (or exceeding) the required technical specifications of the offered equipment according to all Contracting Authority’s requirements.

- The tenderer uploads this Attachment No. 2 in Word and PDF format (including the enclosed means of proof for meeting the technical requirements) to the e-JN information system, section “Other attachments”.

We declare that the offered equipment fully complies with the minimum technical requirements of the Contracting Authority for the subject matter of the public procurement entitled Purchase of a drive test bench system for hardware simulation of interaction between a drive/source and the power grid and that we are prepared to perform the contract according to the following specification:

|  |  | EQUIPMENT |  |  |
| --- | --- | --- | --- | --- |
| Item No. | No. of units/sets | REQUIREMENTS | OFFERED | DATASHEET |
| 1.1 | 1 unit | Three-phase induction machine (IM): |  |  |
|  |  | - It has a wound three-phase rotor and allows double fed operation of the IM | YES NO |  |
|  |  | - Nominal power: from ≥ 30 kW to ≤ 35 kW (S1 – continuous operation); the nominal power must be equal to the nominal power of the three-phase synchronous machine under item 1.2 | YES NO |  |
|  |  | - Nominal stator voltage: 400 V | YES NO |  |
|  |  | - Number of poles: 4 | YES NO |  |
|  |  | - Nominal frequency: 50 Hz | YES NO |  |
|  |  | - Designed for operation under inverter supply | YES NO |  |
|  |  | - Force-ventilated | YES NO |  |
|  |  | - Built-in temperature monitoring element | YES NO |  |
|  |  | - Maximum machine speed: ≥ 3000 rpm | YES NO |  |
|  |  | - Foot-mounted design (B3) | YES NO |  |
|  |  | - Equipped with an incremental encoder – speed/position sensor with ≥ 5000 pulses per revolution. | YES NO |  |
| 1.2 | 1 unit | Three-phase synchronous machine (SM): |  |  |
|  |  | - It has an excitation winding on the rotor | YES NO |  |
|  |  | - Nominal power: from ≥ 30 kW to ≤ 35 kW (S1 – continuous operation); the nominal power must be equal to the nominal power of the three-phase induction machine under item 1.1. | YES NO |  |
|  |  | - Nominal stator voltage: 400 V | YES NO |  |
|  |  | - Number of poles: 4 | YES NO |  |
|  |  | - Nominal frequency: 50 Hz | YES NO |  |
|  |  | - Power factor range: from 0.1 to 1 (inductive/capacitive in both motoring and generating operation). | YES NO |  |
|  |  | - Designed for operation under inverter supply | YES NO |  |
|  |  | - Force-ventilated | YES NO |  |
|  |  | - Built-in temperature monitoring element | YES NO |  |
|  |  | - Maximum machine speed: ≥ 3000 rpm | YES NO |  |
|  |  | - Foot-mounted design (B3) | YES NO |  |
|  |  | - Equipped with an incremental encoder – speed/position sensor with ≥ 5000 pulses per revolution | YES NO |  |
| 1.3 | 1 unit | Controllable excitation voltage source: |  |  |
|  |  | - Nominal power: it must cover the maximum power required to excite the synchronous machine under item 1.2. | YES NO |  |
|  |  | - Voltage range: must cover the voltage range required to excite the synchronous machine under item 1.2. | YES NO |  |
|  |  | - Current range: must cover the current range allowed to excite the synchronous machine under item 1.2. | YES NO |  |
|  |  | - The setpoints of the excitation voltage can be updated in real-time by the system level controller under item 1.13 or by the programmable controller in the power converter of the synchronous machine under item 1.9. | YES NO |  |
|  |  | - Enables operation of synchronous machine in over- and under-excited state within permissible values of stator and rotor currents | YES NO |  |
| 1.4 | 1 unit | Torque sensor: |  |  |
|  |  | - Nominal torque: from ≥ 250 Nm to ≤ 500 Nm | YES NO |  |
|  |  | - Maximum rotation speed: ≥ 5.000 rpm | YES NO |  |
|  |  | - Accuracy class: ≤ 0.1 % full scale | YES NO |  |
|  |  | - Cut-off frequency: ≥ 2500 Hz. | YES NO |  |
| 1.5 | 1 unit | Interface for adjusting torque/speed sensor measurement signals: |  |  |
|  |  | - Analogue outputs: from -10 V to +10 V | YES NO |  |
|  |  | - Analogue low-pass filter that enables for a constant torque control of the machines | YES NO |  |
| 1.6 | 2 units | Mechanical couplings: |  |  |
|  |  | - Elastomer coupling with split clamping, which allows easy disassembly without moving of the machines or the torque sensor | YES NO |  |
|  |  | - Allows for adjusting the damping/flexibility of the couplings by replacing the elastomer elements | YES NO |  |
| 1.7 | 2 units | Shaft and clutch protection elements: |  |  |
|  |  | - Foldable, with the possibility of easy opening to access the clutch/couplings | YES NO |  |
|  |  | - Built-in safety sensors that are connected to the system level controller under item 1.11, whereas in the open/unprotected state of the shaft the operation of the test bench system is disabled | YES NO |  |
| 1.8 | 1 unit | Machine table for the test bench: |  |  |
|  |  | - Welded steel base frame with leveling feet to which a carrier clamping plate is attached, which enables perfect coaxial installation of the synchronous machine under item 1.2, the torque sensor under item 1.3 and the induction machine under item 1.1. | YES NO |  |
|  |  | - Carrier clamping plate (machine table) length: ≤ 2500 mm | YES NO |  |
|  |  | - Carrier clamping plate (machine table) width: ≤ 900 mm | YES NO |  |
|  |  | - Carrier clamping plate thickness: ≥ 50 mm | YES NO |  |
|  |  | - The carrier clamping plate has adequate T-slots in the longitudinal direction, which serve to align both machines, the torque sensor and the shaft protection elements and enables an adequate coaxial mechanical coupling of the shafts of both machines via a torque sensor. | YES NO |  |
|  |  | - Support elements that enable the height-alignment of shafts of the synchronous machine, the torque sensor and the induction machine and the installation of shaft guards. | YES NO |  |
| 1.9 | 2 sets | Modular system of power electronics converters for powering three-phase AC machines: |  |  |
|  |  | - It is based on back-to-back topology of two three-phase bridge converters with an DC-link energy storage | YES NO |  |
|  |  | - Enables 4-quadrant operation | YES NO |  |
|  |  | - Maximum DC bus voltage: ≥ 800 V | YES NO |  |
|  |  | - Nominal power of the system: ≥ 44 kW (continuous operation at a switching frequency of 10 kHz and a maximum DC bus voltage) | YES NO |  |
|  |  | - Nominal current: ≥ 70 Arms (continuous operation at a switching frequency of 10 kHz) | YES NO |  |
|  |  | - Switching frequencies range of the inverter on the drive side: from ≤ 5 kHz to ≥ 10 kHz | YES NO |  |
|  |  | - Maximum grid voltage: ≥ 480 Vrms | YES NO |  |
|  |  | - Nominal line (grid) frequency: od ≤50 Hz od ≥60 Hz | YES NO |  |
|  |  | - Contains an Active Front End (AFE) filter on the grid side that is tuned to the switching frequency of the inverter on the grid side. | YES NO |  |
|  |  | - Grid inverter switching frequency: between ≥5 kHz and ≤25 kHz and tuned to the AFE filter | YES NO |  |
|  |  | - Enables operation with power factor: from -1 to 1 | YES NO |  |
|  |  | - It has forced air cooling of the three-phase converter modules | YES NO |  |
|  |  | - Measurements of all line currents on both the grid side converter and drive side converter (6x in total) | YES NO |  |
|  |  | - Measurement of the instantaneous values of the DC-link voltage | YES NO |  |
|  |  | - Measurement of the instantaneous values of line currents (3x) and of phase voltages (3x) on the grid side of the AFE filter, which enables power factor control on the grid side of the AFE filter | YES NO |  |
|  |  | - It contains a modular control unit for acquisition, processing and generation of signals based on an open-programmable processor and corresponding peripheral units (i.e., AD and DA converters), which enable closed-loop real-time control of the entire system based on measurements. | YES NO |  |
|  |  | - The control unit is connected to the central computer under item 1.11 and enables the implementation of custom control algorithms (i.e., rapid prototyping - adequate control algorithms can be programmed by the user using block programming and loaded on the programmable processor, where the algorithms are executed in real time). | YES NO |  |
|  |  | - The control unit enables block programming of control algorithms (for example in the software environment Matlab Simulink, Plexim Plecs NI LabView or equivalent) | YES NO |  |
|  |  | - The control unit is acquiring in real time (via digital or analog inputs) at least the following signals: - the pulses from the incremental encoder of the corresponding machine under item 1.1 or 1.2, - torque from the torque sensor under item 1.4, - all 3 line currents of the bridge converter on the drive side, - all 3 line currents of the bridge converter on the grid side, - DC-link voltage, - all 3 line currents on the grid side of the AFE filter, - all 3 phase voltages of the grid, - all 4 line currents on the SM stator and rotor windings, - all 4 phase voltages on the SM stator and rotor windings, - all 6 line currents on the IM stator and rotor windings, - all 6 phase voltages on the IM stator and rotor windings, - temperature of both bridge converter modules, - system control interface reference values under item 1.11, - state of the emergency stop button, - status of the contactor for connecting the converter system to the grid, - states of safety functions (errors) of both converter modules. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO 5) YES NO 6) YES NO 7) YES NO 8) YES NO 9) YES NO 10) YES NO 11) YES NO 12) YES NO 13) YES NO 14) YES NO 15) YES NO 16) YES NO |  |
|  |  | - The control unit is generating (via digital or analog outputs) at least the following output signals: - PWM signals for switching of all transistors in both converter modules, - switching on/off the main contactor for connecting the converter system to the grid, - signals for switching of both of the machines to the different configurations of the system according to technical specification under item 1.10, - excitation voltage reference value for the power supply under item 1.3. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO |  |
|  |  | - It has the function of "soft" synchronization to the grid (with adequate pre-charge circuit) | YES NO |  |
|  |  | - It has at least the following basic hardware or software self-protection functions: - adequate fuses in the circuits, - implementation of dead time in transistor driver circuits, which prevents the possibility of a short-circuit of the DC-link in the three-phase converter modules, - over-current protection of three-phase converter modules, - over-temperature protection of three-phase converter modules. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO |  |
|  |  | - Has implemented basic control algorithms (i.e., 4 quadrant drive operation) containing at least: - control of DC-link voltage and grid-side power factor by means of the grid-side converter, - speed/torque field-oriented control (FOC) of the machine under item 1.1 and 1.2 by means of the drive-side converter. | 1) YES NO 2) YES NO |  |
|  |  | - All components are installed in a protective cabinet that has an emergency shutdown button for the converter system | YES NO |  |
| 1.10 | 1 unit | Machine switching, measurement, and safety system: |  |  |
|  |  | - Contains adequate switches/relays, which, by means of switching manipulations, enable at least the following power supply combinations of the individual windings of the induction and synchronous machine under items 1.1 and 1.2: - The stator winding of the synchronous machine is powered by the inverter system under item 1.9 or directly from the grid. - The rotor winding of the synchronous machine is powered by the power supply under item 1.3 or is not powered. - The rotor winding of the induction machine is powered by the corresponding inverter system under item 1.9 and the stator winding of the asynchronous machine is connected directly to the grid. - The stator winding of the induction machine is powered by the corresponding inverter system under item 1.9 and the rotor winding of the induction machine is short-circuited. - The stator winding of the induction machine is powered directly from the grid and the rotor winding of the induction machine is short-circuited. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO 5) YES NO |  |
|  |  | - Includes adequate fuses/safety devices in all circuits of both machines. | YES NO |  |
|  |  | - Contains adequate wired connections for direct measuring of the voltages at all winding terminals of both machines. | YES NO |  |
|  |  | - Contains 10 current sensors that enable simultaneous measurement of instantaneous values of all line currents of the test specimen (i.e., induction or synchronous machine): - 1 sensor for measuring the excitation current of the synchronous machine (the nominal value must include the maximum value of the current allowed to excite the synchronous machine under item 1.2), - 9 current sensors that enable simultaneous measurement of stator and rotor currents of both machines under item 1.1 and 1.2, - sensors must be able to measure instantaneous values of AC and DC currents, - frequency range: 0 to >= 200 kHz at +-3dB attenuation, - accuracy <= 0.5 % - galvanic isolation of the output signal, - non-battery power supply, - measuring ranges of the sensors must cover corresponding maximum current values of both machines under items 1.1 and 1.2. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO 5) YES NO 6) YES NO 7) YES NO 8) YES NO |  |
|  |  | - Contains 10 voltage sensors that enable simultaneous measurement of instantaneous values of all terminal voltages of the test specimen (i.e., induction or synchronous machine): - 1 sensor for measuring the excitation current of the synchronous machine (the nominal value must include the maximum value of the current allowed to excite the synchronous machine under item 1.2), - 9 corresponding current sensors that enable simultaneous measurement of stator and rotor currents of both machines under item 1.1 and 1.2, - frequency range: 0 to >= 100 kHz at +-3dB attenuation, - accuracy <= 0.5 % - galvanic isolation of the output signal, - non-battery power supply, - measuring ranges of the sensors must cover corresponding maximum current values of both machines under items 1.1 and 1.2. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO 5) YES NO 6) YES NO 7) YES NO |  |
|  |  | - Installed in a protective cabinet | YES NO |  |
| 1.11 | 1 unit | Control and data logging /measurement system: |  |  |
|  |  | - Built-in industrial computer with at least the following specifications: - processor Intel Core i7 2.5 GHz or equivalent or better, - ≥ 16 GB DDR3 RAM, - hard drive SSD ≥ 500 GB, - graphic card, - ≥ 2x network interface 1000/100 MBit/s (Ethernet), - ≥ 3x USB, - ≥ 1x serial port. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO 5) YES NO 6) YES NO 7) YES NO |  |
|  |  | - Central industrial computer, based on dedicated software, enables: - centralized control/automation of the entire test bench system at the system level, - programming of processor units in both modular converter systems under item 1.9 (i.e., rapid prototyping of control algorithms). | 1) YES NO 2) YES NO |  |
|  |  | - Contains an adequate system level controller for the automation of the entire test bench system. | YES NO |  |
|  |  | - Based on the system level controller, it enables acquisition of basic measuring signals in real time, which contain at least: - torque measurement based on the torque sensor under item 1.4, - speed measurement based on incremental encoders mounted on machines under items 1.1 and 1.2, - ≥ 4 channel temperature measurement of machines under items 1.1 and 1.2. | 1) YES NO 2) YES NO 3) YES NO |  |
|  |  | - Based on the system level controller, it enables generation of basic control signals in real time, which contain at least: - reference values for the DC-link voltage control of both converter modules on the grid side under item 1.9, - reference values for speed, current or torque control of both converter modules on the drive side under item 1.9, - reference value of the excitation voltage for the power supply under item 1.3, - signals for activation/deactivation of modular converters under item 1.9. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO |  |
|  |  | - It has implemented, wired, and marked connections for parallel measurement of at least the following measuring quantities (low voltage signals from corresponding sensors): - 3x phase voltage on the grid side of the AFE filter of the modular converter supplying the synchronous machine, - 3x line current on the grid side of the AFE filter of the modular converter that supplies the synchronous machine, - 3x line current on the grid side of the modular converter that supplies the synchronous machine, - 1x DC-link voltage of the modular converter that supplies the synchronous machine, - 3x line current on the drive side of the modular converter that supplies the synchronous machine, - 3x line current on the stator side of the synchronous machine (connected to the corresponding current sensors under item 1.10), - 3x phase voltage on the stator side of the synchronous machine (connected to the corresponding voltage sensors under item 1.10), - 1x line current on the rotor side of the synchronous machine (connected to the corresponding current sensor under item 1.10), - 1x phase voltage on the rotor side of the synchronous machine (connected to the corresponding voltage sensor under item 1.10), - 3x line current on the stator side of the induction machine (connected to the corresponding current sensors under item 1.10), - 3x phase voltages on the stator side of the induction machine (connected to the corresponding voltage sensors under item 1.10), - 3x line current on the rotor side of the induction machine (connected to the corresponding current sensors under item 1.10), - 3x phase voltage on the rotor side of the induction machine (connected to the corresponding voltage sensors under item 1.10), - speed measurement based on the incremental encoder mounted on the synchronous machine under item 1.2, - speed measurement based on the incremental encoder mounted on the induction machine under item 1.1, - torque measurement based on the sensor under item 1.4, - 3x phase voltage on the grid side of the AFE filter of the modular converter supplying the induction machine, - 3x line current on the grid side of the AFE filter of the modular converter that supplies the induction machine, - 3x line current on the grid side of the modular converter that supplies the induction machine, - 1x DC-link voltage of the modular converter that supplies the induction machine, - 3x line current on the drive side of the modular converter that supplies the induction machine. | 1) YES NO 2) YES NO 3) YES NO 4) YES NO 5) YES NO 6) YES NO 7) YES NO 8) YES NO 9) YES NO 10) YES NO 11) YES NO 12) YES NO 13) YES NO 14) YES NO 15) YES NO 16) YES NO 17) YES NO 18) YES NO 19) YES NO 20) YES NO 21) YES NO |  |
|  |  | - It has an Ethernet communication connection between the main computer and the control units in both modular converters, which enables adequate programming of the control units (i.e., enables rapid prototyping of control) | YES NO |  |
|  |  | - It is installed in a protective cabinet | YES NO |  |
|  |  | - The cabinet has a dedicated shelf on which it is possible to place an independent measuring interface for parallel acquisition of measuring signals whereas the interface can be easily connected to the marked connections. | YES NO |  |
| 1.12 | 1 set | Power, signal, and communication connection cables: |  |  |
|  |  | - Supply and installation of all necessary power, signal and communication cables within the test bench system | YES NO |  |
|  |  | - The tenderer must provide adequate functional connections for the full operation of the entire test bench system. | YES NO |  |
|  |  | - Signal and communication cables must be installed separately from power cables in such a way that the influence of EM interference on the measuring signals is adequately reduced. | YES NO |  |
| 1.13 | 1 unit | System-level automation/control software: |  |  |
|  |  | - It is installed on a central industrial computer under item 1.11 | YES NO |  |
|  |  | - Operating system: the latest version of MS Windows | YES NO |  |
|  |  | - Software interface for controlling of the entire test bench at the system level and for capturing of measurement signals, which enables at least: - Choice of different control modes (for example, speed or torque control of an individual machine). - Entering of reference values for speed, torque, DC-link voltage control. - Data logging of basic measured signals given under requirement 1.11b). - Selection of various measuring modules (at least continuous operation S1, short-term operation S2, intermittent operation S3, adjustable load cycles and torque characteristics measurements). | 1) YES NO 2) YES NO 3) YES NO 4) YES NO |  |
| 1.14 | 1 unit | Training: |  |  |
|  |  | - The training is conducted at the location of the University of Maribor, Faculty of Electrical Engineering and Computer Science (FERI), Koroška cesta 46, 2000 Maribor. The supplier must provide training for four (4) equipment users in a minimum of 12 working hours (2 days of 6 hours). | YES NO |  |

| Special requirements for the Built-in industrial computer to the fulfilment of the objectives in accordance with the Decree on green public procurement (Official Journal of the Republic of Slovenia RS, Nos. 51/17, 64/19 and 121/21): (built in computer under 1.11a) |  |  |  |
| --- | --- | --- | --- |
|  |  | MEANS OF PROOF |  |
| 1. Computer complies with the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. | YES NO |  |  |
| 2. Computer is designed in such a way that the hard disk and memory are easily accessible using universally available tools (i.e. with a screwdriver, spatula, pliers, or tweezers) and are replaceable. | YES NO |  |  |
| 3. The volume of the computer is measured in accordance with the SIST EN ISO 7779 standard and does not exceed 26 dB at the time of writing to the hard disk. The data must be labelled with the designation LpAm in dB. | YES NO |  |  |

| Place and date: | Stamp: | Authorized signature: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1 of the Operational Program for Implementation of the European Cohesion Policy 2014–2020: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&I infrastructure and capacity to develop excellence in this field and the promotion of competent centres, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge/competences for better national and international cooperation in the knowledge triangle''.

The tenderer must mark whether the offered equipment is meeting (or exceeding) the required technical specifications of the offered equipment according to all Contracting Authority’s requirements.

The tenderer has to indicate the name of the accompanying equipment manufacturer’s catalogue or brochure and number of the page containing the relevant technical specifications. In exceptional cases, the tenderer may also provide a link to the website of the equipment manufacturer which includes an indication of the relevant technical specifications. Where a link to the online document is provided, the tenderer has to provide the number of the page containing the relevant technical specifications. Upon submitting the tender, the tenderer submits together with Attachment No. 2 means of proof on meeting the minimum technical requirements and in case of stating a link to online documents, provides for accuracy and activity of these in Word format.

For item 1.11 a Built-in industrial computer (personal computer) as means of proof it shall indicate the certificate or proof of compliance with the requirements of the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, No. 51/17, 64/19 and 121/21), as specified in the Technical Specifications that are part of these procurement documents and submit it with the tender, together with the rest of the tenderer’s technical documents.