Project title: Upgrading National Research Infrastructures – RIUM

Designation of the public procurement: 302/9 – RIUM19P-67-FERI01/2022

# Technical specifications for a drive test bench system for hardware simulation of interaction between a drive/source and the power grid

- SUBJECT OF THE PUBLIC PROCUREMENT

Subject of the public procurement is the purchase of a drive test bench system for hardware simulation of interaction between a drive/source and the power grid, including delivery, assembly, configuration, installation, and commissioning of devices in accordance with technical specifications.

Delivery time: at the latest 240 days from the day of entry into force of the contract.

Delivery location: University of Maribor, Faculty of Electrical Engineering and Computer Science, Koroška cesta 46, 2000 Maribor, Slovenia, Building G, Room G-027 (user location), in accordance with the delivery times from the technical specifications.

Training: the training is conducted at the location of the University of Maribor, Faculty of Electrical Engineering and Computer Science (FERI), Koroška cesta 46, 2000 Maribor. In exceptional circumstances, the Supplier may, by special agreement with the User, conduct an introductory training for users also in the online form.

The Supplier has to conduct the training within a period of thirty (30) days after the equipment handover at the User's location. The exact dates of the training sessions will be mutually agreed upon by the Supplier and the User.

The Supplier has to ensure that the people providing the training are professionally qualified.

The training is conducted in Slovenian or English.

The guarantee period: At least 3 years from the day of signature of the take-over protocol (applies to all items of equipment except consumables).

Maintenance and support during the warranty period:

Included at least:

- For devices or components that are subject to the warranty regime, repair, or replacement in case of failure within 60 working days after the submission of the request. Delivery of defective and repaired or replacement parts or devices takes place at the address of the user.

- Delivery of software updates as soon as they are available from the manufacturer. The client can, if the provider allows it, download the updates from the server of the manufacturer or provider.

- Electronic technical support.

- Guaranteed access to the manufacturer's technical support center for all equipment the maintenance of which is the subject of the tender.

- All repairs, service packs, upgrades, and support must be from the hardware or software manufacturer.

Prices must be expressed in euros and must include all the elements of which they are composed (e.g., tendered goods and accompanying services, possible discounts, labor costs, transport costs, insurance, delivery costs, installation costs, packaging waste disposal costs, etc.) or all costs related to the implementation of the public procurement in question.

- TECHNICAL REQUIREMENTS

The desired equipment is a drive test bench system for hardware simulation of interaction between a drive/source and the power grid and is schematically shown in Figure 1. It includes a machine table with mechanically connected induction and synchronous machines, two power electronics power supply systems, a switching-measuring system and a central automation and control system of the entire drive system.

The entire test bench system must meet the following general properties:

- Allows flexible 4-quadrant operation of both machines. One of the machines is in the role of a test specimen, whereas the other represents an active mechanical load or source of mechanical energy; the roles of both machines can be easily interchanged.

- The basic parameters of both machines must be given by the manufacturer (e.g., stator and rotor winding data, individual winding resistances, leakage and magnetizing inductances, transient and sub-transient parameters, ...).

- Both machines have its own modular converter power supply system, as shown in Figure 1.

- Each modular converter system (including the machine) has its own supply cable as shown in Figure 1.

- It contains an adequate switching-measuring system which, in accordance with the technical specifications, enables various connections of the machines either to the corresponding modular converter system or directly to the supply cable, as shown in Figure 1.

- Has implemented adequate connections for performing independent parallel measurements of quantities in accordance with technical specifications.

- The modular converter systems contain programmable control units in accordance with technical specifications, which enable rapid prototyping of custom control algorithms by using block programming.

- All subsystems contain adequate self-protection functions.

- The drive test bench is automated at the system level based on a system level control/monitoring unit and a central computer with a dedicated software interface.

- The test bench system enables regenerative energy use (energy from the grid covers only losses in the test bench system).

The bidder must combine the offered equipment subsystems and interconnect it into a single and seamlessly functioning system (i.e., a turnkey solution), in accordance with the technical specifications, and eliminate any errors that may occur due to possible incompatibility or improper connection to a single, flawless system. The offer must include all the small material needed to connect the devices, even if it is not explicitly stated.

The equipment is procured for the needs of the newly established laboratory. The entire subject of the public procurement will be placed in a space measuring 7 m wide x 8 m long x 4 m high. A door 1.9 m wide and 2.3 m high will be installed in the room. An exemplary layout of the equipment is shown in Figure 2. The subject of the public procurement are devices that are colored green and are proposed to be installed in 4 protective cabinets and 1 machine table in accordance with specifications. All other devices (colored red, blue and orange) are already pre-installed in the given space. A 2.7 m wide strip will be available for the installation of all devices that are the subject of this tender, as shown in Figure 2. There is a 3 x 400 V electrical connection in the room, realized with an existing 4 x 70 mm2 conductor system, whereas adequate overhead cable shelves, underfloor pipes and shafts for installation of cables are already realized as presented in Figure 2. The drive assembly will be connected with two cable systems to a connection/distribution cabinet located on the other side of the room, as spatially shown in Figure 2 and schematically shown in Figure 1. The room is cooled with a 10 kW cooling air conditioner.

Figure 1: Schematic representation of the drive test bench system for hardware simulation of interaction between a drive/source and the power grid

Figure 2: Proposed installation of equipment (only devices painted green are a subject of the public procurement)

The minimum technical requirements are given below. The tender must meet or exceed all minimum technical requirements.

Subject of the public procurement: Drive test bench system for hardware simulation of interaction between a drive/source and the power grid

| Item No. | No. of units/sets | REQUIRED |
| --- | --- | --- |
| 1.1 | 1 unit | Three-phase induction machine (IM), with the following properties: - It has a wound three-phase rotor and allows double fed operation of the IM. - Nominal power: from ≥ 30 kW to ≤ 35 kW (S1 – continuous operation); the nominal power must be equal to the nominal power of the three-phase synchronous machine under item 1.2. - Nominal stator voltage: 400 V. - Number of poles: 4. - Nominal frequency: 50 Hz. - Designed for operation under inverter supply. - Force-ventilated. - Built-in temperature monitoring element. - Maximum machine speed: ≥ 3000 rpm. - Foot-mounted design (B3). - Equipped with an incremental encoder – speed/position sensor with ≥ 5000 pulses per revolution. |
| 1.2 | 1 unit | Three-phase synchronous machine (SM), with the following properties: - It has an excitation winding on the rotor. - Nominal power: from ≥ 30 kW to ≤ 35 kW (S1 – continuous operation); the nominal power must be equal to the nominal power of the three-phase induction machine under item 1.1. - Nominal stator voltage: 400 V. - Number of poles: 4. - Nominal frequency: 50 Hz. - Power factor range: from 0.1 to 1 (inductive/capacitive in both motoring and generating operation). - Designed for operation under inverter supply. - Force-ventilated. - Built-in temperature monitoring element. - Maximum machine speed: ≥ 3000 rpm. - Foot-mounted design (B3). - Equipped with an incremental encoder – speed/position sensor with ≥ 5000 pulses per revolution. |
| 1.3 | 1 unit | Controllable excitation voltage source, with the following properties: - Nominal power: it must cover the maximum power required to excite the synchronous machine under item 1.2. - Voltage range: must cover the voltage range required to excite the synchronous machine under item 1.2. - Current range: must cover the current range allowed to excite the synchronous machine under item 1.2. - The setpoints of the excitation voltage can be updated in real-time by the system level controller under item 1.13 or by the programmable controller in the power converter of the synchronous machine under item 1.9. - Enables operation of synchronous machine in over- and under-excited state within permissible values of stator and rotor currents. |
| 1.4 | 1 unit | Torque sensor, with the following properties: - Nominal torque: from ≥ 250 Nm to ≤ 500 Nm. - Maximum rotation speed: ≥ 5.000 rpm. - Accuracy class: ≤ 0.1 % full scale. - Cut-off frequency: ≥ 2500 Hz. |
| 1.5. | 1 unit | Interface for adjusting torque/speed sensor measurement signals, with the following properties: - Analogue outputs: from -10 V to +10 V. - Analogue low-pass filter that enables for a constant torque control of the machines. |
| 1.6 | 2 units | Mechanical couplings, with the following properties: - Elastomer coupling with split clamping, which allows easy disassembly without moving of the machines or the torque sensor. - Allows for adjusting the damping/flexibility of the couplings by replacing the elastomer elements. |
| 1.7. | 2 units | Shaft and clutch protection elements, with the following properties: - Foldable, with the possibility of easy opening to access the clutch/couplings. - Built-in safety sensors that are connected to the system level controller under item 1.11, whereas in the open/unprotected state of the shaft the operation of the test bench system is disabled. |
| 1.8 | 1 unit | Machine table for the test bench, with the following properties: - Welded steel base frame with leveling feet to which a carrier clamping plate is attached, which enables perfect coaxial installation of the synchronous machine under item 1.2, the torque sensor under item 1.3 and the induction machine under item 1.1. - Carrier clamping plate (machine table) length: ≤ 2500 mm. - Carrier clamping plate (machine table) width: ≤ 900 mm. - Carrier clamping plate thickness: ≥ 50 mm. - The carrier clamping plate has adequate T-slots in the longitudinal direction, which serve to align both machines, the torque sensor and the shaft protection elements and enables an adequate coaxial mechanical coupling of the shafts of both machines via a torque sensor. - Support elements that enable the height-alignment of shafts of the synchronous machine, the torque sensor and the induction machine and the installation of shaft guards. |
| 1.9 | 2 sets | Modular system of power electronics converters for powering three-phase AC machines, with the following properties: - It is based on back-to-back topology of two three-phase bridge converters with an DC-link energy storage, as shown in Figure 1. - Enables 4-quadrant operation. - Maximum DC bus voltage: ≥ 800 V. - Nominal power of the system: ≥ 44 kW (continuous operation at a switching frequency of 10 kHz and a maximum DC bus voltage). - Nominal current: ≥ 70 Arms (continuous operation at a switching frequency of 10 kHz). - Switching frequencies range of the inverter on the drive side: from ≤ 5 kHz to ≥ 10 kHz. - Maximum grid voltage: ≥ 480 Vrms. - Nominal line frequency: from ≤50 Hz to ≥60 Hz. - Contains an Active Front End (AFE) filter on the grid side that is tuned to the switching frequency of the inverter on the grid side. - Grid inverter switching frequency: between ≥5 kHz and ≤25 kHz and tuned to the AFE filter. - Enables operation with power factor: from -1 to 1. - It has forced air cooling of the three-phase converter modules. - Measurements of all line currents on both the grid side converter and drive side converter (6x in total), as shown in Figure 1. - Measurement of the instantaneous values of the DC-link voltage, as shown in Figure 1. - Measurement of the instantaneous values of line currents (3x) and of phase voltages (3x) on the grid side of the AFE filter as shown in Figure 1, which enables power factor control on the grid side of the AFE filter. - It contains a modular control unit for acquisition, processing and generation of signals based on an open-programmable processor and corresponding peripheral units (i.e., AD and DA converters), which enable closed-loop real-time control of the entire system based on measurements. - The control unit is connected to the central computer under item 1.11 and enables the implementation of custom control algorithms (i.e., rapid prototyping – adequate control algorithms can be programmed by the user using block programming and loaded on the programmable processor, where the algorithms are executed in real time). - The control unit enables block programming of control algorithms (for example in the software environment Matlab Simulink, Plexim Plecs, NI LabView or equivalent). - The control unit is acquiring in real time (via digital or analog inputs) at least the following signals (schematically shown in Figure 1): - the pulses from the incremental encoder of the corresponding machine under item 1.1 or 1.2, - torque from the torque sensor under item 1.4, - all 3 line currents of the bridge converter on the drive side, - all 3 line currents of the bridge converter on the grid side, - DC-link voltage, - all 3 line currents on the grid side of the AFE filter, - all 3 phase voltages of the grid, - all 4 line currents on the SM stator and rotor windings, - all 4 phase voltages on the SM stator and rotor windings, - all 6 line currents on the IM stator and rotor windings, - all 6 phase voltages on the IM stator and rotor windings, - temperature of both bridge converter modules, - system control interface reference values under item 1.11, - state of the emergency stop button, - status of the contactor for connecting the converter system to the grid, - states of safety functions (errors) of both converter modules. - The control unit is generating (via digital or analog outputs) at least the following output signals (schematically shown in Figure 1): - PWM signals for switching of all transistors in both converter modules, - switching on/off the main contactor for connecting the converter system to the grid, - signals for switching of both of the machines to the different configurations of the system according to technical specification under item 1.10, - excitation voltage reference value for the power supply under item 1.3. - It has the function of "soft" synchronization to the grid (with adequate pre-charge circuit). - It has at least the following basic hardware or software self-protection functions: - adequate fuses in the circuits, - implementation of dead time in transistor driver circuits, which prevents the possibility of a short-circuit of the DC-link in the three-phase converter modules, - over-current protection of three-phase converter modules, - over-temperature protection of three-phase converter modules. - Has implemented basic control algorithms (i.e., 4 quadrant drive operation) containing at least: - control of DC-link voltage and grid-side power factor by means of the grid-side converter, and - speed/torque field-oriented control (FOC) of the machine under item 1.1 and 1.2 by means of the drive-side converter. - All components are installed in a protective cabinet that has an emergency shutdown button for the converter system. |
| 1.10 | 1 unit | Machine switching, measurement, and safety system, with the following properties: - Contains adequate switches/relays (Figure 1) which, by means of switching manipulations, enable at least the following power supply combinations of the individual windings of the induction and synchronous machine under items 1.1 and 1.2: - The stator winding of the synchronous machine is powered by the inverter system under item 1.9 or directly from the grid. - The rotor winding of the synchronous machine is powered by the power supply under item 1.3 or is not powered. - The rotor winding of the induction machine is powered by the corresponding inverter system under item 1.9 and the stator winding of the asynchronous machine is connected directly to the grid. - The stator winding of the induction machine is powered by the corresponding inverter system under item 1.9 and the rotor winding of the induction machine is short-circuited. - The stator winding of the induction machine is powered directly from the grid and the rotor winding of the induction machine is short-circuited. - Includes adequate fuses/safety devices in all circuits of both machines. - Contains adequate wired connections for direct measuring of the voltages at all winding terminals of both machines, as shown in Figure 1. - Contains 10 current sensors that enable simultaneous measurement of instantaneous values of all line currents of the test specimen (i.e., induction or synchronous machine): - 1 sensor for measuring the excitation current of the synchronous machine (the nominal value must include the maximum value of the current allowed to excite the synchronous machine under item 1.2), - 9 current sensors that enable simultaneous measurement of stator and rotor currents of both machines under item 1.1 and 1.2, - sensors must be able to measure instantaneous values of AC and DC currents, - frequency range: 0 to >= 200 kHz at +-3dB attenuation, - accuracy <= 0,5 % - galvanic isolation of the output signal, - non-battery power supply, - the measuring ranges of the sensors must cover corresponding maximum current values of both machines under items 1.1 and 1.2. - Contains 10 voltage sensors that enable simultaneous measurement of instantaneous values of all terminal voltages of the test specimen (i.e., induction or synchronous machine): - 1 sensor for measuring the excitation voltage of the synchronous machine (the nominal value must include the maximum value of the voltage allowed to excite the synchronous machine under item 1.2), - 9 voltage sensors that enable simultaneous measurement of stator and rotor terminal voltages of both machines under item 1.1 and 1.2, - frequency range: 0 to >= 100 kHz at +-3dB attenuation, - accuracy <= 0,5 % - galvanic isolation of the output signal, - non-battery power supply, - the measuring ranges of the sensors must cover corresponding maximum voltage values of both machines under items 1.1 and 1.2. - Installed in a protective cabinet (schematically shown in Figure 1). |
| 1.11 | 1 unit | Control and data logging /measurement system, with the following properties: - Built-in industrial computer with at least the following specifications: - processor Intel Core i7 2,5 GHz or equivalent or better, - ≥ 16 GB DDR3 RAM, - hard drive SSD ≥ 500 GB, - graphic card, - ≥ 2x network interface 1000/100 MBit/s (Ethernet), - ≥ 3x USB, - ≥ 1x serial port. - Central industrial computer, based on dedicated software, enables: - centralized control/automation of the entire test bench system at the system level, - programming of processor units in both modular converter systems under item 1.9 (i.e., rapid prototyping of control algorithms). - Contains an adequate system level controller for the automation of the entire test bench system. - Based on the system level controller, it enables capturing of basic measuring signals in real time, which contain at least (schematically shown in Figure 1): - torque measurement based on the torque sensor under item 1.4, - speed measurement based on incremental encoders mounted on machines under items 1.1 and 1.2, - ≥ 4 channel temperature measurement of machines under items 1.1 and 1.2. - Based on the system level controller, it enables generation of basic control signals in real time, which contain at least (schematically shown in Figure 1): - reference values for the DC-link voltage control of both converter modules on the grid side under item 1.9, - reference values for speed, current or torque control of both converter modules on the drive side under item 1.9, - reference value of the excitation voltage for the power supply under item 1.3, - signals for activation/deactivation of modular converters under item 1.9. - It has implemented, wired, and marked connections for parallel measurement of at least the following measuring quantities (low voltage signals from corresponding sensors, schematically shown in Figure 1): - 3x phase voltage on the grid side of the AFE filter of the modular converter supplying the synchronous machine, - 3x line current on the grid side of the AFE filter of the modular converter that supplies the synchronous machine, - 3x line current on the grid side of the modular converter that supplies the synchronous machine, - 1x DC-link voltage of the modular converter that supplies the synchronous machine, - 3x line current on the drive side of the modular converter that supplies the synchronous machine, - 3x line current on the stator side of the synchronous machine (connected to the corresponding current sensors under item 1.10), - 3x phase voltage on the stator side of the synchronous machine (connected to the corresponding voltage sensors under item 1.10), - 1x line current on the rotor side of the synchronous machine (connected to the corresponding current sensor under item 1.10), - 1x phase voltage on the rotor side of the synchronous machine (connected to the corresponding voltage sensor under item 1.10), - 3x line current on the stator side of the induction machine (connected to the corresponding current sensors under item 1.10), - 3x phase voltage on the stator side of the induction machine (connected to the corresponding voltage sensors under item 1.10), - 3x line current on the rotor side of the induction machine (connected to the corresponding current sensors under item 1.10), - 3x phase voltage on the rotor side of the induction machine (connected to the corresponding current sensors under item 1.10), - speed measurement based on the incremental encoder mounted on the synchronous machine under item 1.2, - speed measurement based on the incremental encoder mounted on the induction machine under item 1.1, - torque measurement based on the sensor under item 1.4, - 3x phase voltage on the grid side of the AFE filter of the modular converter supplying the induction machine, - 3x line current on the grid side of the AFE filter of the modular converter that supplies the induction machine, - 3x line current on the grid side of the modular converter that supplies the induction machine, - 1x DC-link voltage of the modular converter that supplies the induction machine, - 3x line current on the drive side of the modular converter that supplies the induction machine. - It has an Ethernet communication connection between the main computer and the control units in both modular converters, which enables adequate programming of the control units (i.e., enables rapid prototyping of control). - It is installed in a protective cabinet as shown in Figure 1. - The cabinet has a dedicated shelf on which it is possible to place an independent measuring interface for parallel acquisition of measuring signals whereas the interface can be easily connected to the marked connections. |
| 1.12 | 1 set | Power, signal, and communication connection cables, with the following properties: - Supply and installation of all necessary power, signal and communication cables within the test bench system. - The tenderer must provide adequate functional connections for the full operation of the entire test bench system in accordance with the schematic in Figure 1 and the technical specifications. - Signal and communication cables must be installed separately from power cables in such a way that the influence of EM interference on the measuring signals is adequately reduced. |
| 1.13 | 1 unit | System-level automation/control software, with the following properties: - It is installed on a central industrial computer under item 1.11. - Operating system: the latest version of MS Windows. - Software interface for controlling of the entire test bench at the system level and for capturing of measurement signals, which enables: - Choice of different control modes (for example, speed or torque control of an individual machine). - Entering of reference values for speed, torque, DC-link voltage control. - Data logging of basic measured signals given under requirement 1.11b). - Selection of various measuring modules (at least continuous operation S1, short-term operation S2, intermittent operation S3, adjustable load cycles and torque characteristics measurements). |
| 1.14 | 1 unit | Training: - The training is conducted at the location of the University of Maribor, Faculty of Electrical Engineering and Computer Science (FERI), Koroška cesta 46, 2000 Maribor. The supplier must provide training for four (4) equipment users in a minimum of 12 working hours (2 days of 6 hours). |

|  | Special requirements for item 1.11 a) Built-in industrial computer (personal computer) |
| --- | --- |

The offered supplies must also fully meet the following technical specifications which ensure the fulfilment of the objectives in accordance with the Decree on green public procurement (Official Journal of the Republic of Slovenia RS, Nos. 51/17, 64/19 and 121/21):

Technical specifications: PERSONAL COMPUTERS

REQUIRED

- Personal computer shall comply with the latest Energy Star energy efficiency standards in force on the date of publication of the public tender’s notice or on the date of the invitation for submitting the tender, or equivalent standards. Means of proof: The tenderer submits with the tender: - Energy Star certificate, or - certificate stating that the supply has Type I environmental label, or - manufacturer's technical documentation showing that this requirement is met.

- Personal computer must be designed in such a way that the hard disk and memory are easily accessible using universally available tools (i.e. with a screwdriver, spatula, pliers, or tweezers) and are replaceable. Means of proof: The tenderer submits with the tender: - clear disassembly and repair instructions that allow the products to be disassembled without damage when replacing key components or other parts for upgrading or repair (e.g. exploded view image, showing accessible parts that can be replaced and the necessary tools, video), and - certificate stating that the supply has Type I environmental label, or - suitable proof showing the fulfilments of the demand.

3. The volume of the personal computer must be measured in accordance with the SIST EN ISO 7779 standard and must not exceed 26 dB at the time of writing to the hard disk. The data must be labelled with the designation LpAm in dB. Means of proof: The tenderer submits with the tender: - certificate stating that the supply has Type I environmental label, or - suitable proof showing the fulfilments of the demand.

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.

ENERGY STAR standards: http://www.eu-energystar.org/. The version valid at the time of publication of these environmental examples and criteria is 6.1, and updates are available at http://www.eu-energystar.org/specifications.htm.

Hard Drive Operating.

LpAm (Low power AM) is the designation for the sound emission pressure.