## Technical Specifications
Colony counting system; must be independent and self-sufficient system composed of hardware and software ready to use - 1 piece
| The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
Hardware:
monochrome USB3 camera specifically adopted to produce photos of standard Petri dishes
camera must produce high quality video imaging with fixed focus lenses, with resolution 1392 x 1040 pixels, or comparable quality
lighting system for incident, transmitted and dark-field illumination
plate holders for plates up to 90mm diameter (and 6-well) with incident and transmitted illumination
plate holder for plates 90mm in diameter (and 6-well) with dark-field illumination that ensure optimal illumination of the sample and accurate counting of the colonies
user manual and IQ documentation should be included
system computer; Intel Core i5 (10th Gen) 10500T / 2.3GHz (3.8 GHz) (6-core) computer with 256 GB solid state drive, 8 GB RAM, English UK keyboard, mouse and
Windows 10 Pro 64-bit Edition (or younger version) or equivalent (comparable properties)
two 24” monitors (1920x1200 resolution or comparable properties)
Software for instant image analysis and colony counting:
user firendly and intuitive to use and allowing to build a step-by-step colony counting sequence, allowing operator to set the optimal counting of the colony
counting should be fast, repeatable and with high accuracy
it must include features to help distinguish between touching (separation) and overlapping colonies
high speed analysis, 1 second or less per plate
calibration for precise measurement data should be simple
possibility to save different configurations
produced data/results should be transferable to Excel or in specific software for Ames studies which can assemble the final report
software should be compatible with other software that ensure the smooth running, planning of studies and data collection
compatible with Windows 10 (64-bit versions) or younger edition
licence to register for at least to four individuals
GLP and FDA 21 CFR part 11 compliance
comprehensive auditing of a System Administrator editing user accounts and user activities relating to program access, password revision etc.
uses Oracle database or equivalent to securely hold all user account data and audit trails
measurement of individual objects
count items by size classification
sample measurements should include at least: Plate Count, Extrapolated Plate Count, Field Area, Frame Area, No. objects excluded, Field PPM, Field Perimeter, Mean Grey Level, Total Object Area, Total Object Perimeter, Mean Object Grey Level
yes
yes
yes
yes
yes
no
no
no
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
yes
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no

## General Conditions-EN
GENERAL CONDITIONS
The life time of the equipment:
minimum 10 years.
Warranty period:
1 year, including all required services, spare parts, labour costs and travelling costs.
Installation location:
at the NIB location, Večna pot 111, Ljubljana.
Transportation of equipment to the installation location:
should be included.
Installation in the foreseen space:
should be included.
Packing of the equipment:
should be included.
Basic consumables for installation and commissioning of the equipment:
should be provided.
Commissioning and hand over of the equipment:
after the completed training. Before commissionin and hand over of the equipment test of the
operation of the equipment must be performed. Prior to delivery, the selection of test parameters
to demonstrate individual requirements in the technical specifications shall be agreed.
Instructions for use:
should be delivered upon delivery of the equipment.
Training:
On-line or on-site (depending on epidemic situation) training should be provided for at least 5 attendees:
at least 1 days for machine operation.
On-line or on-site (depending on epidemic situation) training should be provided for unlimited number of attendees:
at least 1 days for software,
at least 1 days for theoretical basis of measurements.
Remote support for experimental setup and troubleshooting during warranty period:
free of charge.
Service:
Service must be provided by the selected tenderer during warranty period and life time of the equipment.
During warranty period the service response time must be 2 working days and the instrument repair time 4 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the contracting authority. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
Service is performed by qualified service tehnicians.
When it is necessary to transport the equipment to the service, the transport costs are covered by the selected tenderer within the warranty period.
Replacement with new equipment:
If the repair within the warranty period lasts more than 45 days from the reporting of the error, the selected tenderer undertakes to replace the equipment or its part with the same or adequate new one within 3 months.
Certificates:
the equipment must be CE certified.
Period of availability of spare parts and servicing after warranty period:
at least 9 years.
Free software and updates:
should be provided for the whole life time of the equipment.
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
yes
no
yes
no
yes
no
yes

## General Conditions-SLO
SPLOŠNI POGOJI
Življenjska doba:
najmanj 10 let.
Garancijski rok:
1 leto: v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in potni stroški serviserja.
Lokacija inštalacije:
NIB, Večna pot 111, Ljubljana.
Prevoz opreme do mesta inštalacije:
mora biti vključen.
Inštalacija v predviden prostor:
mora biti vključena.
Pakiranje opreme:
mora biti vključeno.
Osnovni potrošni material za zagon in inštalacijo opreme:
mora biti vključeno.
Primopredaja in prevzem opreme:
po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti preizkus delovanja opreme. Pred dobavo se dogovori o izboru parametrov preizkusa
za dokazovanje posameznih zahtev v tehničnih specifikacijah.
Delovna navodila za uporabo opreme:
morajo biti predana naročniku ob dostavi opreme.
Usposabljanje:
Usposabljanje kadra naročnika za delo z opremo na lokaciji naročnika ali preko video konference (odvisno od situacije epidemije) za vsaj 5 udeležencev:
najmanj 1 dan za usposabljanje o delovanju opreme.
Usposabljanje kadra naročnika za delo z opremo na lokaciji naročnika ali preko video konference (odvisno od situacije epidemije) mora biti zagotovljeno za neomejeno število udeležencev:
najmanj 1 dan za usposabljanje za uporabo programske opreme,
najmanj 1 dan za usposabljanje o teoretičnih osnovah meritev.
Brezplačna oddaljena pomoč za nastavitev poskusov in odpravljanje težav v času garancijske dobe.
Servis:
Servis mora biti zagotovljen s strani izbranega ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme.
V času garancijske dobe mora biti odzivni čas servisiranja največ 2 delovna dni, čas za popravilo opreme pa največ 4 delovne dni od prijave napake, če je možno popravilo na daljavo. Če je potrebno popravilo na lokaciji naročnika, mora biti čas za popravilo opreme največ 14 dni oziroma po dogovoru z naročnikom. Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok.
Servis se izvaja s strani ustrezno usposobljenih serviserjev.
V kolikor je potreben transport opreme na servis, stroški transporta do servisa bremenijo izbranega ponudnika v času garancijske dobe.
Zamenjava z novo opremo:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev.
Certifikati:
Oprema mora imeti CE certifikat.
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe:
vsaj 9 let.
Brezplačna programska oprema in njene posodobitve:
morajo biti zagotovljene tekom celotne življenjske dobe opreme.
da
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
da
da
da
da
da
da
da
da
da
da
da
da
da
da
da
da
ne
da
ne
da
ne
da
ne
da