## Technical Specifications
Flow cytometer - 1 PIECE
|  | The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
Flow cytometer should include:
3 lasers (488 nm +/- 2nm, 640 nm +/- 2nm nm, 405 nm +/- 2 nm) with possibility of later upgrade.
At least 35 detection channels.
Lasers must be designed as spatially separated (non-collinear).
One Forward Scatter (FSC) detector and granulation (SSC) detectors.
Two Side scatter(SSC) detectors available simultaneously with no optical rearrangement needed.
Instrument must be equipped with Avalanche Photodiodes (APDs) detector modules for enhanced sensitivity.
Flow cytometer should enable:
Continuous detection of fluorescence range at least from 420 nm to 820 nm.
Fluorescence detectors of the instrument must allow fluorescence detection of fluorophores with emission spectra in infrared range up to 900 nm.
Simultaneous detection of at least 35 fluorescent parameters on the immune cells.
Minimum SSC resolution must be 100 nm.
Instrument must have the possibility to be upgraded with advanced SSC detector capable of resolving particles in range 70-80 nm.
Carryover < 0,1%.
Maximal speed of detection: min 30000 events/second.
Instrument must allow to measure absolute concentration without a need for calibration beads.
Instrument must be equipped with automated system for adjustment and sensitivity optimization.
Instrument must allow automated cleaning procedures.
Instrument must allow using deionized water as sheath fluid.
System must have an internal reference d-base of fluorescence emission profiles for individual common fluorophores and reference spectral profiles and posibility of adding own profiles.
Software for spectral unmixing must allow distinguishing of very close emission profiles and using fluorochromes with very close emission maxima simultaneously (e.g. FITC and GFP).
Software must have capability to subtract cellular autofluorescence.
Software must allow manual compensation after signal acquisition.
System must allow daily quality check (QC).
Instrument must allow measuring in a standard 5 ml (12x75 mm) polystyrene and polypropylene cytometry tubes.
Instrument must have autosampler, using microwell plates with 96 well format (with wells of U and V shape, flat bottom and deep well plates) which allows temperature control.
Flow cytometer should fit in available space and supplies:
Benchtop instrument, small footprint. Maximum dimensions for cytometer with autosampler: maximum 80x70x100 cm (width, depth, height).
Environmental parameters in the room: temperature range: 15-35 °C, relative humidity: 40-80 %, no condensing environment.
Hardware and software
Power requirements: 100-240V, 50/60 Hz, 2A.
Data must be exportable to format - Flow Cytometry Standard (FCS) 3.1 (including 18-bit format) for securing data compatibility with third party analytical software.
Workstation with Intel Core i7 processor or equivalent, 3,0 GH or equivalent, 16GB, 2400 MHz DDR4 SO-DIMM or equivalent, display 28'' UHD or equivalent must be provied with instrument.
yes
yes
yes
yes
yes
yes
no
no
no
no
no
no
no
yes
yes
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
no

## General Conditions-EN
GENERAL CONDITIONS
The life time of the equipment:
minimum 10 years.
Warranty period:
1 year, including all required services, spare parts, labour costs and travelling costs.
Installation location:
Installation is performed by the supplier at the NIB location, Večna pot 111, Ljubljana.
Transportation of equipment to the installation location:
should be included.
Installation in the foreseen space:
should be included.
Packing of the equipment:
should be included.
Basic consumables for installation and commissioning of the equipment:
should be provided.
Commissioning and hand over of the equipment:
after completed training and prior to commissioning and hand over of the equipment, it is necessary to perform a test
of the equipment opreation (site acceptance test - SAT or other appropriate test) and calibration. Prior to delivery, the selection of test
parameters to demonstrate individual requirements in the technical specifications shall be agreed.
Instructions for use:
should be delivered upon delivery of the equipment.
Training:
On-site training should be provided for at least 5 attendees:
at least 2 days for machine operation.
On-line or on-site (depending on epidemic situation) training should be provided for unlimited number of attendees:
at least 2 days for software,
at least 3 days for theoretical basis of measurements.
Remote support for experimental setup and troubleshooting during warranty period:
free of charge.
Service:
Service must be provided by the selected provider during warranty period and life time of the equipment.
Service is performed by suitably qualified service technicians.
During warranty period the service response time must be 2 working days and the instrument repair time 24-72 hours from the error report if remote repair is possible and at least 3-7 days if on-site (intervention) repair is needed or in agreement with the contracting authority. Spare parts replacement maximum 30 days. For any repair of a defect longer than 7 days, replacement instrument should be provided. For any repair longer then 14 days, the warranty period is extended accordingly.
Service must be available in a Member State of the European Union.
When it is necessary to transport the equipment to the service, the transport costs are covered by the selected provider within the warranty period.
Replacement with new equipment:
If the repair within the warranty period lasts more than 45 days from the reporting of the error, the selected provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Certificate:
The equipment must be CE certified.
Period of availability of spare parts and servicing after warranty period:
At least 9 years.
Free software and updates:
should be provided for the whole life time of the equipment.
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
yes
no
yes
no
yes
no
yes
no
yes

## General Conditions-SLO
SPLOŠNI POGOJI
Življenjska doba:
najmanj 10 let.
Garancijski rok:
1 leto: v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in potni stroški serviserja.
Lokacija inštalacije:
NIB, Večna pot 111, Ljubljana.
Prevoz opreme do mesta inštalacije:
mora biti vključen.
Inštalacija v predviden prostor:
mora biti vključena.
Pakiranje opreme:
mora biti vključeno.
Osnovni potrošni material za zagon in inštalacijo opreme:
mora biti vključen.
Primopredaja in prevzem opreme:
Po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti preizkus delovanja opreme (test sprejemljivosti in polne funkcionalnosti opreme - SAT
ali drug ustrezen preizkus) in kalibracijo. Pred dobavo se dogovori o izboru parametrov testa za dokazovanje posameznih zahtev v tehničnih specifikacijah.
Delovna navodila za uporabo opreme:
morajo biti predana naročniku ob dostavi opreme.
Usposabljanje kadra za delo z opremo:
Usposabljanje kadra naročnika za delo z opremo na lokaciji naročnika za vsaj 5 udeležencev:
najmanj 2 dni za usposabljanje o delovanju opreme.
Usposabljanje kadra naročnika za delo z opremo na lokaciji naročnika ali preko video konference (odvisno od situacije epidemije) mora biti zagotovljeno za neomejeno število udeležencev:
najmanj 2 dni za usposabljanje za uporabo programske opreme,
najmanj 3 dni za usposabljanje o teoretičnih osnovah meritev.
Brezplačna oddaljena pomoč za nastavitev poskusov in odpravljanje težav v času garancijske dobe.
Servis:
Servis mora biti zagotovljen s strani ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme.
Servis se izvaja z ustrezno usposobljenimi serviserji.
V garancijskem obdobju mora biti odzivni čas servisa 2 delovna dneva in čas popravila instrumenta 24-72 ur od poročila o napaki, če je možno popravilo na daljavo, in najmanj 3-7 dni, če je potrebno popravilo na kraju samem (intervencija) ali v dogovoru z naročnikom. Zamenjava rezervnih delov največ 30 dni. Za vsako popravilo okvare, daljše od 7 dni, je treba zagotoviti nadomestni instrument. Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok.
Servis mora biti na voljo v državi članici Evropske unije.
V kolikor je potreben transport opreme na servis, stroški transporta do servisa bremenijo izbranega ponudnika v času garancijske dobe.
Zamenjava z novo opremo:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev.
Certifikat:
Oprema mora imeti CE certifikat.
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe:
vsaj 9 let.
Brezplačna programska oprema in njene posodobitve:
morajo biti zagotovljene tekom celotne življenjske dobe opreme.
da
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
da
da
da
da
da
da
da
da
da
da
da
da
da
da
da
da
ne
da
ne
da
ne
da
ne
da
ne
da