## Technical specifications
| The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
1. The system should be designed and manufactured as a complete system for behavioral research, morphological measurements and cardiology measurements consisting of software and hardware (observation chamber and computer) and should be suitable for using on various developmental stage of fish (embryo, larvae, adults).
2. Observation chamber should be equipped with high quality IR sensitive digital camera, with a Fresnel lens for a perfectly straight view into each well that do not induce distortion of the image and allow tracking the animals in the dark.
3. Camera and lens must be of high quality; HD (1280 x 1024 pixels), 30 fps maximum for live tracking and 60fps for recording, or equivalent. The optical part of the system should be perfect in resolution, shutter speed, and frame rate for recording the fast “burst-like” movement of zebrafish larvae.
4. Tracking of the animals must be reliable with less than 2% missing samples under standard conditions using accurate method such as detection of the whole animal and finding its center point on each frame, or equivalent. Variables such as velocity and distance moved should be calculated frame to frame.
5. The distance between camera and object should be less than 25cm, that does not induce parallax problems or occlusions in 96 well plates.
6. The system should be adaptable to 150 mm petri dishes (100 mm in standard design).
7. In a observation chamber top light unit must be included with the following specifications; white light: > 30.000 lux, 1584 lm, color temperature 4000K, max. voltage 18 Volt; red light: 192 lm, wavelength 623 nm, max. voltage 24 Volt; Green light : 384 lm, wavelength 525 nm, max. voltage 24 Volt; Blue light : 240 lm, wavelength 470 nm, max. voltage 24 Volt, or equivalent.
8. The created light must be even across all wells, preventing over- or underestimate animal movement, and ensuring there is no light gradient.
9. Temperature control Unit with high precision/accuracy should measure temperature inside of the basin, in the observation chamber, in order to ensure that the temperature of the water in which the fish are exposed, is equal to the set temperature.
a. The temperature should be measured in basin, around the well plate, and not in the more distance units
b. The deviation between the temperature set on the Temperature Control Unit and in the basin should not exceed 0.5 °C
c. Maximum heating capacity at least 10 °C above ambient temperature, within the temperature set point range 15-40 °C
d. Maximum cooling capacity 2.5 °C below ambient temperature, within the temperature set point range 15-40 °C
e. Permissible water temperature range 5 - 50 °C
f. Water capacity ± 350 and pump flow rate – water ± 310 ml/min
10. Tapping device to induce a startle reflex should be included into observation unit
a. The intensity on a scale from 1 to 8 can be defined within the video tracking software
b. Free interval selection in the video tracking software for comparison behavior before and after the event
11. The computer that run smoothly the software and operates the observation chamber must be part of the whole system.
a. The computer has to have at least processor Intel Core i7-11700 or equivalent
b. The computer has to include at least graphics card NVIDIA Quadro P2200 or equivalent
c. Has to run operating system Windows 10 Pro or equivalent (but not older version)
12. The video tracking software must be 100% compatible with observation chamber and:
a. must be in form of Floating license model protected with a hardware key and that can be installed on unlimited number of computers
b. has to include unlimited number of free analysis licenses
c. analysis of raw data and visualization of every data point unfiltered should be possible, as well as application of selected filters according to own justification. The software should allow quality check of actual data by time-based comparison of data and video material, up to single sample points.
d. access to track editor for manual adjustment of data points if necessary
e. well-designed, modern and integrated user interface with clear workflow guidance throughout the experiment
f. possibility to track live and from pre-recorded videos
g. fully automatic start and stop of trials possible, for high-throughput testing
h. detailed functional and technical documentation available – elaborate manuals and online help describing tracking methods, analysis parameters and other functionality
i. user friendly with built-in video tutorial
j. minimum of six data acquisition methods: gray scaling, static subtraction, dynamic subtraction, differencing, color marking and pixel change
k. efficient analysis of large numbers of trials
l. at least three track smoothing algorithms available – Lowess, minimal distance moved, maximum distance moved
m. availability of add-on module to interact with external hardware through TTL and other commands, triggered by the animal’s detected behaviors. Visual scripting language for protocols
n. the software should also be used for video tracking of larger (aquatic) species
o. possibility to measure swimming speed, swim bursts, distance to social stimuli, path shape, calculate shoaling behavior, etc. of the adult fish
p. persistent and re-usable data selection profiles to filter acquired data based on complex combinations of independent and dependent variables
q. persistent and re-usable analysis profiles defining selections of output statistics for various purposes
r. strong academic track record – minimum of 2.000 hits in PubMed Central when searching on brand and product name
13. Additional high quality GigE camera with Kowa lens (no fish-eye effect) or equivalent for recording activity of zebrafish embryos and tracking the behavior of adult zebrafish
14. A separate software for analyzing activity of zebrafish embryos. Next to that, software should be suitable for carrying out the cardiology measurements, blood flow, and morphology measurements of zebrafish larvae, and:
a. Include the possibility to analyze multiple subjects at the same time for high throughput data collection and analysis
b. Should be able to automatically detect the zebrafish embryos
c. Acquire data in a batch from a list of video
yes
yes
yes
yes
yes
no
no
no
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
yes
no
yes
yes
yes
yes
yes
no
no
no
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
yes
yes
no
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no

## General conditions
GENERAL CONDITIONS
The life time of the equipment:
minimum 3 years with regular maintenance according to the manufacturer's instructions
Warranty period:
3 years, including all required services, spare parts, labour costs and travelling costs
Installation location:
at the NIB location, Večna pot 111, Ljubljana
Transportation of equipment to the installation location:
should be included
Installation in the foreseen space:
should be included
Packing of the equipment:
should be included
Basic consumables for installation and commissioning of the equipment:
should be provided
Commissioning and hand over of the equipment:
after the completed training. Before commissionin and hand over the equipment must be tested for operation.
Prior to delivery, the selection of test parameters to demonstrate individual requirements in the technical
specifications shall be agreed.
Instructions for use:
should be delivered upon delivery of the equipment
Training:
On-line or on-site (depending on epidemic situation) training should be provided for at least 5 attendees:
at least 1 days for software
at least 1 days for theoretical basis of measurements
Remote support for experimental setup and troubleshooting during warranty period:
free of charge
Service:
Service must be provided by the selected tenderer during warranty period and life time of the equipment.
During warranty period the service response time must be 2 working days and the instrument repair time 10 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the customer. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
At least one qualified service technician must be available in EU.
When it is necessary to transport the equipment to the service, the transport costs are covered by the selected tenderer within the warranty period.
Certificates:
The equipment must be CE certified (EU declaration of conformity)
Period of availability of spare parts and servising:
At least 3 years
Free software and updates:
should be provided for the whole life time of the equipment (at least 3 years)
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
yes
no
yes
no
yes
no
yes

## Splošni pogoji
SPLOŠNI POGOJI
Življenjska doba:
najmanj 3 leta z rednim vzdrževanjem po navodilih proizvajalca
Garancijski rok:
3 leta, v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in potni stroški serviserja
Lokacija inštalacije:
NIB, Večna pot 111, Ljubljana
Prevoz opreme do mesta inštalacije :
mora biti vključen
Inštalacija v predviden prostor :
mora biti vključen
Pakiranje opreme:
mora biti vključen
Osnovni potrošni material za zagon in inštalacijo opreme:
mora biti vključen
Primopredaja in prevzem opreme :
po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti preizkus delovanja opreme. Pred dobavo se dogovori
o izboru parametrov preizkusa za dokazovanje posameznih zahtev v tehničnih specifikacijah.
Delovna navodila za uporabo opreme:
morajo biti predana naročniku ob dostavi opreme
Usposabljanje kadra za delo z opremo
Usposabljanje kadra naročnika za delo z opremo na lokaciji naročnika ali preko video konference (odvisno od situacije epidemije) mora biti zagotovljeno za VSAJ 5 udeležencev:
najmanj 1 dni za usposabljanje za uporabo programske opreme
najmanj 1 dni za usposabljanje o teoretičnih osnovah meritev.
Brezplačna oddaljena pomoč za nastavitev poskusov in odpravljanje težav v času garancijske dobe
Servis:
Servis mora biti zagotovljen s strani ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme.
V času garancijske dobe mora biti odzivni čas servisiranja največ 2 dni, čas za popravilo opreme pa največ 10 dni od prijave napake, če je možno popravilo na daljavo. Če je potrebno popravilo na lokaciji naročnika, mora biti čas za popravilo opreme največ 14 dni oziroma po dogovoru z naročnikom. Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok.
V državi članici Evropske unije mora biti na voljo vsaj en usposobljen serviser.
V kolikor je potreben transport opreme na servis, stroški transporta do servisa bremenijo izbranega ponudnika v času garancijske dobe.
Zamenjava z novo opremo:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev.
Certifikati
Oprema mora imeti CE certifikat.
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe:
vsaj 3 let
Brezplačna programska oprema in njene posodobitve :
morajo biti zagotovljene tekom celotne življenjske dobe opreme
da
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
da
da
da
da
da
da
da
da
da
da
da
da
da
da
da
da
ne
da
ne
da
ne
da
ne
da