## Technical Specifications
High resolution digital microscope - 1 piece
|  | The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
High resolution digital microscope:
The digital microscope's configuration (microscope's main body, camera, and software) has to be optimized for the research of biological samples, which requires a state-of-the-art camera with high resolution and hardware solutions as specified below.
Main configuration:
The digital microscope stage has to be fully motorized in three directions (X, Y, and Z). It allows for fully automated and motorized image capture for image stitching (X, Y) and stacking (Z).
Magnification range at least between 20x - 2500x
Fully motorized magnification change with rotating head (lenses for different magnifications combined into a single unit - revolver), that helps to eliminate human-induced errors and improves automated image capture.
Motorized Z ax movement has to have at least 0.1 µm resolution, and the travel range has to be at least 4 cm to allow for stacking images of larger biological samples
The digital microscope has to allow the observation of a specimen from different angles. It has to have the tilting option of the Z ax and the rotating table with ± 90° rotating angle. During the rotation or tilting, the focus should always be on the specimen – the specimen is in an eucentric position.
Minimal working distance at the high magnification (2500x) has to be at least 5 mm to allow for the manipulation of microscopic samples in the liquid medium. For smaller magnifications (20x – 500x), the minimal working distance has to be over 2 cm.
The digital microscope must be equipped with a smart and easy-to-use control unit for focus and illumination manipulation.
Lenses of the microscope must be plan apochromatic and telecentric.
The LED light source must be a part of the microscope's body (not an external, separate light source) and enable sample observation in the brightfield and darkfield, with backlight, and polarized light (polarized ring light and backlight). The user can also observe a specimen under a mixed illumination.
Solutions to minimize vibrations.
Processing/display unit, camera, and other
Space-saving, single processing/display unit with at least 27" LCD monitor and 1 Tb of internal storage capacity. This unit must have an option to connect to the internet, connect external mouse and keyboard, and have at least two USB 3 ports (and additional USB 2 ports).
Digital camera with very high resolution (at least 12 Megapixels) with large CMOS sensor (1/1.7 inch or better), with HDR option. The camera and software should also record videos and allow playback.
The system should comprise numerous analytical functions. The following are highly important: Antivibration function, camera shaking corrections, ring-reflection and glare removal, 2D and 3D image-stitching and image-stacking, 2D measurements.
The system should also have a user-friendly interface with functions that allow the user to reproduce image settings quickly, to use auto-focus, to control the illumination fully, to set exposure, gamma, gain, etc. and to use the real-time zoom (1x – 10x)
The user must be able to select a full-screen view or split the screen into numerous smaller ones to compare different sample observations. It is also essential that the software has the navigation function that helps with the orientation (which part of the sample you are observing) for example, when we observe a small part of a specimen with high magnification.
The system should have a function to easily generate reports (export at least in Excel) and prepare publication-grade images or multi-image composites. Therefore, the system should have image enhancement functions (e.g., color enhancement for research on micro-plastics, 3D shape correction…), comparison functions (e.g., 3D comparison) and high-quality depth composition.
The system must have advanced functions for a precise surface inspection and finding patterns (Optical shadow effect or similar).
The system must enable unlimited communication between the digital microscope and the user’s own personal computer. For example, the user has to be able to remotely connect to the digital microscope and have unlimited access to images (including metafiles).
The system must run on Windows 10 operating system and must have system protection (antivirus).
We will conduct varying research on the digital microscope so the software has to enable user specific memory (user specific settings).
The system must allow for timer-set image capture. This allows observing changes in biological material in preset time intervals. This kind of image capture must be performed with autofocus.
yes
yes
yes
no
no
no
yes
no
yes
yes
no
yes
no
yes
no
yes
no
yes
no
yes
yes
yes
yes
no
no
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
no
yes
no

## General Conditions
GENERAL CONDITIONS
The life time of the equipment:
minimum 8 years with regular maintenance according to the manufacturer's instructions
Warranty period:
1 year, including all required services, spare parts, labour costs and travelling costs
Installation location:
NIB, Večna pot 111, Ljubljana
Transportation of equipment to the installation location:
should be included
Installation in the foreseen space:
should be included
Packing of the equipment:
should be included
Basic consumables for installation and commissioning of the equipment:
should be provided
Commissioning and hand over of the equipment:
after onsite training and prior to commissioning and hand over of the equipment it is necessary to perform a test of
the operation of the equipment (test of acceptability and full functionality of the equipment - SAT or other appropriate test)
Prior to the delivery of the equipment, the selection of test parameters for proving individual requirements in technical specifications shall be agreed.
Instructions for use:
should be delivered upon delivery of the equipment
Training:
On-site training should be provided for at least 10 attendees:
at least 1 day for machine operation and 1 day for software
Additional on-line or on-site training should be provided:
after major software updates
Remote support for experimental setup and troubleshooting during warranty period:
free of charge
Service:
Service must be provided by the provider during warranty period and life time of the equipment.
Service is performed by qualified service technicians.
During warranty period the service response time must be 2 working days and the instrument repair time 4 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the customer. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
When it is necessary to transport the equipment to the service, the transport costs are covered by the supplier within the warranty period.
Replacement with new equipment:
If the repair within the warranty period lasts more than 45 days from the reporting of the error, the provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Period of availability of spare parts and servicing after warranty period:
7 years
Compatibility:
All components of the system (digital microscope, processing/display unit, camera ...) must be from the same producer to guarantee
compatibility and optimal system functionality. The technical specifications have to be officially available and verifiable.
Free software and updates:
should be provided for the whole life time of the equipment
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
yes
no
yes
no
yes
yes
no

## Splošni pogoji
SPLOŠNI POGOJI
Življenjska doba:
najmanj 8 let ob rednem vzdrževanju po navodilih proizvajalca
Garancijski rok:
1 leto, v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in potni stroški serviserja
Lokacija inštalacije:
NIB, Večna pot 111, Ljubljana
Prevoz opreme do mesta inštalacije:
mora biti vključen
Inštalacija v predviden prostor:
mora biti vključena
Pakiranje opreme:
mora biti vključeno
Osnovni potrošni material za zagon in inštalacijo opreme:
mora biti vključen
Primopredaja in prevzem opreme:
po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti preizkus delovanja opreme (test sprejemljivosti in polne funkcionalnosti opreme - SAT ali drug ustrezen preizkus). Pred dobavo opreme se dogovori o izboru parametrov testa za dokazovanje posameznih zahtev v tehničnih specifikacijah.
Delovna navodila za uporabo opreme:
morajo biti predana naročniku ob dostavi opreme.
Usposabljanje kadra naročnika za delo z opremo:
Na lokaciji naročnika za najmanj 10 udeležencev:
najmanj 1 dan za usposabljanje o delovanju opreme in 1 dan za programsko opremo
Dodatni trening na lokaciji naročnika ali na spletu:
po večjih posodobitvah programske opreme
Oddaljena pomoč za nastavitev poskusov in odpravljanje težav v času garancijske dobe:
brezplačno
Servis:
Mora biti zagotovljen s strani ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme.
Servis se izvaja s strani ustrezno usposobljenih serviserjev.
V času garancijske dobe mora biti odzivni čas servisiranja največ 2 delovna dni, čas za popravilo opreme pa največ 4 delovne dni od prijave napake, če je možno popravilo na daljavo. Če je potrebno popravilo na lokaciji naročnika, mora biti čas za popravilo opreme največ 14 dni oziroma po dogovoru z naročnikom. Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok.
Če je potreben transport opreme na servis, stroški transporta do servisa bremenijo dobavitelja v času garancijske dobe.
Zamenjava z novo opremo:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev.
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe:
7 let.
Kompatibilnost:
Vse komponente sistema (digitalni mikroskop, procesna/prikazovalna enota, kamera...) morajo biti od istega proizvajalca,
da je zagotovljena združljivost in optimalno delovanje sistema. Tehnične specifikacije morajo biti uradno dostopne in preverljive.
Brezplačna programska oprema in njene posodobitve:
morajo biti zagotovljene tekom celotne življenjske dobe opreme.
DA
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
NE
DA
NE
DA
NE
DA
NE
DA
NE