Project title: Upgrading National Research Infrastructures – RIUM

COMPLIANCE DECLARATION WITH MINIMUM TECHNICAL REQUIREMENTS

OF THE OFFERED EQUIPMENT

Name of the tenderer:

Address of the tenderer:

Subject of the public procurement: Purchase of a system for working with semiconductor photonic and other slices – Lot 3

Designation of the public procurement: 302/20 – RIUM8-FERI03.5/2020

| Is the tenderer also the manufacturer of the offered equipment? | YES ☐ | NO ☐* |
| --- | --- | --- |
| *If the answer is NO, enter the equipment manufacturer's name and address: |  |  |

INSTRUCTIONS:

- This Annex No. 2 must be completed and signed. The tenderer must indicate the actual technical specifications of the equipment offered in line with all requirements.

- The tenderer shall upload this Annex No. 2 to the e-JN information system in the section “Other Attachments”.

We declare that the offered equipment fully meets the minimum technical requirements of the client for the public procurement titled "Purchase of a system for working with semiconductor photonic and other slices – Lot 3", and that we are ready to realize the order according to the following specification:

| EQUIPMENT |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  |  |  |  |  | DATASHEET |
| 1 | Basic properties: |  |  |  |  |  |  |  |
|  | - Die bonder size: Bench Top | - YES |  |  | - NO |  |  |  |
|  | - System size ≤ 80x75x50 cm (width, depth, height) | Width: _____________ cm, Depth: _____________ cm, Height: _____________ cm |  |  |  |  |  |  |
|  | - 210 to 230 VAC power supply with EU power plug. | - YES |  | - NO |  |  |  |  |
|  | - Weight ≤ 45 kg | Weight: _____________ kg |  |  |  |  |  |  |
|  | - User interface for adjustment of parameters | - YES |  |  | - NO |  |  |  |
|  | Pick and place: |  |  |  |  |  |  |  |
|  | - Pick and place of workpieces ≤ 200x200 um | - YES |  | - NO |  |  |  |  |
|  | - Integrated or included Vacuum pump | - YES |  |  | - NO |  |  |  |
|  | - Motorized or manual counter-weighted Z-axis movement | - YES |  |  | - NO |  |  |  |
|  |  | Z-axis movement ≥ 15 mm. |  |  | Z-axis movement: ____________ mm |  |  |  |
|  | - Micrometer control for X-axis and Y-axis | - YES |  |  |  | - NO |  |  |
|  |  | X-axis movement ≥ 8 mm |  |  |  | X-axis movement: ____________ mm |  |  |
|  |  | Y-axis movement ≥ 8 mm |  |  |  | Y-axis movement: ____________ mm |  |  |
|  | - Angle alignment | - YES |  | - NO |  |  |  |  |
|  | - Min. bonding force ≥ 2 cN | Force: ___________ cN |  |  |  |  |  |  |
|  | - Max. Bonding force ≤ 200 cN | Force: ___________ cN |  |  |  |  |  |  |
|  | Adhesion: |  |  |  |  |  |  |  |
|  | - Epoxy stamping. | - YES |  | - NO |  |  |  |  |
|  | - Needle dispensing | - YES |  | - NO |  |  |  |  |
|  |  | Holder for syringe | - YES |  |  |  | - NO |  |
|  | Magnifying system: |  |  |  |  |  |  |  |
|  | - Stereo-Microscope or digital microscope with screen | - YES |  |  | - NO |  |  |  |
|  | - Light Illumination of the sample. | - YES |  |  | - NO |  |  |  |

| Initial consumables and tools: (datasheet not required) |  |  |
| --- | --- | --- |
| - Pick and place tool – needle tip with outer diameter ≤ 300 um and hole diameter ≤ 150 um | - YES | - NO |
| - Stamping tool with surface area ≤ 1,5 mm2 | - YES | - NO |
| - Tools for maintenance | - YES | - NO |
| - User Manual | - YES | - NO |

| Place and date: | Stamp: | Authorized signature: |
| --- | --- | --- |

The operation is co-financed by the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1 of the Operational Program for Implementation of the European Cohesion Policy 2014–2020: “International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy”, priority investment 1.1 “Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance”, specific objective 1.1.1 “Efficient use of research infrastructure and development of knowledge/competences for better national and international cooperation in the knowledge triangle”.

The tenderer must indicate the actual technical specifications of the equipment offered in accordance with all requirements.

The tenderer must give the name of the attached catalogue or brochure of the equipment manufacturer and the page number(s), together with the relevant technical specification. In exceptional cases, the tenderer may also provide a link to the equipment manufacturer's website with the relevant technical specification. Where a link to the online document is provided, the tenderer shall indicate a page number, indicating the relevant technical specification.