## Technical Specifications
Digital polymerase chain reaction (dPCR) platform with high multiplexing capability
|  | The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
dPCR platform should enable:
automated partitioning of the reaction
partitioning the reaction in at least 20.000 partitions
multi-channel multiplexing with 6 colour channels
dynamic range of 5-log
use of both hydrolysis probe base chemistry (e.g. TaqMan) and intercalating dye based chemistry
use of non-proprietary mastermixes (i.e. enables the use of mastermixes other than manufacturer of the instrument)
recovery of the sample after PCR
Software should enable:
export of data is at least cvs and Microsoft Excel formats
spillover or colour compensation
visualization of results in 1D, 2D and 3D
full access to raw data, and enables visualization of results down to a single partitions after PCR
Hardware
enables UPS backup
apart from dPCR platform, a computer (PC or laptop) with licence-free software should be included
yes
yes
yes
yes
yes
yes
no
no
no
no
no
no
no
no
yes
no
yes
yes
yes
no
yes
yes
no
no
yes
no

## General Conditions
SPLOŠNI POGOJI/GENERAL CONDITIONS
Granacijski rok/Warranty period:
1 leto, v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in
potni stroški serviserja/1 year, including all required services, spare parts, labour costs and travelling costs
Lokacija inštalacije/Installation location:
NIB, Večna pot 111, Ljubljana, pritličje/ground floor
Prevoz opreme do mesta inštalacije/Transportation of equipment to the installation location:
mora biti vključen/should be included
Inštalacija v predviden prostor/Installation in the foreseen space:
mora biti vključena/should be included
Osnovni potrošni material za zagon in inštalacijo opreme/Basic consumables for installation and commissioning of the equipment:
mora biti zagotovljen/hould be provided
Primorpedaja in prevzem opreme/Commissioning and hand over of the equipment:
po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti preizkus delovanja opreme (test sprejemljivosti in polne funkcionalnosti - SAT ali drug ustrezen preizkus). Pred dobavo opreme se dogovori o izboru parametrov testa. / after onsite training at the installation location. Before hand over, site acceptance test (SAT) or other comparable test should be done in order to assess proper functioning of the equpment. Test parametrs should be chosen/declared prior to instalation.
Delovna navodila za uporabo opreme/Instructions for use:
morajo biti predana naročniku ob dostavi opreme/should be delivered upon delivery of the equipment
Uposabljanje/Training:
Na lokaciji naročnika mora biti usposabljanje zagotovljeno za vsaj 5 udeležencev / on-site training should be provided for at least 5 attendees:
vsaj 2 dni za delovanje opreme, vključno z analizo podatkov / at least 2 day for machine operation
including data analysis
Oddaljena podpora za poskusno nastavitev in odpravljanje težav v garancijskem obdobju/Remote support for experimental setup and troubleshooting during warranty period:
brezplačno/free of charge
Servis/Service:
Servis mora biti zagotovljen s strani ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme. / Service must be provided by the provider during warranty period and life time of the equipment.
Servis se izvaja s strani ustrezno usposobljenih serviserjev. / Service is performed by qualified service technicians.
V času garancijske dobe mora biti odzivni čas servisiranja največ 2 delovna dni, čas za popravilo opreme pa največ 4 delovne dni od prijave napake, če je možno popravilo na daljavo. Če je potrebno popravilo na lokaciji naročnika, mora biti čas za popravilo opreme največ 14 dni oziroma po dogovoru z naročnikom. Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok. / During warranty period the service response time must be 2 working days and the instrument repair time 4 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the customer. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
V kolikor je potreben transport opreme na servis, stroški transporta do servisa bremenijo dobavitelja v času garancijske dobe. / When it is necessary to transport the equipment to the service, the transport costs are covered by the supplier within the warranty period.
Zamenjava z novo opremo/Replacement with new equipment:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev. / If the repair within the warranty period lasts more than 45 days from the reporting of the error, the provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe/Period of availability of spare parts and servicing after warranty period:
5 let/5 years
Programska oprema in njene posodobitve/Free software and updates:
morajo biti brezplačno zagotovljene tekom celotne življenjske dobe opreme/should be provided for the whole life time of the equipment.
ne/no
ne/no
ne/no
ne/no
ne/no
ne/no
ne/no
ne/no
ne/no
ne/no
ne/no
ne/no
ne/no
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
da/yes
ne/no
da/yes
ne/no
da/yes
ne/no