## Technical Specifications
Laboratory Chamber Furnace
|  | The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
Laboratory Chamber Furnace should enable:
suitable for laboratory use and placement on the laboratory counter
The laboratory chamber furnace should fit the available space and have the following requirements:
External dimensions maximum: 600mm x 750mm x 900 mm (width x depth x height)
Weight: Maximum 100 kg
Internal dimensions at least: 320mm x 490mm x 250 mm (width x depth x heigh)
Volume: at least 39 L
Temperature
Maximum temperature: up to 1100 ° C
Heating time to maximum temperature: maximum 100 min
Temperature uniformity in an empty oven at least ± 5 K
Heating in at least two zones
Control and screen
Guaranteed oven control screen with the possibility of setting at least four programs with four adjustable segments
Option to lock the controller
Possibility to set the program to 1 ° C and 1 minute
Possibility of entering ramps
Adjustable program language - at least English
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
no

## General Conditions
GENERAL CONDITIONS
The life time of the equipment:
minimum 10 years
Warranty period:
1 year, including all required services, spare parts, labour costs and travelling costs
Installation location:
At the temporary location of MBP, Fornače 41, 6330 Piran
Transportation of equipment to the installation location:
should be included
Installation in the foreseen space:
should be included
Packing of the equipment:
should be included
Basic consumables for installation and commissioning of the equipment:
should be provided
Commissioning and hand over of the equipment:
after onsite training and prior to commissioning and hand over, it is necessary to perform validation / testing at the
contracting authority's location, so that after installation the analyzer is ready for use in all required applications.
Instructions for use:
should be delivered upon delivery of the equipment
Training:
On-site training should be provided:
at least 1 day for machine operation
Remote support for experimental setup and troubleshooting during warranty period: free of charge
Service:
Service must be provided by the provider during warranty period and life time of the equipment.
The provider must provide authorized service by service technicians at the site of installation of the furnace or at the location of the contracting authority.
Servise is performed by suitably qualified service technicians.
During warranty period the service response time must be 3 working days and the instrument repair time 14 working days from the error report. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
When it is necessary to transport the equipment to the service, the transport costs are covered by the supplier within the warranty period.
Replacement with new equipment:
If the repair within the warranty period lasts more than 45 days from the reporting of the error, the provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Period of availability of spare parts and servicing after warranty period:
7 years
Free software and updates:
should be provided for the whole life time of the equipment
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
yes
no
yes
no
yes
no
yes

## Splošni pogoji
SPLOŠNI POGOJI
Življenjska doba:
najmanj 10 let
Garancijski rok:
1 leto, vključno z vsemi zahtevanimi storitvami, nadomestnimi deli, stroški dela in potnimi stroški
Lokacija inštalacije:
Na začasno lokacijo MBP, Fornače 41, 6330 Piran
Prevoz opreme do mesta inštalacije:
mora biti vključen
Inštalacija v predviden prostor:
mora biti vključena
Pakiranje opreme:
mora biti vključeno
Osnovni potrošni material za zagon in inštalacijo opreme:
mora biti vključen
Primopredaja in prevzem opreme:
po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti validacijo/testiranje na lokaciji naročnika, tako da je po inštalaciji analizator pripravljen za pričetek uporabe v vseh zahtevanih aplikacijah.
Delovna navodila za uporabo opreme:
morajo biti predana naročniku ob dostavi opreme.
Usposabljanje kadra naročnika za delo z opremo:
Usposabljanje kadra na lokaciji naročnika:
najmanj 1 dan za usposabljanje o delovanju opreme.
Oddaljena pomoč za nastavitev poskusov in odpravljanje težav v času garancijske dobe brezplačno
Servis:
Mora biti zagotovljen s strani ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme.
Ponudnik mora zagotoviti pooblaščen servis iz strani serviserjev na mestu inštalacije peči oz. na lokaciji naročnika.
Servis se izvaja s strani ustrezno usposobljenih serviserjev.
V času garancijske dobe mora biti odzivni čas servisiranja največ 3 delovne dni, čas za popravilo opreme pa največ 14 dni od prijave napake. Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok.
Če je potreben transport opreme na servis, stroški transporta do servisa bremenijo dobavitelja v času garancijske dobe.
Zamenjava z novo opremo:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev.
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe:
7 let.
Brezplačna programska oprema in njene posodobitve
morajo biti zagotovljene tekom celotne življenjske dobe opreme.
DA
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
NE
DA
NE
DA
NE
DA
NE
DA