## Technical specifications
Zebrafish aquartium set
|  |  | The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
1. The system should be designed and manufactured for the purpose of keeping zebrafish healthy and reproducing the most ideal water conditions required for the fish.
2. The system should be suitable for the use of various tanks; for example 1.1, 2.4, 3.5 and 8 L tanks separately or altogether according to the need of the user without requiring any change at rack level. Each single level should host 3,5L and 8,0L tanks simultaneously or separately without requiring any change at rack level. Altogether, the system should be consist of 5 standard rows and one additional row dedicated for small tanks (1.1 L or similar)
3. The system must allow the use of a larger tank of 16L in place of 3,5 or 8,0 L tank. This tank must be interchangeable without any modification at rack level. The tank must be used for the long- term housing of adult fish and must be equipped, when needed, with a dedicated sloped breeding insert to promote fish breeding. The system should have gravel bottom for environmental enrichment, movable spawning platform with sloped gradient and embryo tray and water distribution with perforated pipe from the bottom.
4. The system must have dedicated plastic double inlet valves to feed each tank separately (regardless of the used tank volumes) with filtered and disinfected system water. Each valve includes a flow regulation valve (rear) and an inlet water on/off nozzle (front). These valves guarantee the fine tuning of the inlet flow in each tank separately. The on/off nozzle also has a clear visual indicator when the water flow in order to warn about the risk of static water. The water inlet flow valves should not include any kind of hose or plastic pipes.
5. Tanks of 3,5 and 8,0L should provide a continuous tank siphoning system from the bottom of the tank by means of a true plastic siphon.
6. The system must have the technical sump (volume of 90L) in the lower part of the rack. It includes the following points # 7, 8, 9, 10, 11, 12, 13, 14.
7. The technical sump must not transfer any vibrations to the rack structure.
8. The system must have chemical filtration as an option, upgradable at any time in form of activated carbon or equivalent, that must be provided in a pelleted solution already included.
9. The system must include a fine mechanical filtration. The filtration must be performed using a self- cleaning technology (integrated drum filter with a 26µm mesh), which will prevent the use of pre-filters/fine mechanical filters.
10. The system must have a biological filter able to keep Nitrogen Compounds at safe levels using floating plastic bio-chips, that must be included into offer.
11. The system must be able to automatically replace the 10% of the total system volume on a daily basis. Water needs to be gradually and automatically replaced over a 24h period.
12. The system must include UV disinfection providing a minimum radiance power at 5 WCH (Water Changes per Hour) >150,000 µWs/cm2 disinfection rate at the end of bulbs life (9.000 hours).
13. UV lamp vessel must be made of 316L s/s stainless steel.
14. The system should maintain the water temperature at the value needed between 25-29º C independently from the ambient temperature and it should have robust, stainless steel or titanium, not glass heating elements.
15. Water should run at a minimum flow of 5 WCH (Water Changes per Hour) at tank level.
16. The system must be provided with a soft touch screen with a dedicated designed software to perform system settings and track the water quality parameters on site in real time.
17. The system must include automatic dosing pumps and buffering solutions tanks to buffer water quality.
18. The system needs to constantly monitor the main water quality parameters (Temperature, Conductivity and pH), keeping them as constant as possible at the pre-set value. The user needs to be able to set optimal points easily and through a soft-touch screen.
19. Possibility to download data regarding water quality on a USB memory stick from the main panel and remotely send alarm messages as e-mail.
20. Monitoring system should allow for TDG (Total Dissolved Gas) probe integration as an option.
21. The rack should be manufactured from AISI 316 L stainless steel covered with “moisture and water resistant” dedicated epoxy coating or equivalent.
22. Rack structure needs to be equipped with a clear identifier for easy tank identification.
23. Each tank should sit, while on the rack, on dedicated printed enriched runners for promoting fish welfare and proper orientation within the tank.
24. The tank holders positioned on racks should be drawer type and the user should be able to use both hands to access inside of tanks after pulling them out.
25. The tanks bodies should be manufactured from light blue transparent polycarbonate material or equivalent, in order to reduce the risk of algae growth.
26. System tanks should have automatic water drain. The water coming from the tanks should convey automatically to the filter unit by means of the plastic siphon present at the rear side of the tank. The plastic siphon should be made of two pipes: one vertical pipe which acts as an overflow and the bend pipe acting as the true siphon. The siphoning action and interruption of flow is guaranteed by a silicon spring.
27. Each tank should have two different areas, one where the animals swim and one where the siphon is active. The two compartments need to be separated by a dedicated baffle.
28. System tanks should have light blue transparent polycarbonate lids or equivalent. The lids should not hinder the incoming of filtered water.
29. The tanks should be equipped with removable 300, 500 and 800µm net baffles to prevent the escape of young fish into the drain circuit.
30. The system should feature the possibility to be upgraded with an automatic feeding system at a later stage. The automatic feeding system must be able to provide 4 dry diet as well as a liquid live diet.
31. The system must include an embedded solution to allow the use of a dedicated environmental detector able to collect debris and promote the growth of bio-film to be able as alternative to the standard use of sentinel fish.
32. Dedicated breeding boxes of 0,7 – 1,0 – 1,7L should be provided and included. They should be made of polycarobate material, injection molded without any use of metal grids. Breeding tanks with sloped bottom are preferred as they improve the breeding behavior of zebrafish.
33. The system must include cone shaped vessel of 5L volume with lid for breeding of Artemia (fish food) with aerator, air diffuser and heating element.
34. The system must include mass breeding embryo system volume of more than 10 L with spawning chamber with corner effect. The system should have removable fish divider and allow raisable spawning platform to mimic shallow water, to use in static or recirculation mode and to have bottom valve with strainer to collect embryos.
35. The whole system should be installed and delivered in operating condition, delivering extended training.
36. Additional labor hours and materials must also be budgeted for the installation of the unit; at least 3 meters of hose (for inflow and outflow), related connectors and other indirectly connected material needed for safe and appropriate operation.
yes
yes
yes
yes
yes
no
no
no
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
yes
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no

## General conditions
GENERAL CONDITIONS
The life time of the equipment:
minimum 10 years with regular maintenance according to the manufacturer's instructions
Warranty period:
2 years, including all required services, spare parts, labour costs and travelling costs
Installation location:
at the NIB location, Večna pot 111, Ljubljana
Transportation of equipment to the installation location:
should be included
Installation in the foreseen space:
should be included
Packing of the equipment:
should be included
Basic consumables for installation and commissioning of the equipment:
should be provided
Commissioning and hand over of the equipment:
after the completed training. Before commissionin and hand over the equipment must be tested for operation. Prior to delivery,
the selection of test parameters to demonstrate individual requirements in the technical specifications shall be agreed.
Instructions for use:
should be delivered upon delivery of the equipment
Training:
On-line or on-site (depending on epidemic situation) training should be provided for unlimited number of attendees:
at least 1 days for software
at least 1 days for the system operation
Remote support for experimental setup and troubleshooting during warranty period:
free of charge
Service:
Service must be provided by the selected tenderer during warranty period and life time of the equipment.
During warranty period the service response time must be 2 working days and the instrument repair time 4 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the contracting authority. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
At least one qualified service technician must be available in EU.
Replacement with new equipment:
If the repair within the warranty period lasts more than 45 days from the reporting of the error, the provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Certificates:
The equipment must be CE certified (EU declaration of conformity).
Period of availability of spare parts and servicing after warranty period:
at least 8 years
Free software and updates:
should be provided for the whole life time of the equipment
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
yes
no
yes
no
yes
no
yes

## Splošni pogoji
SPLOŠNI POGOJI
Življenjska doba:
najmanj 10 let ob rednem vzdrževanju po navodilih proizvajalca
Garancijski rok:
2 leti: v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in potni stroški serviserja
Lokacija inštalacije:
NIB, Večna pot 111, Ljubljana
Prevoz opreme do mesta inštalacije:
mora biti vključen
Inštalacija v predviden prostor:
mora biti vključen
Pakiranje opreme:
mora biti vključen
Osnovni potrošni material za zagon in inštalacijo opreme:
mora biti vključen
Primopredaja in prevzem opreme:
po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti preizkus delovanja opreme. Pred dobavo se dogovori
o izboru parametrov preizkusa za dokazovanje posameznih zahtev v tehničnih specifikacijah.
Delovna navodila za uporabo opreme:
morajo biti predana naročniku ob dostavi opreme
Usposabljanje kadra za delo z opremo
Usposabljanje kadra naročnika za delo z opremo na lokaciji naročnika ali preko video konference (odvisno od situacije epidemije) mora biti zagotovljeno za neomejeno število udeležencev:
najmanj 1 dan za usposabljanje za uporabo programske opreme,
najmanj 1 dan za usposabljanje o teoretičnih osnovah delovanje sistema
Brezplačna oddaljena pomoč za nastavitev poskusov in odpravljanje težav v času garancijske dobe.
Servis:
Servis mora biti zagotovljen s strani izbranega ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme.
V času garancijske dobe mora biti odzivni čas servisiranja največ 2 delovna dni, čas za popravilo opreme pa največ 4 delovne dni od prijave napake, če je možno popravilo na daljavo. Če je potrebno popravilo na lokaciji naročnika, mora biti čas za popravilo opreme največ 14 dni oziroma po dogovoru z naročnikom. Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok.
V državi članici Evropske unije mora biti na voljo vsaj en usposobljen serviser.
Zamenjava z novo opremo:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev.
Certifikati:
Oprema mora imeti CE certifikat (izjava EU o skladnosti).
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe:
vsaj 8 let
Brezplačna programska oprema in njene posodobitve:
morajo biti zagotovljene tekom celotne življenjske dobe opreme
da
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
da
da
da
da
da
da
da
da
da
da
ne
da
ne
da
ne
da
ne
da
da
da
da
da
da