## Technical Specifications
Colony counting system; must be independent and self-sufficient system composed of hardware and software ready to use
| The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
· Colony counting system must be independent and self-sufficient system ready to use |  |  | f
· HD colour camera
· HD lenses
· Zoom x28 or greater
· camera must produce high quality video imaging with resolution 1.2 megapixels, or comparable quality
· Lighting technology: long-life white LEDs / Dark Field
· Lighting system: automatic 6 combinations, top and/or bottom light, white or black background
· counting should be fast, repeatable and with high accuracy
· high speed analysis: up to 1000 colonies per second
· Minimum size of colony: 0.05 mm
· Inhibition zone measurement accuracy: ± 0.3 mm
· Counting Petri dishes Ø 55 - 90 mm
· Counting on pour plating
· Counting on surface plating
· it must include features to help distinguish between touching (separation) and overlapping colonies
· Counting: automatic with manual control
· user friendly and intuitive to easy to use and allowing to build a step-by-step colony counting sequence, allowing operator to set the optimal counting of the colony
· Inhibition zone reader
· Automatic separation of clustered colonies
· Export of data to printed report, PDF, jpg, png, bmp, Excel™
· Measurement of individual objects
· Count items by size classification
· Licence to register at least four individuals
· produced data/results should be transferable to Excel
· Colour detection: 7 colours on the same dish + 1 colour to exclude
· Creation of polygonal exclusion zones
· Manual control to add or subtract colonies
· Inhibition zone reading on round dishes Ø 90 mm
· Inhibition zones measurement: automatic with manual control
· Number of antibiotic paper disks: up to 7 antibiotics on a Ø 90 mm Petri dish
· Reading time: 7 inhibition zone reading between 1 to 3 seconds or faster
USB connection
· Data security: modified data traceability in conformity with 21 CFR part 11
· Results/traceability: image / sample number & antibiotic name / bacteria name / read diameter / result with category
· In compliance with: 21 CFR part 11, ISO 7218, AOAC 977.27, EUCAST, CLSI and CA-SFM
· user manual and IQ documentation should be included
yes
yes
yes
yes
yes
no
no
no
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
yes
no
yes
yes
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no

## General Conditions-EN
GENERAL CONDITIONS
Warranty period:
3 years, including all required services, spare parts, labour costs and travelling costs
Installation location:
Installation is performed by the supplier at the NIB location, Večna pot 111, Ljubljana.
Transportation of equipment to the installation location:
should be included
Installation in the foreseen space:
should be included
Packing of the equipment:
should be included
Basic consumables for installation and commissioning of the equipment:
should be provided
Commissioning and hand over of the equipment:
after onsite training, site acceptance test (SAT) should consist of declared parameter verification
Instructions for use:
should be delivered upon delivery of the equipment
Training:
On-site training should be provided for at least 5 attendees:
at least 1 days for machine operation or
On-line or on-site (depending on epidemic situation) training should be provided for unlimited number of attendees:
at least 1 days for software
at least 1 days for theoretical basis of measurements or
Remote support for experimental setup and troubleshooting during warranty period:
free of charge
Service:
Service must be provided by the tenderer during warranty period and life time of the equipment.
During warranty period the service response time must be 2 working days and the instrument repair time 4 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the customer.
Replacement with new equipment:
If the repair within the warranty period lasts more than 45 days from the reporting of the error, the tenderer undertakes to replace the device or its part with the same or adequate new one within 3 months.
Certificates:
The equipment must be CE certified.
Period of availability of spare parts and servicing after warranty period:
At least 10 years
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
yes
no
yes
no
yes

## General Conditions-SLO
SPLOŠNI POGOJI
Garancijski rok:
3 leta: v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in potni stroški serviserja
Lokacija inštalacije:
NIB, Večna pot 111, Ljubljana
Prevoz opreme do mesta inštalacije:
mora biti vključen
Inštalacija v predviden prostor:
mora biti vključena
Pakiranje opreme:
mora biti vključeno
Osnovni potrošni material za zagon in inštalacijo opreme:
mora biti vključen
Primopredaja in prevzem opreme:
po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti test sprejemljivosti in polne funkcionalnosti
opreme (SAT), ki mora vsebovati preverbo v tehničnih specifikacijah navedenih parametrov.
Delovna navodila za uporabo opreme:
morajo biti predana naročniku ob dostavi opreme
Usposabljanje kadra za delo z opremo
Usposabljanje kadra naročnika za delo z opremo na lokaciji naročnika za vsaj 5 udeležencev:
najmanj 1 dni za usposabljanje o delovanju opreme ali
Usposabljanje kadra naročnika za delo z opremo na lokaciji naročnika ali preko video konference (odvisno od situacije epidemije) mora biti zagotovljeno za neomejeno število udeležencev
najmanj 1 dni za usposabljanje za uporabo programske opreme in
najmanj 1 dni za usposabljanje o teoretičnih osnovah meritev ali
Brezplačna oddaljena pomoč za nastavitev poskusov in odpravljanje težav v času garancijske dobe - brezplačno
Servis:
Servis mora biti zagotovljen s strani ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme.
V času garancijske dobe mora biti odzivni čas servisiranja največ 2 delovna dni, čas za popravilo opreme pa največ 4 delovne dni od prijave napake, če je možno popravilo na daljavo. Če je potrebno Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok.
Zamenjava z novo opremo:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev.
Certificates:
Oprema mora imeti CE certifikat.
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe:
vsaj 10 let
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
ne
da
da
da
da
da
da
da
da
da
da
da
da
da
ne
da
ne
da
ne
da
da
ne