## Technical Specifications
Elementar analyzer CHNSO
|  | The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
Carrier gass:
Helium as primary carrier gas
Carrier gass consumtion is optimized through various operating modes
Possibility of upgrading from He to Ar carrier gass as optional
Elementar analyzer CHNSO should enable:
Analises of samples of weight range from 0.05 to 50 mg
Build in a rotation autosampler minimum of 30 sample places
Detection limit of respective elements from 0.01 to 100%
Analyses of all analytes, including O, without major setup changes, preferably simultaneously.
Elementar analyzer CHNSO should enable:
Stable and controllable oxidation/reduction oven temperature above 1000 ᵒC.
Programatically regulated furnace temperature
Programmatically regulated- automatic leak test of measuring and reference line
Separation of combustion product on GC Column followed by detection using Supersensitive Thermal Conductivity detector
Software (SW) requirements:
Provided by a SW that meets ISO 9001 standard requirements or equivalent
Fully automatic SW control of all parameters and results reporting by user selected format. Set up guide for latest Windows version
Real time chromatogram display.
Storing an unlimited number of methods.
Ability to reprocess individual runs and batches of samples.
AutoRun & Auto Calculation option provided.
AutoReady & Wake up routine.
Export at least in Excell or ASCII selected format and/or others equivalent.
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
no
yes
no
yes
yes
yes
no
yes
no
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
yes
no
no

## General Conditions
GENERAL CONDITIONS
The life time of the equipment:
minimum 10 years
Warranty period:
2 years, including all required services, spare parts, labour costs and travelling costs
Warranty period for furnace:
10 years
Warranty period TCD (Thermal conductivity detector ) detector:
10 years
Installation location:
NIB, Marine Biology Station Piran, Fornače 41, 6330 Piran, 1. floor (CHNS laboratory)
Transportation of equipment to the installation location:
should be included
Installation in the foreseen space:
should be included
Packing of the equipment:
should be included
Basic consumables for installation and commissioning of the equipment:
should be provided
Commissioning and hand over of the equipment:
after onsite training, implementation of basic equipment operation diagnostics and equipment validation
Instructions for use:
should be delivered upon delivery of the equipment (including for computer control)
Training:
On-site training should be provided for at least 5 attendees:
at least 1 day for machine operation
On-line or on-site (depending on epidemic situation) training should be provided for unlimited number of attendees:
at least 2 days for software
at least 3 days for theoretical basis of measurements
Remote support for experimental setup and troubleshooting during warranty period:
free of charge
Service:
Service must be provided by the provider during warranty period and life time of the equipment.
Provider must provide an authorized official repairer in the Republic of Slovenia.
Service is performed by suitably qualified service technicians.
During warranty period the service response time must be 2 working days and the instrument repair time 4 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the customer. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
When it is necessary to transport the equipment to the service, the transport costs are covered by the supplier within the warranty period.
Replacement with new equipment:
If the repair within the warranty period lasts more than 45 days from the reporting of the error, the provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Period of availability of spare parts and servicing after warranty period:
8 years
Free software and updates:
should be provided for the whole life time of the equipment
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
yes
no
yes
no
yes
yes
no
no
yes
no
yes

## Splošni pogoji
SPLOŠNI POGOJI
Življenjska doba:
najmanj 10 let
Garancijski rok:
2 leti, v tem času morajo biti v garancijo vključeni vsi potrebni servisi, rezervni deli, stroški dela serviserja in potni stroški serviserja
Garancijski rok za peč:
10 let
Garancijski rok za detektor:
10 let
Lokacija inštalacije:
NIB, Morska Biološka postaja Piran, Fornače 41, 6330 Piran, 1. nadstropje (CHNS laboratorij)
Prevoz opreme do mesta inštalacije:
mora biti vključen
Inštalacija v predviden prostor:
mora biti vključena
Pakiranje opreme:
mora biti vključeno
Osnovni potrošni material za zagon in inštalacijo opreme:
mora biti vključen
Primopredaja in prevzem opreme:
po izvedenem šolanju na lokaciji naročnika. Pred primopredajo in prevzemom je potrebno izvesti osnovne diagnostike delovanja opreme in validacijo opreme.
Delovna navodila za uporabo opreme:
morajo biti predana naročniku ob dostavi opreme (vključno za računalniško krmiljenje).
Usposabljanje kadra naročnika za delo z opremo:
Na lokaciji naročnika za najmanj 5 udeležencev:
najmanj 1 dan za usposabljanje o delovanju opreme.
Na lokaciji naročnika ali preko video konference (odvisno od situacije epidemije):
najmanj 2 dni za usposabljanje za uporabo programske opreme
najmanj 3 dni za usposabljanje o teoretičnih osnovah meritev
Oddaljena pomoč za nastavitev poskusov in odpravljanje težav v času garancijske dobe
brezplačno
Servis:
Mora biti zagotovljen s strani ponudnika v času garancijske dobe ter v celotni življenjski dobi opreme.
Ponudnik mora zagotavljati pooblaščenega uradnega serviserja v Republiki Sloveniji.
Servis se izvaja s strani ustrezno usposobljenih serviserjev.
V času garancijske dobe mora biti odzivni čas servisiranja največ 2 delovna dni, čas za popravilo opreme pa največ 4 delovne dni od prijave napake, če je možno popravilo na daljavo. Če je potrebno popravilo na lokaciji naročnika, mora biti čas za popravilo opreme največ 14 dni oziroma po dogovoru z naročnikom. Za vsako odpravo napak, daljšo od 14 dni, se ustrezno podaljša garancijski rok.
Če je potreben transport opreme na servis, stroški transporta do servisa bremenijo dobavitelja v času garancijske dobe.
Zamenjava z novo opremo:
V kolikor popravilo v garancijskem roku traja več kot 45 dni od prijave napake, se izbrani ponudnik zaveže opremo ali njen ustrezni del zamenjati z enako novo opremo ali delom v roku 3 mesecev.
Doba zagotavljanja rezervnih delov in servisiranja po izteku garancijske dobe:
8 let.
Brezplačna programska oprema in njene posodobitve
morajo biti zagotovljene tekom celotne življenjske dobe opreme.
DA
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
NE
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
DA
NE
DA
NE
DA
NE
DA
NE
DA
DA
NE
NE
DA
NE
DA