## Technical Specifications
Semi-automated plant phenotyping system
[5 empty] | The tender includes: | The title of the technical documentation: | Page number in the tehnical documentation:
Semi-automated phenotyping system should enable:
Manual placement of at least 5 trays for plants in individual pots on the platform simultaneously
Automated transportation of plants between imagining modules by conveyor
Sequential imaging of plants with multiple imaging modules
Configurable multi-modal imaging
Throughput: Depending on experimental protocol, from at least 4 trays per hour (dark adaptation, fluorescence imaging, 3D imaging) to more than 20 trays per hour (RGB top view imaging)
The phenotyping system should enable phenotyping of:
Plants with height minimal up to 70 cm (including pot)
Plants with diameter minimal up to 40 cm
Plants planted in individual pots with the volumes approximately:
5 l
2 l
250 ml
Tray weight up to 6 kg
Trays labelled with QR codes or barcode sticker for automatic identification and subsequent analysis
The phenotyping system should fit in available space and supplies:
Available room: 2,9 m x 3,3 m (height: 2,85 m)
Entrance to the room through 2 doors
Environmental parameters in the room: temperature range: 15-35 °C, relative humidity: 40-80 %, no condensing environment
Power supply: 3 phase 400VAC/16A
Disinfection requirements
Disinfection of the phenotyping system against trips and mites should be possible (e.g. 70% alcohol and 10% bleach)
Disinfection of the trays against viruses and bacteria should be possible (e.g. autoclaving)
The tenderer should specify the possibilities for above mentioned disinfection requirements
Semi-automated phenotyping system should consists of:
Imaging Cabinet
Durable construction
Light-insulated
Automated doors
Number of trays provided:
150 trays with lids for approximately 5 l pots
100 trays with lids for approximately 2 l pots
50 trays with lids for approximately 250 ml pots
Number of plant holders suitable for offerd phenotyping system and compatible with offered lids:
100 for approximately 5 l pots
200 for approximately 2 l pots
QR code or barcode reader
QR code or barcode printer
Measuring device for light intensity and spectrum
Light insulated box enabling dark and light adaptation
Sensor modules
RGB Imaging Module
At least 2 cameras (one for top view and one for side view)
Cameras using next generation CMOS chip with at least 12 Megapixel resolution
Spatial resolution: minimum 25 pixels per mm2
Bit depth 8 or 12 Bit
Interface: GigE
LED light source for each camera
Turning platform for 360° image acquisition
RGB Imaging Module should enable:
Measurements of morphometric parameters:
Area
Perimeter
Eccentricity
Roundness
Plant Height
Plant Width
Colour segmentation for plant fitness evaluation
Sequential computed relative leaf growth rates
Comparison of leaf area differences during whole experiment
Greening index
Contribution of individual colours after colour segmentation and characterization
Time course analyses (e.g. for growth rates)
3D Laser Scanning Module
3D Laser Scanning Unit should generate accurate 3D model with at least 1 mm resolution in all directions
Automatic analysis and conversion of point cloud to triangle mesh
Light source: 660 nm laser beam
3D Laser Scanning Module should enable:
Automatic segmentation on individual leaves and stem
Possibility to co-register with RGB texture or parametric images from other imaging modules
Automatic calculations of morphological parameters including
Digital biomass
Leaf/Plant area
Projected area
Leaf angle
Compactness
Plant/Stem height
Number of leaves
Multi-colour Fluorescence imaging including Chlorophyll Fluorescence Imaging module
Chlorophyll Fluorescence Imaging part of the Module
Technology: Pulse amplitude modulation (PAM)
Camera with at least 1,4 MP resolution
Minimum pixel size 6.4 µm x 6.4 µm
Dynamic range: 16 bit or higher
Interface: GigE
Multiple light sources: Red-orange 620 nm LED-based light, Far-Red 735 nm LED-based light, and Cool-white 5700K LED-based light
Light Homogeneity ±5% over imaged area at least 35x35 cm
Saturation pulse intensity at least 6000 µmol/m²s
Actinic Light intensity at least 1500 µmol/m²s
Maximum intensity of the measuring light 0,05 μmol/m2s
Measured parameters at least: Fo, Fm, Fv, Fo', Fm', Fv', Ft
Calculated parameters at least: Fv/Fm, Fv'/Fm', PhiPSII , NPQ, qN, qP, Rfd
Chlorophyll Fluorescence Imaging Module should enable:
Monitoring of abiotic and biotic stresses that affect photosynthetic mechanisms (drought, heat, and light stresses, nutrient deficiencies, heavy metal toxicities, pathogen infections...)
Monitoring of the degree of stress, and recovery from stress, are quantification of stress over time
Multi-colour Fluorescence imaging part of the Module
Additional light sources for multi-colour fluorescence at least: UV, Royal Blue, Blue, Green, Amber, Deep Red.
Camera filter wheel: automatic, at least 7 positions
Multi-colour Module should enable:
Detection of fluorescent proteins and dyes (GFP, YFP, RFP, SYBR Green…)
Multi-colour fluorescence analysis
Measurements of chlorophyll fluorescence excitation spectra
Quantification of pigments (flavanols, anthocyanins)
Hyperspectral SWIR (Short Wavelength Infrared) Imaging Module
Spectral range at least: 900-1700 nm
Spectral resolution at least: 2 nm FWHM
Light source covering whole spectral range
Spatial bands: at least 500
Spectral bands: at least 600
Hyperspectral SWIR Imaging Module should enable:
Quantification of water content in plants
Quantification cellulose and lignin in plants
Quantification of sugar level
Definition of new indices by user
Thermal LWIR (Long Wavelength Infrared) Imaging Module
Detector: uncooled microbolometer FPA
Spectral range at least: 8 to 14 μm
Camera with at least 640x480 pixels resolution
Measurement accuracy: ±2°C or ±2%
Dynamic range: 16 bit
Temperature sensitivity (NETD @30°C): at least 30 mK
Temperature reference sensors
Programmable LED light source
Thermal Imaging Module should enable:
Light-induced stress response
Stomatal conductance research
Hardware and software
Software Package for client application
Software Package for server service
Software Package for image analysis and processing
Software Package for data management and handling of big data
Possibility to access data remotely via VPN client
Software has to enable local reanalysis with different parameters without overwriting original data
Programmable logic controller unit with sensors
Storage capacity: minimum 6 TB, mirrored 1:1
Connection to cloud or external storage with automatic upload and download
UPS backup
Internal diagnostic module with GSM or email alerts
Password protected accounts with different access rights - at least 3 levels (including administrator and basic user)
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
no
yes
no
yes
no

## General Conditions
GENERAL CONDITIONS
The life time of the equipment:
minimum 10 years
Warranty period:
2 years, including all required services, spare parts, labour costs and travelling costs
Installation location:
NIB, Večna pot 111, Ljubljana, 1st floor (2nd level)
Transportation of equipment to the installation location:
should be included
Installation in the foreseen space:
should be included
Packing of the equipment:
should be included
Basic consumables for installation and commissioning of the equipment:
should be provided
Commissioning and hand over of the equipment:
after onsite training, site acceptance test (SAT) should consist of declared parameter verification
Instructions for use:
should be delivered upon delivery of the equipment
Training:
On-site training should be provided for at least 5 attendees:
at least 1 day for machine operation
On-line or on-site (depending on epidemic situation) training should be provided for unlimited number of attendees:
at least 2 days for software
at least 3 days for theoretical basis of measurements
Remote support for experimental setup and troubleshooting during warranty period:
free of charge
Service:
Service must be provided by the provider during warranty period and life time of the equipment.
During warranty period the service response time must be 2 working days and the instrument repair time 4 working days from the error report if remote repair is possible and 14 days if on-site repair is needed or in agreement with the customer. For any repair of a defect longer than 14 days, the warranty period is extended accordingly.
When it is necessary to transport the equipment to the service, the transport costs are covered by the supplier within the warranty period.
Replacement with new equipment:
If the repair within the warranty period lasts more than 45 days from the reporting of the error, the provider undertakes to replace the device or its part with the same or adequate new one within 3 months.
Personnel:
The tenderer should have employed at least one qualified service technician.
Period of availability of spare parts and servicing after warranty period:
8 years
Free software and updates:
should be provided for the whole life time of the equipment
yes
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
no
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
yes
no
no
yes
no
yes