Naziv projekta: Nadgradnja nacionalnih raziskovalnih infrastruktur – RIUM

Project title: Upgrading National Research Infrastructures – RIUM

Tehnične specifikacije opreme za oplemenitenje in preoblikovanje polprevodniških, fotonskih in drugih rezin.

Technical specifications of equipment for refining and reshaping semiconductor, photonic and other slices.

PREDMET JAVNEGA NAROČILA

- SUBJECT OF THE PUBLIC CONTRACT

| Predmet javnega naročila je nabava visoko tehnološke opreme, ki bo omogočila raziskave in razvoj naprednih MEMS, MOEMS, polprevodniških struktur in silicijeve fotonike. Javno naročilo je razdeljeno na 2 (dva) sklopa, in sicer: Sklop 1: Sistem za fotolitografijo Sklop 2: Žaga za razrez rezin Vsak posamezni sklop opreme lahko dostavi različni ponudnik, pri čemer lahko ponudniki ponudbe predložijo za en sklop ali za oba sklopa Ponudniki morajo podati ponudbo, ki vključuje vse postavke opreme, ki je zahtevana v posameznem sklopu. Naročnik ne bo obravnaval ponudb, ki ne ponujajo vseh postavk opreme posameznega sklopa in bo takšne ponudbe zavrnil kot nedopustne. Lokacija dobave opreme za sklop 1: Univerza v Mariboru, Fakulteta za elektrotehniko, računalništvo in informatiko, Koroška cesta 46, 2000 Maribor, Slovenija, Laboratorij za elektro-optične in senzorske sisteme objekt G, čisti prostor G-105. Izbrani ponudnik bo moral opremo dobaviti na zgoraj naveden naslov. Lokacija dobave in vnos opreme za sklop 2: Univerza v Mariboru, Fakulteta za elektrotehniko, računalništvo in informatiko, Koroška cesta 46, 2000 Maribor, Slovenija, Laboratorij za elektro-optične in senzorske sisteme objekt G, čisti prostor G-105. Izbrani ponudnik bo moral opremo vnesti, namestiti, montirati, instalirati, zagnati in testirati. Naročnik bo ponudnikom za sklop 2 omogočil ogled prostora na lokaciji vnosa in namestitve opreme, skladno s točko 11 Navodil ponudnikom. Rok dobave: Kot velja za posamezni sklop. Uposabljanje: Kot velja za posamezni sklop. Garancijski rok: Kot velja za posamezni sklop. Za vso opremo znotraj posameznega sklopa velja, da mora zanjo v času garancijskega obdobja proizvajalec opreme zagotavljati dostop do vseh posodobitev programske opreme brez dodatnih stroškov. Vsi popravki, servisni paketi ter nadgradnje in podpora morajo biti od proizvajalca strojne oziroma programske opreme posameznega sklopa. V času garancijskega obdobja mora ponudnik za vso opremo iz posameznega sklopa nuditi tehnično podporo. Tehnična podpora mora vključevati svetovanje pri nastavitvi parametrov, konfiguracijo sistema iz posameznega sklopa in pomoč pri izvedbi posameznih procesov. Tehnična podpora se lahko izvaja preko oddaljenega dostopa do sistemov iz posameznega sklopa. Ponudnik mora za opremo iz posameznega sklopa zagotoviti: - da so vsi ponujeni moduli/komponente združljivi in medsebojno povezljivi, - da je mogoče vse zahtevane module/komponente povezati v enoten in brezhibno delujoč sistem v skladu s tehničnimi specifikacijami, - odpraviti vse napake, ki bi se pojavile zaradi morebitne medsebojne nezdružljivosti ali nepovezljivosti oz. zaradi neustrezne povezave v enoten, brezhibno delujoč sistem. V ponudbi za posamezen sklop mora biti zajeti ves droben material, ki je potreben za priključitev naprav, tudi če ni izrecno naveden. | The subject of the procurement is a high-tech equipment that will enable the research and development of advanced MEMS, MOEMS, semiconductor structures and silicon photonic devices. The public procurement is divided into two (2) lots , namely: Lot 1: Photolithographic system Lot 2: Dicing saw Each lot of the procurement can be provided by different supplier wherein supplier can provide quotation for individual lot, or both lots. Suppliers must provide quotation which includes all demands of the equipment required in individual part. Client will not consider tenders which does not offer all the items of equipment of individual lot. Such tenders will be rejected as inadmissible. Equipment delivery for lot 1: University of Maribor, Faculty of Electrical Engineering and Computer Science, Laboratory for Electro-Optics and Sensor Systems, Koroška cesta 46, 2000 Maribor, Slovenia, Building G, clean room G-105. The successful tenderer will have to send the equipment on above address. Equipment delivery and placement for lot 2: University of Maribor, Faculty of Electrical Engineering and Computer Science, Laboratory for Electro-Optics and Sensor Systems, Koroška cesta 46, 2000 Maribor, Slovenia, Building G, clean room G-105. The successful tenderer will have to deliver, place, install, assembly, start-up and test the equipment. The contracting authority will enable the tenderers for lot 2 to inspect the clean room on site in line with Point 11 of the Instructions to tenderers. Delivery time: As applicable to the individual lot. Training: As applicable to the individual lot. The warranty period: As applicable to the individual lot. The equipment manufacturer for each lot is obliged to provide access to all software updates at no additional cost during warranty period. All patches, service packs, upgrades and support must be provided by the hardware or software manufacturer of each lot. During the warranty period the supplier for each lot must provide technical support. Technical support must include consulting on setting parameters, use of the machine, configuring the system, and assisting in the implementation of individual processes, etc. Technical support can be provided through remote access for each lot. The provider of the equipment from each lot must guarantee: - that all modules/components offered are compatible and interconnected, - that all required modules/components can be integrated into a single and seamlessly functioning system in line with technical specifications, - to eliminate all errors resulting from possible incompatibility or mismatch, or from improper connection to a single, seamlessly functioning system. The offer of each lot must include all the fine material, necessary for connecting the devices, even if it is not explicitly stated |
| --- | --- |

TEHNIČNE ZAHTEVE

TECHNICAL SPECIFICATIONS

| V spodnjih tabelah so za vsak sklop navedene karakteristike opreme, ki jih zahteva naročnik. Ponudba mora izpolnjevati ali presegati vse minimalne tehnične zahteve. | Tables bellow list the characteristics of equipment for each lot, as required by the customer. The tender must meet or exceed all minimum technical specifications |
| --- | --- |
| Sklop 1: Sistem za fotolitografijo Univerza v Mariboru naroča sistem za fotolitografijo, ki vključuje opremo za nanos tankih plasti smole/laka na različne obilike in velikosti vzorcev, opremo za razvijanje ter opremo za kontrolirano strjevanje/sušenje vzorcev. Sistem naj v grobem sestavljajo programabilna grelna plošča ter dva hitro vrtljiva programirljiva modula; eden za nanos tanke plasti in drugi za namene razvijanja. Obe napravi za nanos tankih plasti naj imasta v osnovi hitro vrtljivi modul za vzorce različnih dimenzih in oblik do premera 150 mm ter naj bo dobavljena skupaj z vakuumskimi držali za standardne dimenzije rezin in laboratorijskih stekelc. Poleg tega naj vsebuje vse potrebne priključke za samostojno delovanje, vključno z vakuumsko črpalko za oba hitro vrtljiva modula. Grelna plošča naj bo opremljena s po višini nastavljivimi držali za točno pozicioniranje vzorcev in naj omogoča segrevanje po naprej določenih receptih do temperature vsaj 230 °C s homogeno porazdelitvijo temperature po celotni temperaturni plošči. Sistem naj bo združljiv z električnim standardom - 230 VAC/50Hz. V nadaljevanju so podane minimalne tehnične zahteve. Ponudba mora izpolnjevati ali presegati vse minimalne tehnične zahteve. | Lot 1: Photolitographic system The University of Maribor is seeking a photolithographic system, which includes equipment for applying thin layers of resin to samples of various dimensions and sizes, equipment for the development process, and equipment for controlled heating/drying of samples. The system should roughly consist of a programmable hotplate and two spin coaters. One spin coater is for applying a thin layer, and another is for the development process. Both spin coaters should basically have a rotating module for samples of various dimensions and shapes up to a diameter of 150 mm and should be supplied together with vacuum holders for standard dimensions of wafers and laboratory slides. In addition, it should contain all the necessary connections for independent operation, including a vacuum pump for each spin coater. The hotplate should be equipped with height-adjustable pins for accurate positioning of samples and should allow heating according to predetermined recipes to a temperature of at least 230 °C with homogeneous temperature distribution over the entire temperature plate. The system should be compatible with the electrical standard - 230 VAC / 50Hz. Minimum technical requirements are given below. The offer must match or exceeded all minimum technical requirements. |
| Št. postavke; Število kosov No. of pieces; ZAHTEVANO; REQUIREMENTS. 1; Sistem za fotolitografijo: - Hitro vrtljivi modul za nanos tankih plasti Maksimalne dimenzije (ŠxDxV) - 300 x 455 x 300 mm Vsebovati mora: - Modul za rotirajoče/centrifugalno nanašanje primerno za do 150 mm velike okrogle rezine ali do 100 mm kvadratne substrate - Vakuumsko prijemalo za standardne rezine do 150 mm - Vakuumsko prijemalo za vzorce majhnih premerov od 12 mm in več - Krmilnik z grafičnim vmesnikom in barvnim zaslonom na dotik - Robusten prozoren pokrov za posodo - Delovno posodo - Zaščita pred odprtjem pokrova - hitrost: programsko nastavljiva v področju:od najmanj 5 rpm do vsaj 10000 rpm (+/- 1 rpm); Photolitographic system: - Spin coater for applying a thin layer Maximum dimensions (WxDxH): - 300 x 455 x 300 mm Must include: - Spinner module for wafers up to 150 mm round or up to 100 mm square substrates - Vacuum chuck for standard wafers up to 150 mm - Vacuum chuck for pieces of small dimensions for sample size diameter 12 mm or larger - Controlling system with GUI and colour touchscreen - Rugged, transparent lid hinge - Process bowl - Cover open interlock - Speed: programable speed setting from at least 5 rpm to at least 10000 rpm (+/- 1 rpm). - pospešek: ≥ 25000 rpm/s - nastavitev časa vrtenja: Programsko od vsaj 1s do najmanj 9999 s. - Možnost shranjevanja vsaj 150 receptov - Možnost priključitve zunanjega USB pogona na krmilnik - Priključek za odvod izpustov - Vključena brezoljna vakuumska črpalka - Dodatno držalo brizg za centralno dovajanje s vsaj tremi priključki - Standardni nastavki za dozirnike; - Acceleration: ≥ 25000 rpm/s - Spinning time: Programmable from at least 1s to at least 9999 s - Program storage for recepies: At least 150 recipes - External connection of USB drive in the controller - Drain connection - Vacuum Pump (Oilfree) - Central syringe holder with at least triple connectors - Standard syringe holders. - Hitro vrtljivi modul za razvijanje Maksimalne dimenzije (ŠxDxV) - 300 x 455 x 300 mm Vsebovati mora: - Modul za rotirajoče/centrifugalno nanašanje primerno za do 150 mm velike okrogle rezine ali do 100 mm kvadratne substrate - Vakuumsko prijemalo za standardne rezine do 150 mm - Vakuumsko prijemalo za vzorce majhnih premerov od 12 mm in več - Krmilnik z grafičnim vmesnikom in barvnim zaslonom na dotik - Robusten prozoren pokrov za posodo - Delovno posodo - Zaščita pred odprtjem pokrova - hitrost: programsko nastavljiva v področju:od najmanj 5 rpm do vsaj 10000 rpm (+/- 1 rpm) - pospešek: ≥ 25000 rpm/s - nastavitev časa vrtenja: Programsko od vsaj 1s do najmanj 9999 s. - Možnost shranjevanja vsaj 150 receptov - Možnost priključitve zunanjega USB pogona na krmilnik - Priključek za odvod izpustov - Vključena vakuumska črpalka (brez olja) - Dodatno držalo brizg za centralno dovajanje s vsaj tremi priključki; - Spin coater for development Maximum dimensions (WxDxH): - 300 x 455 x 300 mm Must include: - Spinner module for wafers up to 150 mm round or up to 100 mm square substrates - Vacuum chuck for standard wafers up to 150 mm - Vacuum chuck for pieces of small dimensions for sample size diameter 12 mm or larger - Controlling system with GUI and colour touchscreen - Rugged, transparent lid hinge - Process bowl - Cover open interlock - Speed: programable speed setting from at least 5 rpm to at least 10000 rpm (+/- 1 rpm) - Acceleration: ≥ 25000 rpm/s - Spinning time: Programmable from at least 1s to at least 9999 s - Program storage for recepies: At least 150 recipes - External connection of USB drive in the controller - Drain connection - Vacuum Pump (Oilfree) - Central syringe holder with at least triple connectors. - Grelna plošča Maksimalne dimenzije (ŠxDxV): - 320x450x200 mm Vsebovati mora: - Namizno grelno ploščo za okrogle vzorce do premera minimalno 150 mm - Vležajen pokrov - Pine za nastavljanje vzorcev po višini - Distančniki za vzorce - Vakumska držala - Temperaturno področje: 50 ≥ T ≥ 230 °C - Temperaturno odstopanje na celotni plošči: - Maksimalno +/- 0.5 °C - Število receptov ≥ 10 - 230 VAC napajanje z EU priključkom Rok dobave - Največ 110 dni od pričetka veljavnosti pogodbe. Garancija - 1 letna garancija od dneva podpisa prevzemnega zapisnika Usposabljanje: - Namestitev opreme in zagon; 2.) Hot plate Maximum dimensions (WxDxH): - 320x450x200 mm Must include: - Table top hot plate for wafers up to 150 mm minimum - Hinged lid - Lifting pins for substrate handling - Proximity pins - Vacuum bake holders - Temperature range: 50 ≥ T ≥ 230 °C - Temperature uniformity: - Max.: +/- 0.5 °C - Number of recipes: ≥ 10 - 230 VAC power supply with EU power plug Delivery time - Maximum 110 days from the day of entry into force of the contract. Warranty - 1-year warranty from the day of signature of the handover record Training: - Equipment installation and commissioning. |
| Sklop 2: Žaga za razrez rezin Univerza v Mariboru naroča zmogljivo žago za razrez rezin, ki bo optimizirana za več kotno rezanje tankih rezin premera vsaj 6” . Žaga mora biti čim bolj vsestranska in mora omogočati rezanje silicijevih (polprevodniških) rezin, silicijevega stekla, keramičnih materialov in silicijevega karbida. Sistem mora vsebovati uporabniku prijazen grafični vmesnik za nastavljanje ter kontrolo naprave. Vključen mora biti sistem za povečavo, ki omogoča optimalno pozicioniranje rezila na vzorcu. Rezila mora biti preprosto zamenjati, vključen mora biti sistem za obnovo rezil. Žaga za razrez rezil mora vsebovati vse komponente, podsisteme ter orodja za takojšnje rezanje rezin po namestitvi. V nadaljevanju so podane minimalne tehnične zahteve. Ponudba mora izpolnjevati ali presegati vse minimalne tehnične zahteve. | Lot 2: Dicing saw University of Maribor is seeking for a versatile system for wafer dicing which will allow for angle dicing of wafers with at least 6”. Dicing saw must be versatile and should allow for dicing of silicon wafers, silica glass, ceramic materials and silicon carbide. System should include user friendly interface that allows for easy control and programming of the device. Included must be vision system that allows for rough and fine positioning of the blade on workpeace. Blades must be easily replaceable and dressing unit shall be included for blade maintenance. Dicing saw shall include all components, sub systems and tools for immediate cutting of dies after installation. Minimum technical requirements are given below. The offer must match or exceed all minimum technical requirements |  |
| Št. postavke; Število kosov No. of pieces; ZAHTEVANO; REQUIREMENTS. 2; 1; Avtomatska žaga za razrez polprevodniških rezin, z naslednjimi značilnostmi: Vreteno - 2 kW vreteno z ≤40000 rpm ali 1.8 kW vreteno z ≤60000 rpm - Vodno hlajeno vreteno Rezilo - velikost rezila 2’’ (dva inča) - nosilec za rezila s prirobnicami - Nosilec za rezila brez prirobnic - Zunanji premer 52 mm - Sistem za brušenje/vzdrževanje (»Dress«) rezila na delovnem primežu - Brezkontaktna nastavitev konice rezila za preverjanje položaja konice in obrabe med avtomatskim delovanjem Vrtljiva miza - Omogoča rotacije ≥320° - Porozna keramična vpenjalna miza primerna za 6" rezine Rezanje - Rezalno področje po x osi: ≥ 155 mm - Največja hitrost premika po X-osi: ≥ 750 mm/s - Največja hitrost vračanja po X-osi: ≥ 750 mm/s - Doseg Y-osi: ≥ 160 mm - Največja hitrost po Y-osi: ≥ 180 mm/s - Ločljivost Y osi: ≤ 5 µm po celotnem območju - Premik po Z osi: ≥ 30 mm - Največja hitrost pomika Z-osi: ≥ 80 mm/s - ponovljivost Z osi: ≤1 µm Sistem vida - Mikroskop z nizko povečavo - Optična povečava ≤ x4 ali povečava na zaslonu ≤ 40x - Avtomatski fokus - Mikroskop z veliko povečavo - Optična povečava ≥ x7 ali povečava na zaslonu ≥ 210x - Avtomatski fokus - Osvetlitev med poravnavo ter pozicioniranjem Velikost sistema - Višina (s signalno lučjo): ≤ 215 cm - Širina: ≤ 50 cm - Globina: ≤ 90 cm Uporabniški vmesnik in nadzorna programska oprema - Grafični uporabniški vmesnik z zaslonom na dotik - Zaslon: ≥ 15“ - Programska oprema za kontrolo in programiranje naprave - Navodila za obnovitev v primeru napake - Hramba več kot 3000 parametrov rezanja (receptov) - Statusni indikatorji naprave - Funkcija za preverjanje vreza - Prikaz položaja rezila med rezanjem - Avtomatsko delovanje - Avtomatsko zaznavanje položaja rezanja - Pol-avtomatsko delovanje - Hitrost in globina se lahko nastavita za vsako posamezno linijo - USB vmesnik za zunanji pomnilnik - Uporabniški vmesnik v angleškem jeziku Zahteve za napajanje, zrak in vodo - Napajanje - 380 VAC, 3-faze - Zrak - Najmanjši tlak ≥ 0.5 MPa - Največji tlak ≤ 0.8 MPa - Povprečna poraba zraka ≤ 180 l/min - Voda za rezanje - Najmanjši tlak ≥ 0.2 MPa - Največji tlak ≤ 0.5 MPa - Povprečna poraba vode ≤ 4 l/min - Voda za hlajenje - Najmanjši tlak ≥ 0.2 MPa - Največji tlak ≤0.5 MPa - pretok vode ≤ 2 l/min Dobavni rok - Največ 110 dni od dneva pričetka veljavnosti pogodbe. Usposabljanje - Instalacija z najmanj 1x 8 urnim usposabljanjem za najmanj 1 (eno) osebo Garancija - 1 letna garancija od dneva podpisa prevzemnega zapisnika (ne velja za začetni potrošni material in orodja). Začetni potrošni material in orodja: Postavka; Količina. Komplet orodja za vzdrževanje naprave; 1. Orodje za menjavo rezila; 1. Rezilo z držalom primerno za silicij; 3. Rezilo brez držala primerno za steklo; 3. Rezilo z držalom primerno za silicijev karbid; 2. Rezilo z držalom primerno za keramične materiale; 2. nosilec za rezila brez prijemal Φ52; 1. Orodje za R tip nosilca; 1. Ploščice za vzdrževanje rezila; 10. 6” ogrodje za rezalno mizo; 5. Vsaj 20 m rola traku za pritrditev obdelovanca; 1.; Automatic dicing saw system minimum technical requirements: Spindle - 2 kW spindle with ≤40000 rpm or 1.8 kW spindle with ≤60000 rpm - Water-cooled spindle Blade - 2’’ (two inch) blade size - Flange for hub blade - Flange for hubbless blade - Outer diameter of 52 mm - System for the Dress of the blade on process chuck table - Non contact blade tip setup to check blade tip position and wear amount during automatic operation Turn table - Degree of rotation ≥320° - Porous ceramic chuck table suitable for 6" wafers Cutting performance - Cutting range on X-axis: ≥ 155 mm - Max X-axis moving (operating) speed: ≥ 750 mm/s - Max X-axis return speed: ≥ 750 mm/s - Y-axis stroke: ≥ 160 mm - Max Y-axis feed speed: ≥ 180 mm/s - Y axis accuracy ≤ 5 µm over max range - Z axis stroke: ≥ 30 mm - Z-axis max feed speed: ≥ 80 mm/s - Z axis repeatability: ≤1 µm Vision System - Microscope with low magnification - Optical magnification ≤ x4 or magnification on monitor ≤ 40x - Auto focus - Microscope with high magnification - Optical magnification ≥ x7 or magnification on monitor ≥ 210x - Auto focus - Light illumination during alignment and positioning System size - Height (including status indicator): ≤ 215 cm - Width: ≤ 50 cm - Depth: ≤ 90 cm User interface and control software - Graphic user interface with touch panel - Screen: ≥ 15” - Software for control and programing of the machine - Recovery guidance in case of an error - Storage of more than 3000 cutting data (recepies) - Machine status indicators - Kerf Check Function - Display of blade position during cutting - Automatic operation - Automatic determination of cutting positions - Semi automatic operation - Feed speed, and height can be set manually for each individual cutting line - USB input for external memory - User interface in English language Power, air, and water requirements - Power - 380 VAC, 3-phase - Air - Min pressure ≥ 0.5 MPa - Max pressure ≤ 0.8 MPa - Average air consumption ≤ 180 l/min - Cutting water - Min pressure ≥ 0.2 MPa - Max pressure ≤ 0.5 MPa - Average flow rate ≤ 4 l/min - Cooling water - Min pressure ≥ 0.2 MPa - Max pressure ≤0.5 MPa - Flow rate ≤ 2 l/min Delivery time - Maximum 110 days from the day of entry into force of the contract. Training - Installation with minimum 1x 8 hours training for at least 1 (one) person included Waranty - 1-year warranty from the day of signature of the handover record (does not apply to initial consumables and tools). Initial consumables and tools: Item; Quantity. Maintenance tool kit; 1. Tools for blade replacement; 1. Hub blade suitable for silicon; 3. Hubless blade suitable for glass; 3. Hub blade suitable for silicon carbide; 2. Hub blade suitable for ceramic materials; 2. R-Type flange for hubbless blade Φ52; 1. Tooling for R-Type flange; 1. Dresser boards; 10. 6" frames for chuck table; 5. At least 20 m tape roll for workpiece fixation; 1.. |

Operacijo sofinancirata Republika Slovenija, Ministrstvo za izobraževanje, znanost in šport in Evropska unija iz Evropskega sklada za regionalni razvoj. Operacija se izvaja v okviru Operativnega programa za izvajanje evropske kohezijske politike 2014-2020 v okviru prednostne osi 1: »Mednarodna konkurenčnost raziskav, inovacij in tehnološkega razvoja v skladu s pametno specializacijo za večjo konkurenčnost in ozelenitev gospodarstva«, prednostne naložbe 1.1 »Krepitev infrastrukture za raziskave in inovacije ter zmogljivosti za razvoj odličnosti na tem področju, pa tudi spodbujanje pristojnih centrov, zlasti takšnih, ki so evropskega pomena«, specifičnega cilja 1.1.1 »Učinkovita uporaba raziskovalne infrastrukture ter razvoj znanja/kompetenc za boljše nacionalno in mednarodno sodelovanje v trikotniku znanja«.

The operation is co-financed by the the Republic of Slovenia, Ministry of Education, Science and Sport and the European Union from the European Regional Development Fund. The operation is carried out under priority axis 1: ''International competitiveness of research, innovation and technological development, in line with smart specialization for increased competitiveness and greening of the economy'', priority investment 1.1 ''Strengthening R&D infrastructure and capacity to develop excellence in this and the promotion of competent centers, especially those of European importance'', specific objective 1.1.1 ''Efficient use of research infrastructure and development of knowledge / competences for better national and international cooperation in the knowledge triangle'' under the Operational Program for implementation of the European cohesion policy 2014-2020.