SPECIFICATIONS

| Contracting Authority | Univerza v Ljubljani, Naravoslovnotehniška fakulteta Aškerčeva cesta 12 1000 Ljubljana who performs the public procurement in his own name and in the name and on behalf of Inštitut za kovinske materiale in tehnologije, IMT Lepi pot 11 1000 Ljubljana |
| --- | --- |
| Public contract reference number | JNOP-03/24-OMM |
| Subject matter of the public contract | Sistem za termomehanske simulacije in preizkušanje materialov – System for Thermomechanical Simulations and Testing of Materials |

TYPE, CHARACTERISTICS, QUALITY AND APPEARANCE OF THE PUBLIC CONTRACT SUBJECT/TENDER

Sample Dimensions

- Uniaxial tensile and compression specimen sizes between 5mm and 20mm diameter

- round or with a cross sectional area of up to 400 mm2 for square or rectangular samples.

- Plane Strain compression testing specimen sizes of up to 10mm x 15mm x 20mm with a

- 5mm plane strain section

Vacuum and Inert Gas Testing

- Vacuum tank surrounding test area, with the possibility of testing under high vacuum or testing in an inert gas environment by evacuating the tank and backfilling with non-flammable gases such as argon, nitrogen, or helium.

- System must include vacuum equipment capable of providing vacuum pressures of from

- atmospheric pressure to 5 x 10–2 mm Hg. Optional high vacuum pressures available from 5 x 10–2 mm Hg to 10–5 mm Hg at an extra cost.

- Ability to quench with air, inert gas or water inside the vacuum tank.

- Electrical and mechanical ports through the vacuum tank for thermocouples, strain transducers, quench media, etc. Transparent viewing window in the vacuum tank to allow viewing of the sample from multiple angles during testing under vacuum conditions.

Thermal System

- ±1 °C resolution of thermal control system.

- Direct resistance heating (based on the Joule effect).

- Capacity to melt and perform controlled solidification after melting.

- Temperature range between ambient temperature and 1700 °C (Based on choice of thermocouples. Optional equipment available to operate to 3000 °C)

- Heating rates of zero (steady state) to 10,000°C/sec (for specific sample types) and controlled cooling rates between 10 and 330 °C/sec at 1000 °C without applying external quenching media. Quench cooling rates of up to 10,000 °C/sec (surface temperature for specific sample types).

- Thermocouple temperature control with the option to utilize optical pyrometer for thermal control.

- Ability to produce isothermal planes through sample thickness direction.

- Ability to quench in-situ without removing the heating system.

- Ability to monitor a minimum of 4 thermal channels simultaneously. Ability to utilize multiple thermocouple types such as type K, S, R, B, or E thermocouples.

- Electrical power of at least 70 KVA.

Mechanical System

- Closed-loop servo hydraulic control.

- Horizontal test frame to minimize temperature variations caused by “chimney effect”.

- Single 60gpm valve or equivalent system to accommodate both fast and slow stroke rates of between 0.01 and 1000 mm/sec.

- Temperature controlled hydraulic fluid system to avoid errors caused by hydraulic fluid temperature changes.

- Static compression force capability of up to 10 metric tons. Static tensile force capability of up to 10 metric tons. Dynamic compressive force at a ram speed of 500 mm/sec of at least 5 metric tons.

- Stroke measurement resolution of 0.002 mm.

- Force measurement resolution of 1 kg.

- Force measurement accuracy of ±1 % of full scale.

- Ability to measure strain either lengthwise or crosswise selectively in the hot zone of the sample.

- Ability to control the system based on; stroke, crosswise strain, lengthwise strain, force, true stress in hot zone, engineering strain, true strain.

- Frame stiffness of at least 7.0 x 10–10 meters/Newton.

Additional equipment

- System must include a portable thermocouple welder.

- The system must include all required computer equipment necessary to run the tests and simulations as well as a computer system and software capable of analyzing the data and making appropriate graphs, charts, etc.

- Supplier must offer a wide selection of grips / anvils, etc. for tension, compression and plane strain compression conditions on different sample types and sizes.

- System must include provisions for mechanical or optical dilatometry & extensometry measurements.

- System must allow for the creation of transformation curves (i.e. CCT curves).

- System must contain provisions for the accurate measurement of both very high forces (consistent with the maximum capacity of the machine) as well as very low forces suitable for low strength tensile testing or the determination of nil-strength temperatures.

- Ability to internally quench samples using water, air, or inert gas media while maintaining external sample vacuum.

- Ability to expand the machine capabilities in the future by additional modules and upgrade conversions, for example: Torsion, Rolling Mill Simulation, Ultra Fine Grain Refinement, Laser Ultrasonic microstructure monitoring.

Computer Control and Data Acquisition

- Computer control of all operations through included software.

- Ability to analyze test data on the system while additional tests are being run.

- Automatic acquisition of data.

- User selectable data acquisition speed from zero to 50 KHz.

- Compatibility with Ethernet / Internet.

I, the tenderer’s representative/authorised signatory hereby declare that the offered goods/services fully comply with the above description.

In     		 , on

Name and surname: