SPECIFICATIONS

| Contracting Authority | Univerza v Ljubljani, Naravoslovnotehniška fakulteta Aškerčeva cesta 12 1000 Ljubljana who performs the public procurement in his own name and in the name and on behalf of Znanstvenoraziskovalni center Slovenske akademije znanosti in umetnosti Novi trg 2 1000 Ljubljana |
| --- | --- |
| Public contract reference number | JNOP-02/24-OG |
| Subject matter of the public contract | Nakup in dobava rentgenskega praškovnega difraktometra (XRD) The purchase of the multipurpose x-ray powder diffractometer (XRD) |

TYPE, CHARACTERISTICS, QUALITY AND APPEARANCE OF THE PUBLIC CONTRACT SUBJECT/TENDER

1. System requirements

Multipurpose X-Ray powder diffractometer for standard powder measurements of regular and bulky samples in reflection and transmission mode, microdiffraction, Grazing Incidence X-Ray Diffraction (GIXRD) and Small Angle X-ray Scattering (SAXS).

2. System enclosure should enable easy access to operator and should contain all required electronics, cables, accessories and supporting systems. The enclosure should allow use Silver (Ag) radiation and fulfil all required X-ray safety regulations.

3. High Voltage Generator

Power: 3 kW or more

High voltage: up to 60 kV or more

Current: up to 60 mA or more

4. Goniometer

Vertical Theta/Theta goniometer, with mechanism for fast exchange of all optical modules and sample stages without need for any re-alignment after exchange.

Angular range at least: -110<2Theta< 164 degrees,

The smallest step: 0.0001degree or less

5. X-ray Source

Long Fine Focus (LFF) sealed ceramic X-ray tubes with Cu, anode.  X-Ray tube and tube housing allowing easy change from line to point focus and vice versa without any realignment, supplies re-connection or special tools.

6. Rotational Sample Stage for reflection and transmission

Rotational sample stage (sample “spinner”) with programmable rotation speed and programmable angular position, ready for operation in reflection and foil transmission mode, suitable for following sample types: standard reflection (“back” or “front” loading); Si single crystal holders (zero background) for small amount of sample and holders suitable for loading small irregularly shaped solid samples. The sample stage should be exchangeable with other sample stages without the need for any alignment

7. Automatic Sample Changer

Automatic sample changer allowing automatic loading of min. 45 samples in one single automatic batch with sample holders, including min. 45 sample holders for reflection geometry, sample holders for samples on substrates (min. 12 pieces) and substrates (min.25 pieces) for filtering of clay samples and 3 sample holders for the samples of irregular shape.

8. Sample platform for irregular and bulky samples

Sample platform with 3 motorized and programmable axis (sample height and 2 sample translation axis). The sample stage should allow loading of samples with dimensions at least 8x8x4 cm and up to 2 kg weight. The range of programmable axis should be:

Translation (X,Y) min. 50 mm,

Height (Z) min. 10 mm)

The stage should be equipped with camera or camera/laser system for easy determination of measurement spot.

The sample stage should be exchangeable with other sample stages without the need for any alignment.

9. Primary monochromator

Primary monochromator for measurements in Bragg-Brentano geometry, to improve “peak to background” ratio and remove CuKbeta line.

10. Incident divergence Slits

Programmable slit system allowing measurements in constant divergence angle (at least from 1/32O to 2O) and constant irradiated length (at least from 1-20mm).

The module should be exchangeable with other modules without the need for any alignment.

11. Small Spot module (Microdiffraction)

Separate module or accessories for other modules enabling measurements in “small spot” mode. It should enable at least 3 different spot sizes in the range from 0.1 to 0.5 mm. The module should be exchangeable with other modules without the need for any alignment

12. Small Angle Scattering (SAXS) module

Separate module or accessories for other modules enabling measurements in SAXS mode.  The module should be exchangeable with other modules without the need for any alignment.

13. Grazing incidence diffraction (GI XRD) module

Separate module or accessories for other modules enabling measurements in GI XRD mode. The module should be exchangeable with other modules without the need for any alignment

14. Detector Slits

Programmable slit system allowing measurements in constant divergence angle (at least from 1/32O to 2O) and constant irradiated length (at least from 1-20 mm). The module should be used in a combination with fast linear (1D) detector.

The module should be exchangeable with other modules without the need for any alignment

15. Other diffracted beam modules and accessories

The complete sets of modules/accessories required for Microdiffraction, SAXS and GI XRD to be used in combination with above listed incident beam modules and sample stages and below listed detector or detectors. The system should include at least 2 sets of Soller slits with different divergence angles and all accessories enabling both high resolution and high intensity mode.

All modules should be exchangeable with other modules without the need for any alignment

16. Solid state energy resolution fast detector

Scanning and static measurements in “point detector” mode (“0D”) and linear mode (“1D”)

The channel (strip) size: 90 micron or less

Number of channels (strips): 100 or more

Energy resolution for Cu Kalpha radiation: less than 400 eV

17. System accessories

System should include all necessary support systems like X-ray tube cooling device, compressed air supply, personal computer, monitor, etc.

18. Instrument control and experiment setup software

Software module for operation of the system, data collection strategy preparation, automatic processing, and data conversion.

19. Analytical software

Analytical functionality required below can be combined into one or more software packages or modules. License for use at least on 3 independent PCs should be included for each offered software package/module.

- General data treatment: General graphics data manipulation, profile fitting peak search, data import/export, smoothing, alpha2 stripping, etc.

- Qualitative and semi-quantitative analysis: Search/Match (phase identification) and semi-quantitative analysis module.

- Crystallographic analysis: Unit cell search and refinement and full profile refinement (Le Bail and/or Pawley).

- Rietveld analysis: Rietveld analysis (quantitative and structural).

- Data Clustering: Cluster analysis of high number of scans without reference databases.

20. ICDD PDF 4/Minerals database

I, the tenderer’s representative/authorised signatory hereby declare that the offered goods/services fully comply with the above description.

In     		 , on

Name and surname: