NABAVA NAPRAVE: SISTEM ZA TEKOČINSKO KROMATOGRAFIJO Z VISOKO LOČLJIVOSTJO - HPLC

Tehnične specifikacije in kriteriji za nabavo:

| Zaporedna številka | Zahtevane strokovne in tehnične lastnosti za HPLC | Opis je v katalogu na strani |
| --- | --- | --- |
| Splošne zahteve |  |  |
| 1. | Uporabniku prijazen visokozmogljiv sistem za tekočinsko kromatografijo. |  |
| 2. | Omogočati mora delovanje do 700 bar. |  |
| 2 | Delovati mora v pH območju med 1-12 ali širše. |  |
| 3 | Vsebovati mora senzorje za puščanje tekočine. |  |
| 4 | Sistem mora biti opremljen s programom, ki omogoča upravljanje instrumenta, obdelavo podatkov in kreiranje poročil. |  |
| 5 | Sistem mora biti kontroliran s sistemskim kontrolerjem z barvnim ekranom na dotik, ki omogoča pregled nad delovanjem sistema, prijavo v sistem, zagon in zaustavitev sistema, auto-purge, kreiranje metode ter ogled kromatograma. |  |
| Črpalka |  |  |
| 6 | Črpalka mora imeti vgrajeno razplinjevanje mobilnih faz ter omogočati mešanje štirih mobilnih faz hkrati. |  |
| 7 | Omogočati mora pretoke do 10 ml/min in s točnostjo pretoka +/- 1 %. |  |
| 8 | Preciznost pretoka mora biti vsaj ≤0.1% RSD ali +/- 0.03 min SD ali boljše. |  |
| Avtomatski vzorčevalnik |  |  |
| 9 | Omogočati mora injiciranja med 0.1 in 50.0ul z možnostjo nadgradnje do 1000.0 ali več. |  |
| 10 | Omogočati mora termostatiranje vzorcev od 4°C do 40 °C ali več. |  |
| 11 | Točnost injiciranja mora biti vsaj ≤+/- 1 %. |  |
| 12 | Imeti mora prostora za vsaj 160 1.5/2 ml vial. |  |
| 13 | Linearnost injektorja mora biti ≥0.9999 |  |
| 14 | Carryover manj kot 0.0015% |  |
| Termostat za kolone |  |  |
| 15 | Omogočati mora 6 kolon dolžine do 100mm ali 3 kolone do 30cm. |  |
| 16 | Omogočati mora termostatiranje v območju: 10°C pod sobno temperaturo do 80°C ali več. |  |
| 17 | Točnost temperature mora biti +/-0.8°C ali boljše s stabilnostjo +/-0.3°C ali boljše |  |
| Detektor na niz diod |  |  |
| 18 | Število niza fotodiod mora biti do 1024. |  |
| 19 | Območje valovnih dolžin od 190 do 800nm |  |
| 20 | Delovati mora v ph območju 1-12 ali širše |  |
| 21 | Točnost valovne dolžine mora biti +/-1nm |  |
| 22 | Celica mora imeti 10mm optično pot ter mora zdržati pritiske do 100 barov ali več. |  |
| Programska oprema |  |  |
| 23 | Mora omogočati upravljanje z vsemi komponentami HPLC sistema in obdelavo podatkov. |  |

Druge zahteve:

|  | Ponudba mora vključevati ustrezen računalnik z monitorjem, tipkovnico in miško. |  |
| --- | --- | --- |
|  | Ponudba mora vključevati instalacijo sistema |  |
|  | Ponudba mora vključevati izobraževanje uporabnikov |  |
|  | Ponudba mora vključevati kromatografsko kolono C18 dimenzije 150x4.6 mm, 2.6um s pripadajočo predkolono in nosilcem za predkolono. |  |
|  | Ponudba mora vključevati zaščitne pokrove s filtri za mobilne faze in odpadna topila. |  |
|  | Servisna podpora in tehnična pomoč v Republiki Sloveniji |  |
|  | Garancija: vsaj2 leti. Garancijsko obdobje prične teči po dobavi opreme. ** |  |
|  | Ob predaji opreme priložena po 2 izvoda navodil za uporabo in vzdrževanje, garancijskega lista in ostale dokumentacije, v slovenskem ali angleškem jeziku. |  |
|  | Dobavni rok: najkasneje do 15. 12. 2024. |  |
|  | Usposabljanje: na srednjem nivoju: pomoč pri zagonu aparature, izvedbi prve analize, nadalje pa dostopnost za pomoč/tehnično podporo na daljavo, ko jo bomo potrebovali. Število terminov usposabljanj: vsaj 2x in sicer ob dobavi in po vpeljavi v delo z aparaturo. |  |
|  | Dobavljivost rezervnih delov, potrošnega materiala najmanj 10 let |  |

** Z dnem dostave opreme v prostore naročnika oprema preide v last naročnika. O dobavi se sestavi prevzemni zapisnik. Naročnik si pridržuje pravico, da montažo in zagon ter edukacijo uporabnikov naroči po poteku določenega časovnega obdobja po dobavi opreme. Dobavitelj bo moral na pisni poziv naročnika montirati in zagnati opremo ter izvesti edukacijo na rok, določen v pozivu.

Po izvedeni montaži, zagonu ter edukaciji se sestavi zapisnika o montaži in zagonu opreme. Garancijsko obdobje prične teči naslednji dan po izdelavi zapisnika o montaži in zagonu opreme.

DODATNA MERILA ZA TOČKOVANJE

Naročnik bo izbral ekonomsko najugodnejšo ponudbo skladno z merili za izbor. Izbran bo ponudnik, ki bo prejel najvišje število točk. Ponudnik mora ponuditi opremo, ki ustreza minimalnim tehničnim zahtevam določenimi v poglavju Tehnične specifikacije. V kolikor bo ponudnik ponudil opremo, ki presega minimalne tehnične specifikacije, bo v elementih, ki so merilo za izbor, prejel dodatne točke. Skupno število točk, ki jih prejme ocenjevana ponudba, se izračuna tako, da se za vsako posamezno merilo določi število točk, ki jih prejme ocenjevana ponudba, nato pa se točke, ki jih ponudba prejme za vsako posamično merilo, seštejejo.

Izbrana bo ponudba, ki bo zbrala največje število točk.

V primeru dveh ali več dopustnih ponudb z enakim številom točk, bo izbrana tista ponudba, ki bo prejela največ točk v okviru merila B. Če bosta še vedno dopustni dve ali več ponudb z enakim številom točk, bo izbrana prej prispela ponudba.

Merilo ekonomsko najugodnejše ponudbe je sestavljeno iz spodaj navedenih meril:

Merilo A - ponudbena cena z DDV (max 60 točk)

Ocenjuje se skupna ponudbena vrednost v EUR brez DDV iz ponudbenega predračuna. Cenovno najugodnejša ponudba – Pmin prejme največje število točk (60 točk), vse ostale prejmejo glede na najcenejšo ponudbo sorazmerno manjše število točk na podlagi prikazanega izračuna:

P = Pmin/Px*60

kjer je:

P – dodeljeno število točk.

Pmin – vrednost ponudbe, ki je najnižja od vseh in je zadovoljila minimalne tehnične zahteve.

Px – vrednost ponudbe, ki se točkuje.

60 – utežni faktor

Merilo B Možnost nadgradnje z MS detektorjem

|  | Merilo | Možnost nadgradnje | Število točk | Opis je v katalogu na strani |
| --- | --- | --- | --- | --- |
| B | Sistem omogoča nadgradnjo z MS detektorjem | DA | 5 |  |
|  |  | NE | 0 |  |

Merilo C Karakteristike črpalke

|  | Merilo | Karakteristika črpalke | Število točk | Opis je v katalogu na strani |
| --- | --- | --- | --- | --- |
| C | Vzporedno vezan dvojni bat črpalke omogoča izjemno nizko pulzacijo in stabilnost bazne linije. | DA | 5 |  |
|  |  | NE | 0 |  |

Merilo D Hitrost injiciranja

|  | Merilo | Hitrost injiciranja | Število točk | Opis je v katalogu na strani |
| --- | --- | --- | --- | --- |
| D | Avtomatski vzorčevalnik mora omogočati ultra hitro injiciranje | 8 s ali manj | 10 |  |
|  |  | 10 s ali manj | 5 |  |
|  |  | 12 s ali manj | 0 |  |

Merilo E Prenos vzorca (Carryover)

|  | Merilo | Prenos (carryover) vzorca | Število točk | Opis je v katalogu na strani |
| --- | --- | --- | --- | --- |
| E | Prenos (carryover) vzorca avtomatskega vzorčevalnika mora biti zelo nizek, saj ima visok prenos (carryover) vzorca negativne učinke na natančnost in zanesljivost analiznih rezultatov | Prenos vzorca ≤ 0.0003 % | 5 |  |
|  |  | Prenos vzorca ≥ 0.0003 % | 0 |  |

Merilo F Nadgradnja nadzornega modula za mobilno faz

|  | Merilo | Nadgradnja nadzornega modula za mobilno fazo | Število točk | Opis je v katalogu na strani |
| --- | --- | --- | --- | --- |
| F | Sistem mora omogočati nadgradnjo z nadzornim modulom za mobilno fazo, ki meri težo mobilne faze in sporoča podatke o preostanku mobilne faze preko računalnika ali prenosne naprave. | DA | 5 |  |
|  |  | NE | 0 |  |

Merilo G Termostat za kolone

|  | Merilo | Termostat za kolone | Število točk | Opis je v katalogu na strani |
| --- | --- | --- | --- | --- |
| G | Prisilna cirkulacija zraka omogoča homogeno temperaturo znotraj celotne pečice. | DA | 5 |  |
|  |  | NE | 0 |  |

Merilo H Avtodiagnostika, odprava napake in nadaljevanje analize

|  | Merilo | Avtodiagnostika, odprava napake in nadaljevanje analize | Število točk | Opis je v katalogu na strani |
| --- | --- | --- | --- | --- |
| H | Avtomatsko spremljanje pulzacije sistema, menjava mobilne faze v sistemu v primeru zaznane pulzacije in nadaljevanje analize po uspešno opravljeni napaki. | DA | 5 |  |
|  |  | NE | 0 |  |

PURCHASE OF DEVICE: HIGH-PERFORMANCE LIQUID CHROMATOGRAPHY SYSTEM - HPLC

Technical Specifications and Procurement Criteria:

| Serial Number | Required Professional and Technical Specifications for HPLC | Description in the catalog on page |
| --- | --- | --- |
| General Requirements |  |  |
| 1. | User friendly high-performance liquid chromatography system |  |
| 2. | Must operate up to 700 bar. |  |
| 2. | Must operate in a pH range between 1-12 or wider. |  |
| 3. | Must include sensors for liquid leakage. |  |
| 4. | The system must be equipped with software that enables instrument control, data processing, and report creation. |  |
| 5. | The system must be controlled by a system controller with a color touch screen that allows system monitoring, login, start and stop of the system, auto-purge, method creation, and chromatogram viewing. |  |
| Pump |  |  |
| 6. | The pump must have built-in degassing for mobile phases and allow the mixing of four mobile phases simultaneously. |  |
| 7. | Must support flow rates up to 10 ml/min with a flow accuracy of +/- 1%. |  |
| 8. | Flow precision must be at least ≤0.1% RSD or +/- 0.03 min SD or better. |  |
| Automatic Sampler |  |  |
| 9. | Must support injections between 0.1 and 50.0 µl, with upgradeability to 1000.0 µl or more. |  |
| 10. | Must support sample thermostatting from 4°C to 40°C or higher. |  |
| 11. | Injection accuracy must be at least ≤+/- 1%. |  |
| 12. | Must have space for at least 160 1.5/2 ml vials. |  |
| 13. | Injector linearity must be ≥0.9999. |  |
| 14. | Carryover must be less than 0.0015%. |  |
| Column Thermostat |  |  |
| 15. | Must support 6 columns of up to 100 mm length or 3 columns of up to 30 cm length. |  |
| 16. | Must support thermostatting in the range: 10°C below room temperature to 80°C or higher. |  |
| 17. | Temperature accuracy must be +/-0.8°C or better, with stability of +/-0.3°C or better. |  |
| Diode Array Detector |  |  |
| 18. | Number of photodiodes must be up to 1024. |  |
| 19. | Wavelength range from 190 to 800 nm. |  |
| 20. | Must operate in a pH range of 1-12 or wider. |  |
| 21. | Wavelength accuracy must be +/-1 nm. |  |
| 22. | The cell must have a 10 mm optical path and withstand pressures up to 100 bar or more. |  |
| Software |  |  |
| 23. | Must enable control of all HPLC system components and data processing. |  |

Other requirements:

|  | The offer must include a suitable computer with monitor, keyboard and mouse. |  |
| --- | --- | --- |
|  | The tender must include the installation of the system |  |
|  | The tender must include user training |  |
|  | Training: intermediate level: assistance with apparatus setup, performing the first analysis, and further remote support/technical assistance as needed. Number of trainings: 2 times; upon delivery and some time after familiarisation with working on the apparatus. |  |
|  | The tender must include a C18 150x4,6 mm, 2,6um chromatographic column with a suitable pre-column and pre-column holder. |  |
|  | The tender must include protective covers with filters for mobile phases and waste solvents. |  |
|  | Service support and technical assistance in the Republic of Slovenia. |  |
|  | Warranty: at least 2 years. The guarantee period starts upon delivery of the equipment. ** |  |
|  | User manuals, maintenance instructions, warranty, and other documentation are provided upon delivery of equipment, 2 copy of each and in Slovenia nor English. |  |
|  | Delivery deadline: no later than 15th of December 2024. |  |
|  | Availability of spare parts and consumables for at least 10 years. |  |

** Upon delivery of the equipment to the customer's premises, the equipment becomes the property of the customer. A delivery report will be prepared. The customer reserves the right to schedule the installation, commissioning, and user training for a period after the delivery of the equipment. The supplier will be required to install and commission the equipment, as well as conduct the training, within the timeframe specified in the written request from the customer.

After the installation, commissioning and training, a report on the installation and commissioning of the equipment will be prepared. The warranty period will commence the day after the report on the installation and commissioning of the equipment is completed.

ADDITIONAL CRITERIA FOR SCORING

The client will select the most economically advantageous offer in accordance with the selection criteria. The provider that receives the highest number of points will be chosen. The provider must offer equipment that meets the minimum technical requirements specified in the Technical Specifications chapter. If the provider offers equipment that exceeds the minimum technical specifications, they will receive additional points in the elements that are the selection criteria. The total number of points received by the evaluated offer is calculated by determining the number of points for each individual criterion and then summing the points received for each individual criterion.

The offer that accumulates the highest number of points will be selected.

In the case of two or more acceptable offers with the same number of points, the offer that receives the most points under Criterion B will be selected. If there are still two or more acceptable offers with the same number of points, the offer that was received earlier will be selected.

The criterion for the most economically advantageous offer consists of the following criteria:

Criterion A - offered price with VAT (max 60 points)

The total offered value in EUR with VAT from the offer estimate is evaluated. The most economically advantageous offer – Pmin receives the highest number of points (60 points), while all others receive proportionately fewer points based on the lowest offer according to the following calculation:

P = Pmin/Px*60

where:

P – allocated number of points.

Pmin – the value of the offer that is the lowest of all and has met the minimum technical requirements.

Px – the value of the offer being scored.

60 – weighting factor

Criterion B: Upgradeability with MS Detector

|  | Criterion | Upgradeability | Number of Points | Description in the catalog on page |
| --- | --- | --- | --- | --- |
| B | System allows upgrade with an MS detector | YES | 5 |  |
|  |  | NO | 0 |  |

Criterion C: Pump Characteristics

|  | Criterion | Pump Characteristic | Number of Points | Description in the catalog on page |
| --- | --- | --- | --- | --- |
| C | Parallel double piston pump enables extremely low pulsation and baseline stability. | YES | 5 |  |
|  |  | NO | 0 |  |

Criterion D: Injection Speed

|  | Criterion | Injection Speed | Number of Points | Description in the catalog on page |
| --- | --- | --- | --- | --- |
| D | Automatic sampler must allow ultra-fast injection | 8 s or less | 10 |  |
|  |  | 10 s or less | 5 |  |
|  |  | 12 s or less | 0 |  |

Criterion E: Sample Carryover

|  | Criterion | Sample Carryover | Number of Points | Description in the catalog on page |
| --- | --- | --- | --- | --- |
| E | Sample carryover of the automatic sampler must be very low, as high carryover negatively affects the accuracy and reliability of analytical results. | Carryover ≤ 0.0003% | 5 |  |
|  |  | Carryover ≥ 0.0003% | 0 |  |

Criterion F: Upgrade of Mobile Phase Control Module

|  | Criterion | Upgrade of Mobile Phase Control Module | Number of Points | Description in the catalog on page |
| --- | --- | --- | --- | --- |
| F | The system must allow an upgrade with a control module for the mobile phase, which measures the weight of the mobile phase and reports the remaining mobile phase via computer or portable device. | YES | 5 |  |
|  |  | NO | 0 |  |

Criterion G: Column Thermostat

|  | Criterion | Column Thermostat | Number of Points | Description in the catalog on page |
| --- | --- | --- | --- | --- |
| G | Forced air circulation ensures homogeneous temperature throughout the oven. | YES | 5 |  |
|  |  | NO | 0 |  |

Criterion H: Self-Diagnosis, Error Correction, and Analysis Continuation

|  | Criterion | Self-Diagnosis, Error Correction, and Analysis Continuation | Number of Points | Description in the catalog on page |
| --- | --- | --- | --- | --- |
| H | Automatic monitoring of system pulsation, changing the mobile phase in the system in case of detected pulsation, and continuation of analysis after successfully resolving the issue. | YES | 5 |  |
|  |  | NO | 0 |  |