NABAVA NAPRAVE: ANALITSKA TEHTNICA

Tehnične specifikacije in kriteriji za nabavo:

| Zaporedna številka | Zahtevane strokovne in tehnične lastnosti | Opis je v katalogu na strani |
| --- | --- | --- |
| 1. | Analitska tehtnica, ki omogoča tehtanje v širokem tehtalnem razponu z visoko stopnjo natančnosti. |  |
| 2. | Tehtanje v dveh merilnih območjih: do 52 g in do 120 g oz. v enem območju: do 120 g, pri čemer se posledično pri specifikacijah pod točko 2,4, 5 in 6 upošteva 1. navedena vrednost (strožji kriterij za celoto merilno območje). |  |
| 3. | Primerna za tehtanje majhnih količin analitskih standardov v praškasti ali tekoči obliki, direktno v mikrocentrifugirkah. |  |
| 2. | Ločljivost [d]: ± 0,01/0,1 mg |  |
| 3. | Minimalna zatehta: 1 mg ali manj |  |
| 4. | Ponovljivost: 0,02/0,1 mg ali boljše |  |
| 5. | Linearnost do 0,05/0,2 mg ali boljše |  |
| 6. | Hitrost stabilizacije: 8/2 s ali boljše |  |
| 7. | OLED zaslon |  |
| 8. | Velikost tehtalne plošče naj bo premera vsaj 91 mm |  |
| 9. | Imeti mora samodejno kalibracijo v primeru spremembe temperature ≥ 1 °C ali časovno nadzorovano interna kalibracijo vsake 4 ure |  |
| 10. | Vgrajeno naj bo avtomatsko odpiranje tehtalne celice s prostoročnim upravljanjem |  |
| 11. | Vgrajen naj bo ionizator za odstranjevanje statične elektrike |  |
| 12. | Vgrajen naj bo po višini nastavljiv notranji ščitnik |  |
| 13. | Možnost shranjevanja uporabniških nastavitev za do 10 uporabnikov |  |
| 14. | Možnost brezstičnega tariranja (preko IR senzorja) |  |

Druge zahteve:

| 1. | Skupaj s tehtnico je treba zagotoviti po en izvod tehnične dokumentacije (navodila za uporabo, navodila za vzdrževanje, garancijo in drugo dokumentacijo), namestitev tehtnice in prikaz njenega delovanja. Vsa dokumentacija je v slovenščini ali angleščini. |  |
| --- | --- | --- |
| 2. | Avtoriziran servis mora biti zagotovljen v Republiki Sloveniji, prav tako tudi dodatne storitve kalibracije. |  |
| 3. | Tehnična podpora na daljavo v primeru težav. |  |
| 4. | Analitska tehnica mora biti kalibrirana. |  |
| 5. | Servis in kalibracija v Republiki Sloveniji. |  |
| 6. | Analitska tehnica ne potrebuje biti overjena s strani proizvajalca. |  |
| 7. | Garancija: najmanj 3 leta. Garancijsko obdobje prične teči po dobavi opreme. ** |  |
| 8. | Dobavni rok: najkasneje do 15. 12. 2024 |  |

** Z dnem dostave opreme v prostore naročnika oprema preide v last naročnika. O dobavi se sestavi prevzemni zapisnik. Naročnik si pridržuje pravico, da montažo in zagon naroči po poteku določenega časovnega obdobja po dobavi opreme. Dobavitelj bo moral na pisni poziv naročnika montirati in zagnati opremo na rok, določen v pozivu.

Po izvedeni montaži in zagonu se sestavi zapisnika o montaži in zagonu opreme. Garancijsko obdobje prične teči naslednji dan po izdelavi zapisnika o montaži in zagonu opreme.

PURCHASE OF DEVICE: ANALYTICAL BALANCE

Technical Specifications and Procurement Criteria:

| Serial Number | Required Professional and Technical Features | Description in the catalog on page |
| --- | --- | --- |
| 1. | Analytical balance that combines the highest level of precision with large weighing ranges. |  |
| 2. | Weighing in two measuring ranges: up to 52 g and up to 120 g, or in one range: up to 120 g, with the consequence that, for the specifications under points 2, 4, 5 and 6, the first value indicated (the stricter criterion for the whole measuring range) is taken into account. |  |
| 3. | Suitable for weighing the small amounts of analytical standards in powder or liquid form in microcentrifuge tubes. |  |
| 2. | Resolution [d]: ± 0.01/0.1 mg. |  |
| 3. | Minimum load: 1 mg or less. |  |
| 4. | Repeatability: 0.02/0.1 mg or better. |  |
| 5. | Linearity up to 0.05/0.2 mg or better. |  |
| 6. | Stabilization speed: 8/2 s or better. |  |
| 7. | OLED display. |  |
| 8. | The size of the weighing plate should be at least 91 mm in diameter. |  |
| 9. | It must have automatic calibration in case of temperature change ≥ 1°C or time-controlled internal calibration every 4 hours. |  |
| 10. | Built-in automatic door opening of the weighing cell with hands-free operation. |  |
| 11. | Built-in ionizer for removing static electricity. |  |
| 12. | Built-in height-adjustable internal shield. |  |
| 13. | Ability to store user settings for up to 10 users. |  |
| 14. | Ability for contactless taring (via IR sensor). |  |
| 15. | Service and calibration in the Republic of Slovenia. |  |

Other requirements:

| 1. | All technical documentation, 1 copy of each (User manuals, maintenance instructions, warranty, and other documentation), installation of the weighing scale, and demonstration of its operation must be provided along with it. All documentation is in Slovenian or English. |  |
| --- | --- | --- |
| 2. | Authorized service must be provided in the Republic of Slovenia, as well as additional calibration for the weighing scale. |  |
| 3. | Remote technical support in case of issues. |  |
| 4. | Analytical balance must be calibrated. |  |
| 5. | Service and calibration in the Republic of Slovenia. |  |
| 6. | Analytical balance does not have to be verificiated by manufacturer. |  |
| 7. | Warranty: at least 3 years. The guarantee period starts upon delivery of the equipment. ** |  |
| 8. | Delivery deadline: no later than 15th of December 2024 |  |

** On the day of delivery of the equipment to the purchaser's premises, the equipment becomes the property of the purchaser. A delivery acceptance report will be prepared. The purchaser reserves the right to order the installation and commissioning after a specified period following the delivery of the equipment. The supplier will be required to install and commission the equipment upon written request from the purchaser, within the timeframe specified in the request.

After the installation and commissioning are completed, an installation and commissioning report will be prepared. The warranty period shall commence the day after the installation and commissioning report is prepared.