| Zaporedna številka | Zahtevane strokovne in tehnične lastnosti | Opis je v katalogu na strani |
| --- | --- | --- |
|  | Oprema omogoča aplikacijo terapije preko inhalacij. |  |
|  | Oprema podpira možnost ventilacije tako za otroke kot odrasle. Naročnik dopušča ponudbo opreme, ki omogoča ventilacijo tudi pri novorojenčkih. |  |
|  | Oprema mora biti prenosna z možnostjo fiksne montaže, teža ventilatorja pa brez stojala ne sme presegati 7 kg. |  |
|  | Oprema mora vsebovati premično stojalo na kolesih za fiksno montažo. |  |
|  | Oprema mora delovati tudi brez priključka na kisik in elektriko. Vsebovati mora integrirano baterijo in baterijo, ki jo je mogoče enostavno menjati med delovanjem. Ventilator mora biti sposobni delovati avtonomno vsaj 8 ur. |  |
|  | Oprema mora omogočati vse vrste mehanske ventilacije (invazivna, ne-invazivna, podpora s tlakom in volumenska podpora itd.), kot tudi terapijo s kisikom. |  |
|  | Oprema mora biti primerna tudi za uporabo na terenu. |  |
|  | Oprema mora imeti močno ohišje, ki je odporno na udarce ter padce. |  |
|  | Zaslon je primeren za zunanjo uporabo. |  |
|  | Zaslon deluje na različnih temperaturnih območjih med okvirno -15 do +50. |  |
|  | Oprema mora vsebovati senzorje za SpO2 in Co2. Ponudba naj vključuje naprstni -ščipalkasti senzor, ki je primeren za odrasle. Naročnik dopušča možnost priklopa otroškega senzorja. |  |
|  | Zaželeno je, da oprema vsebuje možnost priklopa vlažilca. |  |

NABAVA NAPRAVE: VENTILATOR ZA MEHANSKO VENTILACIJO PLJUČ PACIENTOV

Tehnične specifikacije in kriteriji za nabavo:

Druge zahteve:

| 1. | Zagotavljanje rezervnih delov, nadgradnje programske opreme in servisa za obdobje najmanj 7 let. |  |
| --- | --- | --- |
| 2. | V ponudbo je vključeno tudi enkratno usposabljanje za delo z opremo, v slovenščini ali angleščini. |  |
| 3. | Ob predaji opreme priložena po 1 izvod navodil za uporabo in vzdrževanje, garancijskega lista in ostale dokumentacije, v slovenskem ali angleškem jeziku. |  |
| 4. | Garancija: 2 leti. Garancijsko obdobje prične teči po dobavi opreme.** |  |
| 5. | Dobavni rok: najkasneje do 15. 12. 2024 |  |

** Z dnem dostave opreme v prostore naročnika oprema preide v last naročnika. O dobavi se sestavi prevzemni zapisnik. Naročnik si pridržuje pravico, da montažo in zagon ter edukacijo uporabnikov naroči po poteku določenega časovnega obdobja po dobavi opreme. Dobavitelj bo moral na pisni poziv naročnika montirati in zagnati opremo ter izvesti edukacijo na rok, določen v pozivu.

Po izvedeni montaži, zagonu ter edukaciji se sestavi zapisnika o montaži in zagonu opreme. Garancijsko obdobje prične teči naslednji dan po izdelavi zapisnika o montaži in zagonu opreme.

PURCHASE OF DEVICE: VENTILATOR FOR MECHANICAL VENTILATION OF PATIENT’S LUNGS

Technical Specifications and Procurement Criteria:

| Serial Number | Required Professional and Technical Features | Description in the catalogue on page |
| --- | --- | --- |
|  | The equipment allows for the application of therapy via inhalation. |  |
|  | The equipment supports ventilation for both children and adults. The client allows for the offer of equipment that also enables ventilation for newborns. |  |
|  | The equipment must be portable with the option for fixed mounting, and the weight of the ventilator without the stand must not exceed 7 kg. |  |
|  | The equipment must include a movable stand on wheels for fixed mounting. |  |
|  | The equipment must also operate without a connection to oxygen and electricity. It must include an integrated battery and a battery that can be easily replaced during operation. The ventilator must be capable of autonomous operation for at least 8 hours. |  |
|  | The equipment must support all types of mechanical ventilation (invasive, non-invasive, pressure support, volume support, etc.), as well as oxygen therapy. |  |
|  | The equipment must also be suitable for use in the field. |  |
|  | The equipment must have a strong casing that is resistant to impacts and falls. |  |
|  | The display is suitable for outdoor use. |  |
|  | The display operates within a temperature range of approximately -15 to +50 degrees. |  |
|  | The equipment must include sensors for SpO2 and Co2. The offer should include a fingertip clip sensor suitable for adults. The client allows for the possibility of connecting a pediatric sensor. |  |
|  | The equipment is preferred to allow for the connection of a humidifier for inhaled air. |  |

Other requirements:

| 1. | Provision of spare parts, software upgrades, and service for a minimum of 7 years. |  |
| --- | --- | --- |
| 2. | One-time training for equipment operation, in Slovenian or English is included in the offer. |  |
| 3. | User manuals, maintenance instructions, warranty, and other documentation are provided upon delivery of equipment, 1 copy of each and in Slovenia nor English. |  |
| 4. | Warranty: 2 years. The guarantee period starts upon delivery of the equipment.** |  |
| 5. | Delivery deadline: no later than 15th of December 2024 |  |

** Upon delivery of the equipment to the customer's premises, the equipment becomes the property of the customer. A delivery report will be prepared. The customer reserves the right to schedule the installation, commissioning, and user training for a period after the delivery of the equipment. The supplier will be required to install and commission the equipment, as well as conduct the training, within the timeframe specified in the written request from the customer.

After the installation, commissioning and training, a report on the installation and commissioning of the equipment will be prepared. The warranty period will commence the day after the report on the installation and commissioning of the equipment is completed.

| Zaporedna številka | Zahtevane strokovne in tehnične lastnosti | Opis je v katalogu na strani |
| --- | --- | --- |
| 1. | Oprema omogoča samodejni izračun pretoka in doze zdravila za aplikacijo. |  |
| 2. | Oprema deluje tudi brez neposrednega napajanja na električno omrežje. Naročnik zahteva ponudbo infuzijskih črpalk, ki delujejo avtonomno pri pretoku 25ml/h vsaj 8 ur. |  |
| 3. | Oprema alarmira ob okluzijah, zračnih mehurčkih v sistemu in ob končani aplikaciji zdravila (svetlobno alarmiranje in prikazovanje napak ali sporočil na zaslonu). |  |
| 4. | Oprema ima funkcijo samodejnega odzračevanja sistema za aplikacijo zdravil. |  |
| 5. | Oprema ima v sistemu shranjeno obsežno knjižico zdravil. |  |
| 6. | Oprema omogoča možnost montaže na infuzijsko stojalo. |  |
| 7. | Oprema omogoča povezavo z informacijskim sistemom ustanove. |  |
| 8. | Črpalka z baterijo ne presega skupno težo 2 kg |  |
| 9. | Oprema omogoča možnosti zaklepanja kontrol ob aplikaciji zdravil. |  |
| 10. | Oprema omogoča pretočnost vsaj od 0,1 ml/h do 1200 ml/h. |  |
| 11. | Oprema omogoča aplikacijo bolusa. Oprema ima ločeno tipko z možnostjo aktivacije te funkcije in posebno tipko za zaustavitev vseh načinov aplikacije zdravila. |  |
| 12. | Oprema vključuje varnostne mehanizme ob pojavu morebitne okluzije v sistemu. |  |
| 13. | Infuzijska črpalka z uporabniškim vmesnikom v slovenščini ali angleščini, a se v tem primeru slednja ob prihodnji nadgradnji nadgradi v slovenščino. |  |

NABAVA NAPRAVE: INFUZIJSKE ČRPALKE

Tehnične specifikacije in kriteriji za nabavo:

Druge zahteve:

| 1. | Zagotavljanje rezervnih delov, nadgradnje programske opreme in servisa za obdobje najmanj 7 let. |  |
| --- | --- | --- |
| 2. | V ponudbo je vključeno tudi enkratno usposabljanje za delo z opremo, v slovenščini ali angleščini. |  |
| 3. | Ob predaji opreme priložena po 1 izvod navodil za uporabo in vzdrževanje, garancijskega lista in ostale dokumentacije, v slovenskem ali angleškem jeziku. |  |
| 4. | Garancija: vsaj 1 leto. Garancijsko obdobje prične teči po dobavi opreme. ** |  |
| 5. | Dobavni rok: najkasneje do 15. 12. 2024 |  |

** Z dnem dostave opreme v prostore naročnika oprema preide v last naročnika. O dobavi se sestavi prevzemni zapisnik. Naročnik si pridržuje pravico, da montažo in zagon ter edukacijo uporabnikov naroči po poteku določenega časovnega obdobja po dobavi opreme. Dobavitelj bo moral na pisni poziv naročnika montirati in zagnati opremo ter izvesti edukacijo na rok, določen v pozivu.

Po izvedeni montaži, zagonu ter edukaciji se sestavi zapisnika o montaži in zagonu opreme. Garancijsko obdobje prične teči naslednji dan po izdelavi zapisnika o montaži in zagonu opreme.

PURCHASE OF DEVICE: INFUSION PUMPS

Technical Specifications and Procurement Criteria:

| Serial Number | Required Professional and Technical Features | Description in the catalog on page |
| --- | --- | --- |
| 1. | The equipment enables the automatic calculation of flow rate and drug dosage for administration. |  |
| 2. | The equipment operates without direct connection to the electrical grid. The client requires the offer of infusion pumps that can autonomously operate at a flow rate of 25 ml/h for at least 8 hours. |  |
| 3. | The equipment alarms in case of occlusions, air bubbles in the system, and upon the completion of drug administration (with visual alarms and error or message displays on the screen). |  |
| 4. | The equipment features an automatic system venting function for drug administration. |  |
| 5. | The equipment has an extensive drug library stored within the system. |  |
| 6. | The equipment allows for mounting on an infusion stand. |  |
| 7. | The equipment allows for connection to the institution's information system. |  |
| 8. | The pump, including the battery, does not exceed a total weight of 2 kg. |  |
| 9. | The equipment allows for the locking of controls during drug administration. |  |
| 10. | The equipment supports a flow rate of at least 0.1 ml/h to 1200 ml/h. |  |
| 11. | The equipment allows for bolus administration. It includes a separate button for activating this function and a dedicated button for stopping all drug administration modes. |  |
| 12. | The equipment includes safety mechanisms in case of possible occlusion in the system. |  |
| 13. | The infusion pump has a user interface in Slovenian or English. The English interface has the option for future upgrades to Slovenian. |  |

Other requirements:

| 1. | Provision of spare parts, software upgrades, and service for a minimum of 7 years. |  |
| --- | --- | --- |
| 2. | One-time training for equipment operation, in Slovenian or English is included in the offer. |  |
| 3. | User manuals, maintenance instructions, warranty, and other documentation are provided upon delivery of equipment, 1 copy of each and in Slovenia nor English. |  |
| 4. | Warranty: at least 1 year. The guarantee period starts upon delivery of the equipment. ** |  |
| 5. | Delivery deadline: no later than 15th of December 2024 |  |

** Upon delivery of the equipment to the customer's premises, the equipment becomes the property of the customer. A delivery report will be prepared. The customer reserves the right to schedule the installation, commissioning, and user training for a period after the delivery of the equipment. The supplier will be required to install and commission the equipment, as well as conduct the training, within the timeframe specified in the written request from the customer.

After the installation, commissioning and training, a report on the installation and commissioning of the equipment will be prepared. The warranty period will commence the day after the report on the installation and commissioning of the equipment is completed.

| Zaporedna številka | Zahtevane strokovne in tehnične lastnosti | Opis je v katalogu na strani |
| --- | --- | --- |
|  | Oprema omogoča samodejni izračun pretoka in doze zdravila za aplikacijo. |  |
|  | Oprema deluje tudi brez neposrednega napajanja na električno omrežje. |  |
|  | Oprema alarmira ob okluzijah in zračnih mehurčkih v sistemu (prikazan na zaslonu). |  |
|  | Oprema ima funkcijo uporabe infuzijskih sistemov ki omogočajo možnost samodejnega odzračevanja za aplikacijo zdravil . |  |
|  | Oprema ima v sistemu shranjeno knjižico najbolj pogosto uporabljenih zdravil v kliničnem okolju. |  |
|  | Oprema omogoča možnost montaže na infuzijsko stojalo. |  |
|  | Oprema omogoča povezavo z informacijskim sistemom ustanove. |  |
|  | Oprema omogoča možnosti zaklepanja kontrol ob aplikaciji zdravil. |  |
|  | Oprema ima največ 0,5 – 2% odstopanja pri mehanski natančnosti. |  |
|  | Perfuzor z uporabniškim vmesnikom v slovenščini ali angleščini, a se v tem primeru slednja ob prihodnji nadgradnji nadgradi v slovenščino. |  |

NABAVA NAPRAVE: PERFUZOR

Tehnične specifikacije in kriteriji za nabavo:

Druge zahteve:

| 1. | Zagotavljanje rezervnih delov, nadgradnje programske opreme in servisa za obdobje najmanj 7 let. |  |
| --- | --- | --- |
| 2. | V ponudbo je vključeno tudi enkratno usposabljanje za delo z opremo, v slovenščini ali angleščini. |  |
| 3. | Ob predaji opreme priložena po 1 izvod navodil za uporabo in vzdrževanje, garancijskega lista in ostale dokumentacije, v slovenskem ali angleškem jeziku. |  |
| 4. | Garancija: vsaj 1 leto. Garancijsko obdobje prične teči po dobavi opreme. ** |  |
| 5. | Dobavni rok: najkasneje do 15. 12. 2024 |  |

** Z dnem dostave opreme v prostore naročnika oprema preide v last naročnika. O dobavi se sestavi prevzemni zapisnik. Naročnik si pridržuje pravico, da montažo in zagon ter edukacijo uporabnikov naroči po poteku določenega časovnega obdobja po dobavi opreme. Dobavitelj bo moral na pisni poziv naročnika montirati in zagnati opremo ter izvesti edukacijo na rok, določen v pozivu.

Po izvedeni montaži, zagonu ter edukaciji se sestavi zapisnika o montaži in zagonu opreme. Garancijsko obdobje prične teči naslednji dan po izdelavi zapisnika o montaži in zagonu opreme.

PURCHASE OF DEVICE: SYRINGE PUMP

Technical Specifications and Procurement Criteria:

| Serial Number | Required Professional and Technical Specifications | Description in the catalog on page |
| --- | --- | --- |
|  | The equipment enables the automatic calculation of flow rate and drug dosage for administration. |  |
|  | The equipment operates without direct connection to the electrical grid. |  |
|  | The equipment alarms in case of occlusions and air bubbles in the system (displayed on the screen). |  |
|  | The equipment is compatible with the use of systems that can automatically vent for drug administration. |  |
|  | The equipment has a library of the most commonly used drugs in clinical settings stored within the system. |  |
|  | The equipment allows for mounting on an infusion stand. |  |
|  | The equipment allows for the locking of controls during drug administration. |  |
|  | The equipment has a mechanical accuracy deviation of at most 0.5 – 2%. |  |
|  | The perfusor has a user interface in Slovenian or English. The English interface has the option for future upgrades to Slovenian. |  |

Other requirements:

| 1. | Provision of spare parts, software upgrades, and service for at least 7 years. |  |
| --- | --- | --- |
| 2. | One-time training for equipment operation, in Slovenian or English is included in the offer. |  |
| 3. | User manuals, maintenance instructions, warranty, and other documentation are provided upon delivery of equipment, 1 copy of each and in Slovenia nor English. |  |
| 4. | Warranty: at least 1 year. The guarantee period starts upon delivery of the equipment. ** |  |
| 5. | Delivery deadline: no later than 15th of December 2024 |  |

** Upon delivery of the equipment to the customer's premises, the equipment becomes the property of the customer. A delivery report will be prepared. The customer reserves the right to schedule the installation, commissioning, and user training for a period after the delivery of the equipment. The supplier will be required to install and commission the equipment, as well as conduct the training, within the timeframe specified in the written request from the customer.

After the installation, commissioning and training, a report on the installation and commissioning of the equipment will be prepared. The warranty period will commence the day after the report on the installation and commissioning of the equipment is completed.