TECHNICAL SPECIFICATIONS

Package 1

Plant phenotyping system with two 3D linear scanners and multispectral sensors (RGB and NIR) mounted on a mobile metal frame:

- metal frame on wheels (height at least 2 meters; length at least 2.8 meters, width at least 0.8 m)

- scanning speed not less than 75mm/s; Full scan cycle/ Total scan cycle in less than 60 seconds

- programmable scanning intervals from 1 minute to 4 days

- working range/Operating temperature range temperature 5 – 40 °C; Operating humidity range humidity 0-90 % (non-condensing)

- sensor with the following wavelengths Red (624-634 nm) Lime green (530-540 nm) Blue (465-485 nm) Near infrared (720-750 nm) Laser (935-945 nm); internal temperature control of optical components

- must be able to evaluate plants in sunlight, artificial light and darkness

- software: automated 3D imaging and multispectral processing, machine learning using multivariate linear regression

- automated calculation of at least 20 plant parameters (e. g.Plant Height Max, Plant Height Average, Canopy Light Penetration Depth, 3D Leaf Area, Projected Leaf Area, Digital Biomass, Convex Hull Area, Convex Hull Area Coverage, Convex Hull Circumference, Convex Hull Maximum Width, Convex Hull Aspect Ratio, Voxel Volume Total, Surface Angle Average, HUE average (Color), Lightness average, Saturation average, Normalized difference vegetation index (NDVI),Normalized Pigment Chlorophyll ratio index (NPCI), Plant Senescence Reflectance Index (PSRI), Green Leaf Index (GLI)).

Package 2

Stand-alone phenotyping cabinet, i.e., growth chamber with a sensor system for detailed multispectral imaging and fluorometry:

- 12 position filter-wheel

- the system allows measurements of Pulse Amplitude Modulated (PAM) and Inductions curve (OJIP) chlorophyll fluorescence in combination with absorption of chlorophyll, multispectral, color and morphological parameters.

- all data are captured by the same multispectral camera (Resolution 12.1 Mp) and lens system

- camera should have a frame rate of 30 frames per second when binned to 2x2, resulting in a 3.1 MP resolution and resulting in the extra parameter Fi/Fm when measuring the OJIP curve

- large imaging field (e.g. 70 x 70 cm)

- software controlled focusing system able to focus from 15 cm up to 100cm from the lens

- the device allows measurements of dual green fluorescent protein reporter system (GFP, 405 nm and 470 nm) or function as a GFP and red fluorescent protein (RFP) system

- excitation sources for GFP  should at least produce 150 umol of light at 70 centimeters form the lens

- measurements can be performed under controlled/manipulated environmental conditions (light intensity and spectrum, temperature, humidity, and CO2 concentration):

  - temperature range (10 to 40 °C)

  - humidity control between 30 % and 90 % (90 % should still be reached at 26 °C)

  - light intensity (up to 1000 umol of light)

  - at least 4 spectra Blue Red White and Far red

  - CO2 control between 400 and 1000 PPM