PUBLIC-TENDER DOCUMENTATION

OPEN PROCEDURE

Article 40

(type of procedure)

»SUPPLY AND INSTALLATION OF AN INTEGRATED LOW-TEMPERATURE

HIGH-MAGNETIC-FIELD MEASUREMENT SYSTEM FOR PHYSICAL PROPERTIES«

(subject of the public procurement)

JN52/2025

(internal reference number)

August 2025

CONTENTS

- Invitations to tender

- Instructions for the bidders

- Technical documentation of the public tender

- The bid (Proforma invoice) – FORM 1

- Details about the bidder – FORM 2

General data about the bidder and the consortium – FORM 2.1

Acting with subcontractors/without subcontractors – FORM 3

The subcontractor’s authorisation to direct payments made by the contracting authority to the subcontractor(s) and consent – FORM 3.1

- Forms for establishing the capacities of a bidder

- ESPD form

- Sample contract – FORM 4

- Appendixes:

APPENDIX 1 – A statement with details about the natural and legal entities owned by the bidder

APPENDIX 2 – Acceptance certificate for public tender

The public tender titled »SUPPLY AND INSTALLATION OF AN INTEGRATED LOW-TEMPERATURE

HIGH-MAGNETIC-FIELD MEASUREMENT SYSTEM FOR PHYSICAL PROPERTIES« is co-financed by the following entities:

- »JOŽEF STEFAN« INSTITUTE,

Jamova cesta 39, 1000 Ljubljana,

represented by the Acting Director Prof. Dr. Leon Cizelj (hereinafter: IJS)

- UNIVERSITY OF LJUBLJANA, FACULTY OF MATHEMATICS AND PHYSICS

Jadranska ulica 19, 1000 Ljubljana

represented by Dean Prof. Dr. Janez Bonča (hereinafter: UL FMF)

- INSTITUTE OF MATHEMATICS, PHYSICS, AND MECHANICS (IMFM)

Jadranska ulica 19, 1000 Ljubljana

represented by Director Prof. dr. Peter Šemrl (hereinafter: IMFM)

»JOŽEF STEFAN« INSTITUTE is authorized to carry out the public procurement.

# 1.  INSTRUCTIONS FOR THE BIDDERS

# 1.1 THE CONTRACTING AUTHORITY

The contracting authority Jožef Stefan Institute, Jamova cesta 39, 1000 Ljubljana, tax number SI55560822, registration number 5051606, bank account: 01100-6030344242 registered with the Public Payment Administration (PPA) Ljubljana, has published on the public procurement portals information about public tender »SUPPLY AND INSTALLATION OF AN INTEGRATED LOW-TEMPERATURE HIGH-MAGNETIC-FIELD MEASUREMENT SYSTEM FOR PHYSICAL PROPERTIES«.

The Contracting Authority hereby invites all interested tenderers to submit a tender in compliance with the requirements from the tender documentation. The tender has to be prepared in line with the tender documentation, fulfilling all the conditions for the participation in this public tender.

# 1.2 GENERAL INFORMATION

| Subject: | SUPPLY AND INSTALLATION OF AN INTEGRATED LOW-TEMPERATURE HIGH-MAGNETIC-FIELD MEASUREMENT SYSTEM FOR PHYSICAL PROPERTIES |
| --- | --- |
| Short description: | The subject of the public tender is the supply and installation of an integrated low-temperature high-magnetic-field measurement system which we intend to use for the characterization of newly-discovered materials. For a good characterization of properties, the system needs to be able to regulate the sample temperature (at least on 1.9 K to 400 K) and apply a magnetic field (up to at least 14 T). |
| Internal reference: | ZN25_08541-MT |
| Type of the bidding procedure: | Open procedure in accordance with Article 40 of PPA-3 |
| Type of purchasing: | Equipment |
| Division into lots: | The public tender is not divided into lots. |
| Variants: | Variants are not permitted. |
| The language of the bids: | The Bidders must submit their bids in Slovene or English. |
| Requested validity of the tender: | 30 November 2025 |
| Deadline for queries: | 4 September 2025 do 16:00 |
| Deadline for submission of bids: | 11 September 2025 till 10:00 |
| Date and time of the public opening: | 11 September 2025 till 12:00 |

Detailed specification of the tender can be found from the technical specifications - point 2 of these instructions.

# 1.3 implementation of the tender

The Contracting Authority carries out the process of awarding a public contract on the basis of the applicable law and by-laws governing public procurement, in accordance with the applicable legislation governing the field of public finance and the field that is the subject of public procurement.

An entity that is subject to an absolute business ban based on the provision of Article 35 of the current Act on Integrity and Prevention of Corruption may not participate in the execution of a public contract.

If the tenderer submits a false statement or gives untrue information about the stated facts, this results in the inadequacy of the offer or the nullity of the contract.

# 1.4 access to the tender documentation

The forms from the tender documentation (TD) are available at the Public-Procurement Portal (https://www.enarocanje.si/). The tender documentation can be accessed free of charge

# 1.5 preparation, submission and delivery of the bids

Any individual or legal entity registered for the service that is the subject of this tender can submit a bid to this tender.

To prepare their bids, the Bidders must give all the details required with the attached specification. The subject of the bid must meet at least the minimum technical requirements specified in the technical specification, which is an integral part of this tender documentation.

The tenderer submits the tender documentation in such a way that, after registration or logging into the ePonudbe system at the address: http://eponudbe.si , for the relevant public procurement, he selects the option "Login and submit a tender", which opens the page for preparing the tender.

Detailed instructions are available in the General Terms and Conditions of the ePonudbe portal and published at address https://eponudbe.si/en/General-Terms-of-Use   (Article 5)

# 1.6 the bidding price

All prices must include all costs for the realisation of the relevant public contract. The final price must include all its constituent elements. VAT must be stated separately (for the bidders from Slovenia). If the price is not expressed in EUR, the contracting authority will convert the offered price into EUR at the exchange rate of the Bank of Slovenia on the day of bid opening.

# 1.7 a joint bid

A bid may be submitted by a group of contractual partners. This group shall submit a Partnership Agreement stipulating how the public contract awarded within this public-procurement procedure will be executed, specifying, in particular, the leading partner authorized to assume, from the Contracting Authority, the obligations, instructions, and possibly also the payments on behalf of, and for the account of, all the partners, as well as the share and type of services/goods to be provided by each individual partner. The Agreement should clearly stipulate that, in relation to the Contracting Authority, all the partners are collectively and individually liable for the realisation of the entire commitment and each part of it, in accordance with paragraph 3 of Article 81 of the PPA-3.

All the partners in a group have to meet the conditions for the recognition of their competencies with respect to their legal statuses (they must meet all the conditions in section 1.13).

All the bidders in the joint bid must fill out the ESPD individually and provide all the required information.

Form 1 "Proforma invoice" is submitted by all providers participating in the joint offer together (one form signed by at least one of the providers participating in the joint offer).

In the case of a joint offer, legal entities should submit in form "ESPD" and provide all participants in the joint offer (Section A Part II ESPD). Providers participating in the joint offer may also specify one of the legal entities with whom the client will communicate until a decision on the order is made, otherwise the client will address all documents to all providers participating in the joint offer.

# 1.8 a bid involving subcontractors

In a subcontracting relationship the main Contractor transfers the execution of the whole or a part of the awarded public contract to a third party, namely, a Subcontractor. The definition of a Subcontractor is given in point 1 of Article 94 of the PPA-3. The Bidder can perform the award of the contract alone or together with Subcontractor(s).

If the tenderer offers the execution of the contract by Subcontractor(s), it is required to:

- provide the list of all the Subcontractors (each Subcontractor must meet all the requirements under point 1.13) and the type of work that the Subcontractor will be undertaking,

- provide the list of contact information and the legal representatives of the Subcontractors,

- complete Declaration of compliance with the conditions for the Economic Operator for each Subcontractor,

- complete FORM 2.2 – The Subcontractor’s authorisation relating to direct payments made by the Contracting Authority to the Subcontractor(s) and consent, if Subcontractor required.

The main Contractor will be required during the realization of the public tender to inform about any changes to the existing information specified in the preceding paragraph and to send information about new Subcontractors, that are planned to be subsequently integrated into carrying out the work no later than 5 days after the change. In the case of the involvement of new Subcontractors the main Contractor must together with the notice provide the information and documents, specified in the second, third and fourth indents of the preceding paragraph and meet all the requirements under point 1.13

In the event that the Subcontractor in accordance with and in such a way specified in the second and third paragraph of Article 94 of the PPA-3 requires direct payment, a direct payment to the Subcontractor shall be deemed to be obligatory in accordance with this Act and bound by an obligation of the Contracting Authority and the main contractor.

Where the tenderer intends to carry out the contract with a Subcontractor, which requires a direct payment in accordance with this Article:

- the main contractor in the contract shall authorize the Contracting Authority, on the basis of an approved invoice or situation by the main contractor to pay directly to the Subcontractor,

- the Subcontractor submits the consent on the basis of which the Contracting Authority instead of the tenderer settles the Subcontractor's claim against the tenderer,

- the main contractor to its invoice attaches an invoice of a Subcontractor, which was previously approved.

If a direct payment to the Subcontractor is not required, the Contracting Authority of the main contractor will require that not later than 60 days from the payment of the final invoice the main contractor sends a written statement and a written statement from the Subcontractor that the Subcontractor has received payment for the supplied goods directly related to the subject of the contract.

If the main contractor does not comply with Article 94 of PPA-3, the Contracting Authority shall submit a proposal to the National Review Commission about the introduction of a minor offence procedure referred to in point 2 of the first paragraph of Article 112 of the PPA-3.

In the event that the tenderer acts with subcontractors, he indicates in FORM 3 that he is applying with subcontractors and at the same time submits FORM 3.1 - Information and consent of the subcontractor for direct payments.

# 1.9 EXAMINATION AND EVALUATION OF THE BIDS

Based on the conditions and criteria specified in the tender documentation, the Contracting Authority will select the Bidder with whom he will enter into a contract for the supply of goods/performance of the service.

Before awarding the contract, the Contracting Authority will verify the existence and content of data or other information indicated in the tender of the tenderer to which it has decided to award the contract. The Contracting Authority will perform an examination and evaluation in line with Article 89 of the PPA-3.

# 1.10 submission and delivery of the bid

Bidders must submit their bids through the IT ePonudbe.si system at the web address http://eponudbe.si/en

Before submitting their bids, Bidders must register at the web address http://eponudbe.si/ . If the Bidder is already registered in the IT ePonudbe.si, they must sign in at the same address.

A user, that is in the ePonudbe.si information system and authorised to submit tenders for an Economic Operator, submits a tender by clicking the »Submit tender« button. The ePonudbe.si information system records the identity and the time of the submission of tender. The user by submitting the tender expresses and states its will to submit a binding tender in the name of the Economic Operator (Article 18 of the Obligations Code, Official Gazette of the Republic of Slovenia, no. 97/07 – official consolidated version, 64/16 – dec. US and 20/18 – OROZ631). With its submission, the tender is binding for the time stated in the tender except if the user withdraws or changes it before the time for submitting the tenders has expired.

A bid is deemed to be submitted on time if the Contracting Authority receives it via the ePonudbe.si system http://eponudbe.si/en  by 11 September 2025 at 10.00 AM. A bid is deemed submitted if it is marked with the status "SUBMITTED" (“ODDANO”) in the IT ePonudbe.si system.

Bidders may withdraw or change their bids by the deadline for submission of the bids. Withdrawal of the bid can be performed on the ePonudbe.si system in the profile of the Bidder that has submitted the bid to the system, when the bid is submitted, with the optional statement about the reason for the withdrawal, the change of the bid is performed within the same system first by withdrawal of already submitted bid, and afterwards with submission of a new bid before expiry of the deadline.

It is not possible to submit any bid after the expiry of the deadline for the submission of bids.

The page for submitting an e-bid in this e-procurement procedure can be accessed here http://eponudbe.si/en.

Instructions for foreign bidders:

We kindly ask foreign providers to start the process at least three days before the deadline. Please follow the instructions https://eponudbe.si/en/General-Terms-of-Use (Article 5)

# 1.11  THE PUBLIC OPENING OF THE BIDS

Bids shall be opened automatically in the IT ePonudbe.si system on 11 September 2025 and will start at 12.00 PM at the web address http://eponudbe.si/en.

The opening takes place in the following way: at the hour set for the public opening of bids, the IT ePonudbe.si system automatically displays data on the Bidder and variants, if requested or allowed, and enables access to the PDF file that the Bidder uploaded in the ePonudbe.si system under the »Proforma Invoice« (»Predračun«) section. The public display automatically ends after 60 minutes. The Bidders that submitted bids have these data available in the IT ePonudbe.si system under the "Minutes on the bid opening" (»Zapisnik o odpiranju ponudb«) section.

# 1.12 QUERIES RELATING TO THE TENDER DOCUMENTATION

Communication with the Bidders concerning queries about the tender and the preparation of a bid shall be organised through the Public-Procurement Portal of the Official Gazette of the Republic of Slovenia: www.enarocanje.si.

Queries about the tender documentation or any other query about the call for tender placed through the Public-Procurement Portal by 16:00 on 4 September 2025 shall be considered as timely. After this time, the Contracting Authority shall not answer any queries about the tender.

Information provided by the Contracting Authority or through the procurement portal shall be, in accordance with the paragraph 2 of Article 67 of PPA-3 understood as an amendment, supplement or clarification regarding the award of the contract, if the content of this additional information results in changing this documentation, or if the explanation eliminates the ambiguity of the statements in this document or the public announcement.

# 1.13 GROUND FOR EXCLUSION OF THE BID

The economic operator shall verify compliance with the conditions by submitting the completed FORM ESPD.

| P1: Reasons related to criminal convictions | Proof |
| --- | --- |
| The Economic Operator must be excluded by the Contracting Authority from participation in a public procurement procedure in the case it is found that the Economic Operator itself or any person who is a member of its administrative, management or supervisory body or has powers of representation, decision or control therein has been the subject of a conviction by final judgment which has the elements of the criminal offenses that are defined in the first paragraph of Article 75 of the PPA-3 In the case that the Economic Operator is in the situation referred to in the above paragraph, in accordance with paragraph 9 of Article 75 of the PPA-3, the Contracting Authority may submit evidence that it has taken sufficient measures to prove its reliability despite the existence of reasons for exclusion. Payment or commitment to pay compensation for all damage caused by a criminal act or violation, active cooperation with investigative authorities for the complete clarification of facts and circumstances, and the adoption of concrete technical, organizational, and personnel measures appropriate to prevent further criminal acts or violations are considered sufficient measures. When evaluating the measures taken by the economic entity, the client takes into account the seriousness and special circumstances of the criminal act or violation. If the client assesses that the evidence submitted by the business entity is sufficient, the business entity shall not be excluded from the public procurement procedure, regardless of the first paragraph of Article 75 of the Public Procurement Act-3. If the client assesses that the measures are not sufficient, it sends the justification for such a decision to the economic entity. | Completed ESPD form. (»Part III: Exclusion grounds, Section A: Grounds relating to criminal convictions«) for all economic operators involved in procedure and Section D: Purely national exclusion grounds«) for all economic operators involved in procedure. |
| P2: Reasons related to the payment of taxes or contributions for social security Proof |  |
| The Economic Operator must be excluded by the Contracting Authority from participation in a public procurement procedure in the case that it does not fulfil the obligations related to taxes and other monetary non-fiscal obligations pursuant to the law regulating the financial administration, collected by the tax authority in line with the provisions of the country in which it has its our registered office or in the Member State of the Contracting Authority and have no outstanding liabilities as on the day of submitting the bid worth EUR 50 or more. On the day of submitting the bid, the Economic Operator must submit all accounts of tax deductions for the incomes arising from employment relationships for the period of the last 5 years. | Completed ESPD form. (»Part III: Exclusion grounds, Section B: Grounds relating to the payment of taxes or social security contributions«) for all economic operators involved in procedure. |
| P3: National provision – Register of economic operators with negative references Proof |  |
| The Economic Operator must be excluded by the Contracting Authority from participation in a public procurement procedure in the case it is on the deadline for the submission of bids excluded from the public award procedures due to being included in the record of entities with negative references, as defined in point a) of the fourth paragraph of Article 75 of the PPA-3 | The proof: Completed ESPD form. (»Part III: Exclusion grounds, Section D: Purely national exclusion grounds«) for all economic operators involved in procedure. |
| P4: National provision – offence related to payment for work | Proof |
| The Economic Operator must be excluded by the Contracting Authority from participating in a public procurement procedure it has been in the 3 years prior to the expiry of the deadline for submitting the bids find twice for an offence related to the payment for work, working hours, on rest, on the performance of work on the basis of civil law contracts, despite the existence of elements of an employment relationship or in connection with the employment of illegal workers by way of a final decision of the competent authority of the Republic of Slovenia or another Member State or a third country. In the case that the Economic Operator is in the situation referred to in the above paragraph, in accordance with paragraph 9 of Article 75 of the PPA-3, the Contracting Authority may submit evidence that it has taken sufficient measures to prove its reliability despite the existence of reasons for exclusion. Payment or commitment to pay compensation for all damage caused by a criminal act or violation, active cooperation with investigative authorities for the complete clarification of facts and circumstances, and the adoption of concrete technical, organizational, and personnel measures appropriate to prevent further criminal acts or violations are considered sufficient measures. When evaluating the measures taken by the economic entity, the client takes into account the seriousness and special circumstances of the criminal act or violation. If the client assesses that the evidence submitted by the business entity is sufficient, the business entity shall not be excluded from the public procurement procedure, regardless of the first paragraph of Article 75 of the Public Procurement Act-3. If the client assesses that the measures are not sufficient, it sends the justification for such a decision to the economic entity. | Completed ESPD form. (»Part III: Exclusion grounds, Section D: Purely national exclusion grounds«) for all economic operators involved in procedure. To the extent that your answer in this case is YES, and you are exercising the redress mechanism, in the “Describe them” field, write the violations and measures that you can use to provide your reliability despite the existence of grounds for exclusion. |

In all cases 1 copy of a proof is sufficient, even when a particular proof is mentioned several times!

With respect to all the above documents, which can be submitted as photocopies, the contracting authority can, at a later stage, request the original copies that the bidder has to submit for viewing to the contracting authority in a period of 3 days. Irrespective of the required age of the documents, they have to disclose the status of the bidder that is legally relevant with respect to the conditions on the day of submitting the bid.

The contracting authority can also consider the data from the official records that the bidder obtained or submitted during the other public-tender procedures if these records are not older than 4 months.

If the country where the bidder is established cannot issue the proofs, or documents, required by the tender documentation, the bidder can, in line with Article 42 of PPA-2, submit an appropriate sworn statement given by the bidder or by a witness in the presence of a judicial or administrative authority, a notary, or an authorised body representing the commercial entities in the bidder’s country.

# 1.15 COMPLETION OF THE BID – ADMISSIBLE TENDER

»Admissible tender« shall mean a tender that is submitted by a tenderer for which there are no grounds for exclusion and which meets the selection criteria, which meets the needs and requirements of the Contracting Authority set out in the technical specifications and the procurement documents, which was received in due time, for which there is no evidence of collusion or corruption, which has not been found by the Contracting Authority to be abnormally low, and the price does not exceed the Contracting Authority’s budget.

Tender documentation must consist of the forms submitted in the following order:

  - FORM 1 – The bid (Proforma invoice)

  - FORM 2 – Details about the Bidder

  - FORM 3 – Acting with subcontractors/without subcontractors

  - FORM 4 – the completed, signed and stamped sample contract initialled on each page

  - Completed »ESPD« form

  - Appendix 1 – Statement with details about the natural and legal entities owned by the bidder

  - Evidence relating to the fulfilment of the requirements of the technical specifications (i.e. brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included)

Only when acting as consortium and/or with Subcontractors the Bidder must submit the following forms:

  - Partnership Agreement

  - FORM 1 – The bid (Proforma invoice)

  - FORM 2.1 – General data about the Bidder and the consortium

  - FORM 3 – Acting with subcontractors/without subcontractors

  - FORM 3.1 – The subcontractor’s authorisation to direct payments made by the contracting authority to the subcontractor(s) and consent

  - FORM 4 – the completed, signed and stamped sample contract initialled on each page

  - Completed »ESPD« form

  - Appendix 1 – Statement with details about the natural and legal entities owned by the bidder

  - Evidence relating to the fulfilment of the requirements of the technical specifications (i.e. brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included)

# 1.16 FORM »BID (PROFORMA INVOICE)«

The Bidder must offer all listed items from the Form 1 – Bid (Proforma invoice), considering the technical description and the specification of the equipment, which is part of the Tender Documentation.

The Bidder completes all the listed items in the Proforma invoice, to a maximum of two decimal places.

If the Bidder declares a price of zero (0) EUR, it is understood this item is offered free of charge.

The Bidder must not change the content of Form 1. The price must include all discounts and costs.

Where the Contracting Authority, during the examination and evaluation of tenders, finds that a calculation error has occurred, it may correct it in line with Article 89, paragraph 7, of the PPA-3.

Bidder uploads in the IT ePonudbe.si system under the section »Pro-forma Invoice« (»Predračun«) in a pdf file the Proforma Invoice.

# 1.17 FORM »ESPD« FOR ALL ECONOMIC OPERATORS

European Single Procurement Document (ESPD) is a formal statement by the economic operator that the relevant grounds for exclusion do not apply, that the relevant selection criteria are fulfilled and that it will provide the relevant information as required by the contracting authority or contracting entity. The ESPD form shall also include a formal statement that the contracting authority may ask any tenderer at any moment during the procedure to submit all or part of the required certificates and supporting documents where this is necessary to ensure the proper conduct of the procedure.

The entries in the ESPD form and/or the evidence submitted by the economic operator must be valid.

The economic operator imports XML.file prepared by the contracting authority to the ESPD file on the website: https://ejn.gov.si/espd and directly enters the requested data into it.

Where groups of economic operators (including temporary associations, economic operators on whose capacity is referred to by the economic operator and subcontractors) participate together in the procurement procedure, a separate ESPD must be given for each of the participating economic operators.

The bidder who submits the bid in the ePonudbe.si system uploads its ESPD in the section »ESPD-Bidder« section and places the other ESPDs in the »ESPD-Other Participants« section. A bidder who submits an offer in the ePonudbe.si system shall upload the electronically signed ESPD into xml. format or unsigned ESPD in xml. in the latter case, in accordance with the General Terms and Conditions of use of the ePonudbe.si information system, a legally binding document having the same validity as is signed shall be deemed to have been submitted.

For the other participating economic operators from the group, the bidder attaches signed ESPDs in the pdf. form or electronically signed xml. in the »ESPD – Other Participants« section.

# 1.18  CRITERIA FOR ASSESSING AND EVALUATING THE BIDS

The criterion for the best Bidder selection is the most economically advantageous bid, determined on the basis of the lowest bidding price in euros without the VAT, provided that all the requirements are fulfilled.

In the case of two or more Bidder having equally economically advantageous bid, the Bidder with the shorter delivery date and exact requested specifications will be selected. If the bids are tied after this, the bidder with the most favourable warranty terms and conditions will be selected.

# 1.19 ANTI-CORRUPTION CLAUSE

In the process of awarding a public contract, the client and bidders may not initiate and perform actions that would predetermine the selection of a specific offer, or that would cause the contract not to enter into force or not be fulfilled. Any kind of lobbying in public procurement procedures is prohibited.

# 1.20 CONTRACT AWARD NOTICE

The Contracting Authority will publish the signed decision on awarding the contract on the public procurement portal. The decision is considered to have been served on the date of its publication on the public procurement portal.

# 1.21 RIGHT OF WITHDRAWAL FROM THE PUBLIC CONTRACT

The Contracting Authority may, according to the eighth paragraph of Article 90 of the ZJN-3, after the decision to award the contract has been made, withdraw from the performance of the public contract until the conclusion of the contract for justified reasons, that they no longer need the object of the public contract or that they do not have guaranteed funds for it, or that a valid reason arises with the client suspicion that the content of the contract was or could be the result of a committed criminal act or that other extraordinary circumstances arose which the client could not influence and foresee and which made the performance of the public contract with the selected provider impossible. In this case, the contracting authority will notify the bidders in writing of its decision and the reasons for withdrawing from the execution of the public contract.

# 1.22 SIGNING OF THE CONTRACT

In line with the 6th paragraph of Article 14 of the Integrity and Prevention of Corruption Act (Official Gazette of the Republic of Slovenia, No. 69/11–official consolidated version and 158/20), the selected Bidder is obliged to submit, prior to signing the contract and at the request of the Jožef Stefan Institute, the awarding authority, a statement or details about the natural or legal entities owned by the selected Bidder, including the silent partners and business entities that are considered to be associated with the selected Bidder in line with the provisions of the law governing companies. If the Bidder submits a false statement or gives untrue information about the required details, the contract shall be annulled.

The selected bidder shall submit to the Contracting Authority, within 8 days of the receipt of the request, the details about:

- its founders, partners, shareholders, limited partnerships or other owners, as well as the equity shares of these entities;

- business entities that are considered to be associated with the selected bidder on the basis of the law governing companies.

The selected Bidder shall sign the contract within 8 working days of the receipt of the contract signed by the Contracting Authority.

Prior to signing, the contract shall be modified depending on whether the selected Bidder submits a joint bid, a bid with Subcontractors, a bid consisting of different realisation phases and the similar.

# 1.23 LEGAL INSTRUCTION

A review claim referring to the contents of the call and/or the tender documentation, may be filed in 10 working days from the date of publication of the call for tender or the notice of supplementary information, information on the pending procedure or correction if this is a notice to amend or supplement requirements or criteria for selecting the most advantageous tenderer from the tender documentation or previously published contract notice; the review claim may refer to the amended, supplemented or clarified content of the publication or the tender documentation or a directly related reference to it in the original publication or tender documentation. A review claim may not be submitted after the deadline for submission of the bid, unless the deadline submission of the bid is less than 10 working days. In this case, a review claim may be filed within 10 working days of the date of publication of the contract notice.

A request for a review shall be submitted via the eRevizija (eReview) portal (https://www.portalerevizija.si).

The applicant shall pay a tax of 4.000,00 euros to the bank account of the Ministry of Finance, No. SI56 0110 0100 0358 802, opened at the Bank of Slovenia, Slovenska 35, 1505 Ljubljana, Slovenia, SWIFT CODE: BS LJ SI 2X; IBAN: SI56011001000358802 – a tax for a revision of public procurement.

# 2. TECHNICAL DOCUMENTATION OF THE PUBLIC TENDER

The subject of the public tender is the supply and installation of an integrated low-temperature high-magnetic-field measurement system for physical properties. More precisely, the aim is to buy a measurement system, which will enable us to measure the heat capacity, electrical transport (including sample rotation; including a modular approach to allow custom measurements) and thermal transport for small pieces of newly-discovered materials in a wide range of temperatures and in applied magnetic fields up to at least 14 T.

TECHNICAL SPECIFICATIONS:

Integrated measurement system – general requirements

- The automated integrated measurement system should be capable of varying the sample temperature and magnetic field during measurements. The software supplied with it should allow control of all parts of the measurement system – the electronics, the hardware (including the superconducting magnet, temperature regulation and the atmosphere in the sample chamber) and also the data acquisition,

- Simple automation of complex measurement procedures (including changing the temperature and magnetic field during the measurement procedure) should be possible via the provided software. It should be possible to save and reuse the programs, scripts or sequences, which control the measurement system during automated operation.

- At least 15 measurement modules should be available for purchase, which include all the hardware and software so that high-quality measurements of physical properties can begin immediately, without the need to develop our own hardware or software.

- The design of the system should also be such that custom user experiments can be performed, where our own hardware or software is used for part of the experiment without too much effort.

- The system should be modular. Namely, it should be designed in such a way that samples are mounted on small sample measurement inserts for insertion into the system. The measurement inserts (when not used with extra cryostats or rotators) should couple directly to at least 12 wires built into the cryostat of the integrated measurement system (and thus well thermally anchored) so that most experiments can be performed reliably without extra top-to-bottom wiring that could cause unneeded thermal leaks. The actual measurement hardware located outside of the dewar (e.g. a resistance measurement module) then uses the 12 or more wires coming out of the cryostat to perform the measurements.

- The measurement system should include a microprocessor, which controls the basic functioning of the system. With this microprocessor, the system should be able to perform at least basic electrical measurement (current and voltage sourcing, resistance measurements) without extra hardware.

- Sample chamber of at least 25 mm diameter, so that custom experiments can be designed easier.

- The measurement system should be capable of performing measurements during slow temperature sweeps or other faster changes of the sample temperature (both during heating and warming), as well as during sweeps and changes of the magnetic field (both those increasing and decreasing the magnetic field).

- The helium dewar of the system should be nitrogen-jacketed to decrease the helium evaporation rate.

- The integrated measurement system should be a “wet” system – that is a system that requires periodic refill of cryogens and does not contain a cryocooler, e.g. a pulse tube cryocooler.

- The integrated measurement system should contain a vacuum pump that is able to quickly purge the atmosphere in the sample chamber (via a purge cycle of the form vacuum-gas-vacuum-gas).

- The software supplied with the system should be a professionally written Windows application (and not just a simple collection of Labview instruments or simple scripts). Access to automatization and scripting should be available without a licen se or should be included in a license supplied together with the measurement system. It should be possible to use the software also on other computers to analyse the measurements, postprocess them or export them, either without a license or with a license included with the purchase of the measurement system.

- The system should enable interfacing with external instruments – to some extent via the provided software, to a larger extent so that at least the programming environments of LabView and Python can be used to control the experiments.

- The integrated measurement system should be physically and in software aspects compatible with the already owned Quantum Design He3 and Quantum Design Dilution Refrigerator cryostats.

- Temperature range for the sample chamber and therefore temperature range for the measurements should be at least 1.9 K to 400 K. The cryostat of the measurement system should be able to stabilize all temperatures in the measurement range, including temperatures below 4.2 K, continuously (that is for infinitely long time) and also during measurements.

- Typical temperature stability 0.2% or better for temperatures below 20 K. Typical temperature stability 0.02% or better for temperatures above 20 K.

- Temperature accuracy 1% or better.

- Typical temperature rates: 6 K/minute or better during cooling; 10 K/minute or better warming. The typical time for cooling one or more samples from 300 K to 1.9 K should be 40 min or better, whereby this cooling actually occurs with the cooling of the cryostat down from room temperature and not with the insertion of the sample or samples into an already precooled cryostat – so that it is possible to perform test measurements of physical properties already during cooling.

- Appropriate design of the cryostat so that changes of the magnetic field do not influence the sample temperature.

- Appropriate design of the cryostat so that temperature sweeps throughout the whole temperature range are smooth and monotonic, including potentially problematic temperatures such as 4.2 K and switch points from high- to low-temperature operation.

- The cold helium gas, which provides cooling to the cryostat (and therefore to the sample), flows around the sample space (so through some type of cooling sleeve or »annulus«) and not directly through the sample space. Therefore, the thermal gradients on the sample and in its vicinity are minimized.

- The implementation of the temperature regulation of the measurement system should function without mechanic parts at low temperature (e.g. cryogenic needles valves). An appropriate design is one with multiple impedance – e.g. one used exclusively for low temperature regulation and a larger one for regulation at higher temperatures.

- The measurement system should use for determining the sample temperature a combination of a thermometer which is accurate at high temperature (e.g. a Pt100) and a thermometer which is accurate at low temperature (e,g. a Cernox). These thermometers should be combined into a smooth temperature dependence without any steps at any temperature within the temperature range of operation of the system.

- Another thermometer should be located at a location higher in the sample chamber, so that in conjunction with the sample thermometer it is possible to estimate the temperature gradients in the sample chamber.

- The measurement system should support the usage of additional thermometers (»user thermometers«), which are thermometers that can be located nearer to the sample – e.g. integrated into the measurement insert (small puck-like ones or larger ones such as a rotator) or into any additional cryostat, which can be inserted into the integrated measurement system.

- Maximum magnetic field at least 14 T (or in cgs units 140 000 Oe), that can be applied in the direction of up or down. The magnetic field is provided via a superconducting magnet to ensure low noise in the magnetic field.

- Variation of magnetic field at most 0.1% over a distance of 5.5 cm on-axis at the sample location

- The magnetic field is reported with a resolution of 22 µT or better (or in cgs units 0.22 Oe or better).

- Maximum field change rate at least 0.01 T/s (or in cgs units at least 100 Oe/s)

- A typically achievable minimum magnetic field change rate of 10 μT/s or smaller (or in cgs units 0.1 Oe/s or smaller)

- Supports at least three drive modes for the superconducting magnet: an approach to the final value with a constant rate of change (a »linear« ramp), an approach which slows down near the target magnetic field (that is a non-overshooting approach), and an approach which allows to get rid of trapped flux in the superconducting magnet by oscillating the magnetic field around the target magnetic field (starting with large oscillations and going to ever smaller ones)

- The sample chamber can be sealed with helium gas either at normal pressure (1000 mbar), at reduced pressure / low vacuum (somewhere in the range 1 to 6 mbar) or put in a state of high-vacuum.

- High vacuum pump and appropriate sealing of the sample chambers: the pressure in the sample chamber can be reduced to 0.1 mTorr or smaller. If the high vacuum pump is a cryopump, guided regeneration procedures should be available via the software.

- With the integrated measurement system, a baffled rod (a rod with horizontal plates) is included, which can be inserted into the upper part of the sample chamber during measurement. The purpose of the baffled rod is to reduce the amount of convection in the sample space and the amount of thermal radiation that can travel down the sample chamber. The rod therefore makes cooling the sample easier by reducing the unwanted thermal currents towards the sample. At the end of the baffled rod, there is a removable piece of active charcoal, which further lowers the pressure in the vicinity of the sample.

Heat capacity module/option – module requirements

- Microcalorimeter for small samples based on small inserts/pucks. These microcalorimeter inserts have a small sample platform, which is suspended on thin wires to thermally decouple it from the surroundings. Both the heater as well as the thermometer needed for the measurements of the sample heat capacity are attached to the platform.

- The measurement takes place in high vacuum inside the sample chamber.

- Capable of operating in whole temperature range 1.9 – 400 K, at all magnetic field 0 T – 14 T.

- Thermal pulse method for extracting heat capacity – operates by recording the temperature response and a fitting algorithm based on a thermal circuit model.

- Additional analysis method available based on the slope of the temperature response during pulses, by which high-resolution data can be obtained in post processing.

- Programs, commands or scripts prepared to measure the temperature dependence of heat capacity of a sample, commands prepared to calibrate and properly measure also at constant applied magnetic field.

- Fully automated operation – data for one or several temperature sweeps can be collected automatically without operator presence.

- Automatic adaptation of heat pulse duration based on the sample time constant. Heat pulse duration is automatically adapted to properly reflect changes of heat capacity with temperature so that accurate data can be obtained.

- Automated background (addenda) acquisition and automatic background subtraction during sample measurements.

- The fitting algorithm in the heat pulse method is able to account for finite thermal contact between the sample and calorimeter to increase accuracy when the coupling is not perfect.

- Provided mounting station which limits motion of the calorimeter sample platform so that samples can be mounted easily and without damage to the calorimeter.

- Software enables choice of units, e.g. to report heat capacity.

- Software enables guided setup of the calibration of microcalorimeters with regards to the main system thermometer and automatic execution of the procedure.

- Available as add-on / later purchase: Adapter to change orientation of the sample so that the magnetic field is no longer perpendicular to the mounting/attachment plane of the sample but can be applied in-plane

- Accuracy: ± 5% or better for 2 K to 300 K; ± < 2% typical.

- Resolution: 10 nJ/K at 2 K or better.

- Calorimeter platform area (space available for mounting samples): 3 mm x 3 mm or larger, typical range of samples size is 1 to 500 mg or larger.

- Typical background (addenda) magnitude 0.2 μJ/K at 2 K; 15 mJ/K at 400 K.

- At least 2 microcalorimeters included, at least one additional spare for the sensitive parts of the calorimeter (e.g. wires and platform).

- Compatibility with existing Quantum Design He3 refrigerator so that the temperature range can be extended to 0.4 K to 350 K with He3; operation in magnetic fields.

- Compatibility with existing Quantum Design DR refrigerator so that measurements can be performed in the temperature range 50 mK – 4 K; operation in magnetic fields.

- Available as add-on / later purchase: He3 calorimeter inserts with background 10 nJ/K, 2.25 μJ/K, 340 μJ/K and 11 mJ/K or smaller at temperatures 0.4 K, 4 K, 35 K and 350 K, respectively.

- Available as add-on / later purchase: dilution refrigerator calorimeter inserts with background 2.5 nJ/K, 225 nJ/K and 1.5 μJ/K or smaller at temperatures 0.05 K, 2 K and 4 K, respectively.

Thermal transport module/option – module requirements

- Can simultaneously and with a high sampling frequency measure the sample electrical resistivity, thermal conductivity and Seebeck coefficient, so that the temperature dependence of the thermal figure of merit ZT can be determined in one temperature sweep.

- Includes thermal transport measurement insert/puck with two thermometers and a heater, which allows proper attachment of the sample.

- The measurement of the thermal transport it performed in vacuum (<0.1 mTorr, with a cryopump, turbopump or similar) to guarantee its accuracy.

- The module allows for a non-stationary measurement during a temperature sweep so that good thermal stability is not required for each measurement point. A thermal pulse is applied to the sample and the sample temperatures and Seebeck voltage are recorded during the pulse. Then the sample thermopower and thermal conductivity are extracted via a model, which properly accounts for both the thermal circuit and the temperature sweep/ramp. An electrical resistivity measurement is performed between the heat pulses.

- The module also allows for a stationary measurement without the advanced temperature sweep and heat pulse method – one in which thermal stability of the sample is first acquired and then measurements are performed.

- Automated operation – the measurement module and software can acquire data for one or several temperature sweeps automatically without operator intervention or presence. If adjustment of measurement parameter is needed during the measurement, the parameters are adjusted automatically. The automated algorithm is adaptive and capable of measurements during temperatures ramps so that thermal stability during measurement is not required.

- The measurement of all the thermal transport properties (thermal conductivity, electrical resistivity, Seebeck coefficient) is performed in the software in a simple manner – e.g. with a one-line command. The software can also automatically determine the figure of merit ZT and its uncertainty in this one-line command or other simple method.

- Raw data of the temperature pulses (including Seebeck voltage, temperatures, electrical resistivity etc.) is collected and stored, so that quality of the measurement can be assessed during the measurement itself of so that the data may be reprocessed – either by the thermal transport software or custom user software – later.

- The modules hardware is designed in such a way that two- and four-point thermal transport measurements can be performed.

- Any required mounting hardware, lead material and at least one appropriate conductive epoxy glue can be purchased from the supplier of the integrated measurement system.

- A spare set of thermometers and heater is included in the purchase. Spare thermometers and heaters should be available for later purchase from the supplier of the integrated measurement system.

- An appropriate mounting station should be provided, so that the puck can be fixed in place and tilted for easier mounting of the sample

- A nickel reference sample should be provided, so that the measurement module can be checked for proper operation at any time.

- Operating range 1.9 K to 400 K, also in magnetic fields up to 14 T if the measurement insert is properly calibrated.

- Accuracy of the determined thermal conductance has an upper bound of either 5% (or smaller) or fixed values listed later (or smaller), whichever is greater. The fixed values are: 2 μW/K for temperatures below 15 K, 20 μW/K for temperatures 15 to 200 K, 0.5 mW/K for temperatures 200 to 300 K and 1 mW/K for temperatures higher than 300 K.

- The estimated dynamic ranges for the thermal conductance of samples with normal aspect ratios should be as specified or greater than: 1 to 25 mW/K for high temperatures (e.g. 400 K); 100 μW/K to 100 mW/K for temperatures in the vicinity of 50 K; 10 μW/K to 1 mW/K for low temperatures (e.g. 1.9 K)

- The estimated dynamic range for thermal conductivity for samples with normal aspect ratios should be 0.1 to 250 W/m∙K (or greater) at T = 300 K. The module should allow mounting both of flat samples (“pancake”-like) and long samples (“rod”-like) to increase the range of potentially measurable thermal conductivities.

- The accuracy of the Seebeck coefficient measurements should be better than the following: The accuracy should have an upper bound given by 5% of the Seebeck coefficient value, 0.5 μV/K or as determined by an error in the Seebeck voltage of 2 μV.

- The estimated dynamic range for the Seebeck coefficient for samples with a normal aspect ratio should be at least 1 μV/K to 1 V/K.

- The precision of the measurement of the electrical resistance should be 0.01% or better for 1 Ω at an electrical current of 200 μA.

- The estimated dynamical range for the electrical resistance should be 10 μΩ to 5 MΩ for samples with a normal aspect ratio.

- The typical accuracy of the determined thermoelectric figure of merit should be 15 % or better

- The measurement module should support thermal transport measurements at typical slew rates of ± 0.5 K/min for temperatures greater than 20 K and ±0.2 K/min at temperatures smaller than 20 K.

Horizontal rotator – module requirements

- The measurement module/option allows the measurement of the angular dependence of the electrical resistivity of samples via a motorized rotator.

- The rotation is around an axis perpendicular to the applied field of the integrated measurement system.

- Automated operation, where the measurement system can record accurate angular positions (via an automated encoder) and temperatures (via an on-board thermometer near to the samples). The system can collect the data for several angular sweeps at different temperatures automatically without operator intervention or presence.

- The software allows for simple programming of several angular sweeps at different temperatures or for several temperature sweeps at different angles.

- The horizontal rotator should work (that is enable simple measurements) with the included electrical resistivity options provided with the integrated measurement system.

- The hardware is designed in such a way that besides the complete rotator measurement insert, the module has a small sample holder onto which the samples can be easily mounted before mounting to the rotator

- The rotator has an integrated thermometer, which is in good thermal contact with the sample holders.

- The materials of the insert should be chosen so that the effects of the temperature and magnetic field on the angle and measurement are minimised.

- At least two types of sample holders are provided which are compatible with the electrical resistivity measurement options of the integrated measurement system. One holder type should orient the samples in such a way so that the axis of rotation is in plane with regards to the mounted samples. The other holder type should orient the samples is such a way that the axis of rotation is perpendicular to the plane of the mounted samples. These sample holders should provide at least two electrical measurement channels for samples and these two channels should be four-terminal.

- A more general extra holder type should be available for purchase to facilitate custom experiments.

- The range of angles accessible should be at least -10° to 370° relative to the magnetic field.

- The angular step resolution should be 0.0133°/step or better when consecutive steps are driven in the same direction.

- The backlash should be smaller than 10°.

- The rotator module should be able to perform measurements at least for temperatures between 1.9 K and 400 K.

- The rotator module should enable measurements for magnetic fields up to 14 T.

DC resistance/resistivity – module requirements

- The module/measurement option should enable measurement of at least three four-terminal channels in its standard configuration with standard hardware; there should be an additional four-terminal measurement channel available for custom experiments. Each channel may be turned on or off from the software.

- The measurement option includes at least two measurement inserts that enable use of three four-terminal channels without remounting the samples.

- Channels measured in series via multiplexing and not simultaneously; channels are disconnected while they are not being measured.

- The resistance bridge channels can be configured individually – setting a current, voltage and power limit for each channel separately, so that damage to sensitive samples can be avoided.

- Polarity switching (using an internal bipolar power supply) can be turned on for each channel, so that DC offsets in the resistance/resistivity can be eliminated.

- The measurement module/option is fully integrated into the main software of the system, so that it can be easily used. Enables automatic calculation of resistivity from resistance from sample dimensions (for block shaped samples). Can also report conductivity and conductance.

- The module offers automated current-voltage (“V-I”) curve measurement with appropriately adjustable scanning of the excitation.

- Square-wave-shaped DC current excitation (synchronized with the oscillations of the powerline to increase accuracy); polarity switching is automatically taken into account into the results without postprocessing.

- It is possible to set the number of acquired datapoints for the electrical resistance (or resistivity) that are averaged together to get a single resistance (or resistivity). The acquisition rate (without initial calibration with the help of internal resistors) should be at least 8 datapoints per second.

- Available current range at least 10 nA to 5 mA, current ranges in reasonable steps, with automatic switching between ranges.

- 95 mV or larger compliance voltage.

- Measurement range: 10 μΩ to 4 MΩ or larger.

- Typical sensitivity: specified as noise in the measured voltage of 20 nV RMS or better or equivalently noise in resistance of 4 μΩ or better at 5 mA of current.

- Can operate in the whole temperature range (1.9 K to 400 K) and full range of magnetic fields (0 to 14 T).

- Compatibility with existing Quantum Design He3 refrigerator so that the temperature range can be extended to 0.4 K to 350 K with the He3 refrigerator; operation in magnetic fields.

- The option can be used to read the thermometers of the instrument.

- Is compatible with the horizontal rotator that is also being bought so that angular-dependent resistivity measurements may be performed.

- Calibration resistors available inside the resistivity unit, which can be used periodically (on the scale of minutes) during the measurement, so that the reported sample resistance/resistivity values are properly calibrated. This option can be used automatically through the software and also disabled if needed.

A modular set of additional measurement hardware for electric measurements

- As part of the integrated measurement system we want to purchase also an extension – a set of modular units, which are aimed at custom user experiments.

- The main module out of this set of modular units should be able to control at least 3 source modules (for voltage or current) and at least 3 measure units (for voltage or current). It should be possible to use any combination of these sources and measurement units.

- The modular set should be provided as a package of at least 3 current source modules, 3 voltage measurement modules, 2 voltage source modules and 2 current measurement modules.

- The modules should use a common clock for ADC/DAC conversion, so that they are time-wise synchronized.

- All modules, both sources and measurement units, should be capable of operating in DC or AC mode (to 100 kHz), in some regimes they may be capable also operating in a combination of both.

- Voltage measurement module: single-ended or differential operation; DC or AC input coupling; measurement ranges at least 10 V, 1 V, 100 mV, 10 mV; automatic switching between measurement ranges; best sensitivity 1nV or better in the 10 mV range, 10 s measurement; input hardware low- and high-pass filters with 20dB or 40dB per decade with filter frequencies available of at least 10 Hz, 30 Hz, 100 Hz, 300 Hz, 1 kHz, 3 kHz; input impedance at least 10 GΩ, 120 pF at DC coupling; Accuracy for 10 mV range 0.1% + 20 μV or better for DC,  0.1% or better for lock-in measurements to 1000 Hz; Accuracy in differential mode in the 10 mV range 0.05% + 3 μV or better for DC, 0.05% or better for lock-in measurements to 1000 Hz. Noise – without filters, inputs shorted: 3.2 nV/√Hz or better for the 10-mV range at 1 kHz and 4.1 nV/√Hz or better in v differential operation.

- Current source module: can generate signals from DC to AC 100 kHz; output ranges available at least 100 mA, 10 mA, 1 mA, 100 μA, 10 μA, 1 μA, 100 nA, 10 nA; automatic switching between output ranges; compliance voltage is 20 V or more differential mode and 10 V or more in single-ended mode; supports balanced or differential operation so that for example ground loops can be eliminated; can control via an additional input, which point in the circuit is at zero potential and in this way help reduce the effect of the resistance of wires; contains a guard option that may be turned on or off; the output coupling can be set to be either DC or AC; The accuracy at DC operation is 0.5% + 120 pA or better on the1 μA range, 0.5% + 60 pA or better on the 100 nA range and 0.5% + 30 pA or better on the 10 nA range; The accuracies during AC lock-in operation on the output range 1 μA at frequencies up to 1 kHz, on the output range 100 nA up to 100 Hz and in the output range 10 nA up to 30 Hz are 0.5% or better; The noise density on the output ranges 1 μA and 100 nA at 1 kHz is 100 fA/√Hz or less and on the output range 10 nA at 100 Hz it should be 100 fA/√Hz or less; the noise over the band of frequencies DC to 100 MHz is 2mV RMS or less; the output impedance should be greater than 1 MΩ for the output range 10 mA, larger than 10 MΩ for the output range 1 mA, larger than 100 MΩ for the output range 100 μA and larger than 1 GΩ for the output ranges 10 μA, 1 μA, 100 nA and 10 nA.

- Current measurement module: can measure currents from DC to at least AC 100 kHz; allows determination of amplitude, phase and higher harmonics; has a bias voltage feature up to ±10 V for use in experiments, e.g. of photodiodes, where this is necessary; input offset voltage 150 μV or less; available measurement ranges at least 100 mA, 10 mA, 1 mA, 100 μA, 10 μA, 1 μA, 100 nA, 10 nA, 1 nA; sensitivity 10 fA or better in the measurement range 1 nA at a 10 second acquisition; hardware filters of low-pass and high-pass type and attenuation of 20 dB or 40 dB per decade available at least with filter frequencies 10 Hz, 30 Hz, 100 Hz, 300 Hz, 1 kHz, 3 kHz; Accuracy at DC operation on the input range 100 nA and at 20 Hz equal to 0.05% + 10 pA or better, on the input ranges 10 nA and 1 nA and at 10Hz equal to 0.1% + 5 pA or better; Accuracy at AC operation in the input range 100 nA up to a frequency of 20 Hz equal to 0.1% or better, in the input ranges 10 nA and 1 nA at 10 Hz equal to 0.5% or better; Typical noise density in the input ranges 10 nA and 1 nA equal to 4.3 fA/√Hz at 77 Hz or better

- Voltage source module: can generate a combination of a DC and AC signal up to 100 kHz, using a different output range for each of these components; available output ranges at least 10 V, 1 V, 100 mV, 10 mV; automatic switching between output ranges; current limit settable up to  100 mA in DC mode; output impedance less than 150 mΩ; accuracy in DC operation in the range 10 mV equal to 0.15% + 450 μV or better; accuracy in AC operation to 1 kHz equal to 0.15% or better; Typical noise density in the 10 mV range at 1 kHz equal to 30 nV/√Hz or less; wide-band noise for DC to 100 MHz equal to 2mV RMS or less.

- Breakout box or adapter for connecting the wiring of the integrated measurement system to the modular set of additional measurement hardware for electric measurements (3 connections to source modules and 3 connections to measurement modules should be possible).

- A license for software control of the additional measurement hardware should be provided enabling users to build a system configuration of all the modules, use a simple programming interface to program the data acquisition, temperature control, magnetic field control, etc. during the experiment. A driver for communication with the integrated measurement system should be provided.

OTHER REQUIREMENTS:

- The equipment must be CE compliant.

- The equipment must operate in the electrical grid 220/400 V.

- Documentation and manuals (HTML or PDF in English) should be provided, including the maintainance documentation to achieve all technical characteristics.

- Price should include delivery charges and complete installation of the system, including all travel and lodging costs of the provider’s crew

- At or shortly after installation of the equipment, we require a brief training on the correct usage and maintenance of the equipment. The cost of the training should be included in price.

- Delivery and installation: by 25th October 2026 (due to requirements of the funding agency)

- The provider is responsible for technical support and service.

- Free post-training technical support and help by phone, electronic mail and remote access to the instrument, direct contact with application engineer

- Warranty and Out-of-warranty support:

- Warranty at least 12 months

- Response time at most 7 working days after the problem notice.

- After expiration of the warranty the bidder must provide payable after-warranty support for at least 5 years after the initial warranty period

# 3. TENDER DOCUMENTATION

The Bidder prepares the offer by entering all the required data in the forms that are an integral part of the tender documentation and attaching all other required documentation. The offer must be submitted on the forms from the tender documentation or on the bidder's forms, which must not differ significantly in content from the attached forms.

All required tender documents must be completed, signed by the legal representative and stamped.

Tender documentation must consist of the forms submitted in the following order:

  - FORM 1 – The bid (Proforma invoice)

  - FORM 2 – Details about the Bidder

  - FORM 3 – Acting with subcontractors/without subcontractors

  - FORM 4 – the completed, signed and stamped sample contract initialled on each page

  - Completed »ESPD« form

  - Appendix 2 – Statement with details about the natural and legal entities owned by the bidder

  - Evidence relating to the fulfilment of the requirements of the technical specifications (i.e. brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included)

The Bidder submitting the offer accepts criminal and material responsibility that all data and documents included in the tender are true. Otherwise, the Bidder shall be liable to the Contracting Authority for any damage caused to it.

BID (PROFORMA INVOICE)

FORM 1

| Name of the Bidder: |  |
| --- | --- |
| Address of the Bidder: |  |
| Contact person: |  |
| Contact person's E-mail address: |  |

On the basis of your public tender, we are pleased to submit the following bid:

| BID No. |  |
| --- | --- |

| Type of purchasing: |  | Service | X | Equipment |  | Const |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| No. | Description of the goods | Quantity | Price | Discount % | Value |
| --- | --- | --- | --- | --- | --- |
| 1 | SUPPLY AND INSTALLATION OF AN INTEGRATED LOW-TEMPERATURE HIGH-MAGNETIC-FIELD MEASUREMENT SYSTEM FOR PHYSICAL PROPERTIES | 1 Set |  |  |  |

| TOTAL |  |
| --- | --- |
| Discount |  |
| VALUE without VAT |  |

| Payment terms: |  |
| --- | --- |
| Delivery time in weeks: |  |
| Delivery term: | DAP Ljubljana (price includes delivery to the client's address) |
| Validity of the bid: | 30 November 2025 |
| Warranty period: | at least 12 months |

MANDATORY ENCLOSURE:

Brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included (please, enclose it to the second part of the tender documentation)

|  |  | The Bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

DETAILS ABOUT THE BIDDER

FORM 2

| Name of the Bidder: |  |
| --- | --- |
| Address of the Bidder: |  |
| Legal representative(s) of the Bidder: |  |
| VAT identification number: |  |
| Name of the bank, SWIFT, IBAN: |  |
| Contact person: |  |
| Phone: |  |
| E-mail: |  |
| Responsible person for signing the contract: |  |

On the basis of a public tender published on the Public Procurement Portal we are applying to your invitation to tender and enclose our offer documents in accordance with the Instructions for the bidders.

Acting as a Bidder, we are submitting the bid:

|  | Individually, without contract partners or subcontractors |
| --- | --- |
|  | As a joint bid as a consortium, involving contract partners |

|  | As a bid involving subcontractors |
| --- | --- |

|  | As a bid using the capabilities of other operators |
| --- | --- |

|  |  | The Bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The Contracting Authority’s note: Form must be completed and attached to the bid, regardless the fact whether you submit the bid individually, as a joint bid as a consortium or with a joint bid with subcontractors.

GENERAL DATA ABOUT THE BIDDER AND THE CONSORTIUM

FORM 2.1

| Name of the Bidder: |  |
| --- | --- |
| Address of the Bidder: |  |
| Legal representative(s) of the Bidder: |  |
| VAT identification number: |  |
| Name of the bank, SWIFT, IBAN: |  |
| Contact person: |  |
| Phone: |  |
| E-mail: |  |
| Responsible person for signing the contract: |  |

On the basis of a public tender published on the Public Procurement Portal we are applying to your invitation to tender and enclose our offer documents in accordance with the Instructions for the bidders.

Contract partners in a joint bid and description of work:

| No. | Name of the partner | Description of the works | Value | Share in % |
| --- | --- | --- | --- | --- |
| The group leader in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |

|  |  | The Bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The Contracting Authority’s note: Form must be completed and attached to the bid, only whether you submit the joint bid as a consortium.

ACTING WITH SUBCONTRACTORS/WITHOUT SUBCONTRACTORS

FORM 3

Acting as a bidder, we are submitting the bid:

|  | a) individually, without subcontractors |
| --- | --- |
|  | b) with the following subcontractors: |

| No. | Subcontractors |
| --- | --- |
| 1. |  |
| 2. |  |
| 3. |  |
| 4. |  |
| 5. |  |

The bidder shall complete the table »Part of the execution of the contract, which will be performed by a subcontractor« for each subcontractor as many times as different parts of the execution of the contract will be taken by each subcontractor.

Part of the execution of the contract that will be performed by a subcontractor:

| Subcontractor |  |
| --- | --- |
| Description of work |  |
| Quantity |  |
| Value |  |
| Place |  |
| Time of completion |  |

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The contracting authority’s note: Form must be completed and attached to the bid, regardless the fact if you

submit bid individually or with a joint bid with subcontractors

DATA ABOUT THE SUBCONTRACTORS AND SUBCONTRACTOR'S AUTHORISATION (relating to direct payments)

FORM 3.1

| Name of the Subcontractor: |  |
| --- | --- |
| Address of the Subcontractor: |  |
| Legal representative of the Subcontractor: |  |
| VAT identification number: |  |
| Phone: |  |
| E-mail: |  |

| Part of the execution of the contract that will be performed by a Subcontractor |  |
| --- | --- |
| Description of work |  |
| Subcontractor share in total contract (%) |  |
| Value (EUR without VAT) |  |
| Time of completion |  |
| Place |  |

In accordance with paragraph 5 of Article 94 of PPA-3 we declare (tick):

|  | YES, we request direct payments by the Contracting Authority and give our consent, and on the basis of the fourth indent of the second paragraph of Article 94, we request that the Contracting Authority "JOŽEF STEFAN" INSTITUTE for the public contract , will settle our claims against the Bidder (name of the Bidder) instead of the Bidder directly to us, to our transaction account. |
| --- | --- |
|  | We do NOT require direct payments from the Contracting Authority. In this case, we will receive payment for the work performed by the selected Bidder. |

Subcontractors who authorize the Contracting Authority to make direct payments and above tick YES, by signing this statement give consent to the Contracting Authority to pay the Subcontractors' claims against the Bidder.

If the Bidder will carry out the contract with Subcontractors, it must provide a list of all the Subcontractors and the type of work that the Subcontractors will be undertaking, contact information and legal representatives of the Subcontractors and complete an ESPD form of Subcontractors in accordance with 79. Article of PPA-3 and accompanied by a request for direct payment of the Subcontractor if requested by the Subcontractor.

|  |  | The Subcontractor: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The Contracting Authority’s note: Form must be completed and attached to the bid, only whether you submit the joint bid as a consortium.

ESPD for all economic operators

European Single Procurement Document (ESPD) is a formal statement by the economic operator that the relevant grounds for exclusion do not apply, that the relevant selection criteria are fulfilled and that it will provide the relevant information as required by the contracting authority or contracting entity. The ESPD form shall also include a formal statement that the contracting authority may ask any tenderer at any moment during the procedure to submit all or part of the required certificates and supporting documents where this is necessary to ensure the proper conduct of the procedure. The entries in the ESPD form and/or the evidence submitted by the economic operator must be valid.

The economic operator imports XML.file prepared by the contracting authority to the ESPD file on the website: https://ejn.gov.si/espd and directly enters the requested data into it.

Where groups of economic operators (including temporary associations, economic operators on whose capacity is referred to by the economic operator and subcontractors) participate together in the procurement procedure, a separate ESPD must be given for each of the participating economic operators.

The bidder who submits the bid in the ePonudbe.si system uploads its ESPD in the section »ESPD-Bidder« section and places the other ESPDs in the »ESPD-Other Participants« section. A bidder who submits an offer in the ePonudbe.si system shall upload the electronically signed ESPD into xml. format or unsigned ESPD in xml. in the latter case, in accordance with the General Terms and Conditions of use of the ePonudbe.si information system, a legally binding document having the same validity as is signed shall be deemed to have been submitted.

For the other participating economic operators from the group, the bidder attaches signed ESPDs in the pdf. form or electronically signed xml. in the »ESPD – Other Participants« section.

SAMPLE CONTRACT

FORM 4

CONTRACT

concluded between

| The Client: | The Supplier: |
| --- | --- |
| JOŽEF STEFAN INSTITUTE Jamova cesta 39 1000 Ljubljana represented by the director , director | (company name) (address) represented by the director , director |
| VAT identification number: SI55560822 | VAT identification number: |

INTRODUCTORY PROVISIONS

Article 1

The contracting parties have established that, in line with the Public Procurement Act, the Client awarded the contract as the most cost-effective Bidder with regard to the realisation of the supply requirement expressed in the public tender published on the Public Procurement Portal of the Republic of Slovenia under the reference number ……………..

With this contract the Client and the Supplier agree on their mutual rights and obligations, necessary for the smooth operation of purchasing and selling.

SUBJECT OF THE CONTRACT

Article 2

Subject-matter of the contract is the

SUPPLY AND INSTALLATION OF AN INTEGRATED LOW-TEMPERATURE HIGH-MAGNETIC-FIELD MEASUREMENT SYSTEM FOR PHYSICAL PROPERTIES

specified, in detail, in the Quotation/pro-forma invoice with the ref. number ………..... as of ………….... that is an integral part of this contract.

An integral part of this contract the client's tender and the entire tender documentation, including all technical documentation.

VALUE OF THE CONTRACT

Article 3

The value of the above equipment is …….………….. . The price does not include the VAT.

DELIVERY TERMS

Article 4

The delivery and installation of the equipment that is the subject of this contract must be performed by …………………….. The Supplier must by this date fulfil all contractual obligations, including installation; qualification of the equipment or preparation of appropriate initial settings for the equipment; and the qualification of the equipment or verification of all the specifications of the equipment.

The Supplier shall provide, transportation including transport insurance of the equipment to the Client’s address. The costs for transportation including insurance is included in the price as set out in article 3 of this agreement. The Supplier shall notify the Client of the intended delivery by e-mail, in writing or by telephone, at least 3 working days prior to the delivery.

The Client shall sign the handover note after the establishing that the correctly delivered equipment has been properly installed and that the Supplier presented to the Client the warranty certificates.

The Supplier shall provide the delivery of the equipment with his its own professionally qualified staff or with the subcontractors, listed in tender documentation. If the Client finds that the work is carried out by the subcontractor not listed in the tender, he is entitled to cancel the contract. The Supplier reserves the right to verify any time subcontractor employees performing work. All employees are required to give credible information.

In case the Supplier does not deliver and install the equipment by the date in the first paragraph of this article, the Client has a right to withdraw from the order.

In case the Client suffers damages or costs due to the Supplier not fulfilling, incorrectly fulfilling the obligations set in this contract or fulfilling them with delay, the Supplier must reimburse all the additional costs that the Client has incurred due to the Supplier not fulfilling, incorrectly fulfilling or fulfilling the contractual obligations with delay.

The Supplier must, if called upon by the Client, together with the Client participate as a client in any potential disputes, triggered by third parties in the event of the Supplier not fulfilling, incorrectly fulfilling or fulfilling the contractual obligations with delay.

PAYMENT TERMS

Article 5

The Supplier shall issue separate invoices, addressed respectively to the Client and to each co-financier contributing to the purchase of the subject matter of the contract, as follows:

- Invoice for the the Client: Jožef Stefan Institute (hereinafter IJS), Jamova cesta 39, 1000 Ljubljana. The supplier will issue an invoice for 50% of the contractual amount upon delivery of the equipment.

- Invoice for the co-financier: University of Ljubljana, Faculty of Mathematics and Physics (hereinafter UL FMF), Jadranska ulica 19, 1000 Ljubljana. The supplier will issue an invoice in the amount of EUR 40.983,60 excluding VAT, technical acceptance.

- Invoice for the co-financier: Institute of Mathematics, Physics and Mechanics (hereinafter IMFM), Jadranska ulica 19, 1000 Ljubljana. The supplier will issue an invoice in the amount of EUR 16.393,44 excluding VAT, technical acceptance.

- Invoice for the Client: Jožef Stefan Institute (hereinafter IJS), Jamova cesta 39, Ljubljana, 1000 Ljubljana.  The supplier will issue an invoice in the amount corresponding to the remaining contract value, i.e. _________________EUR excluding VAT, technical acceptance.

The supplier must indicate the following on the invoice:

- The VAT ID number of their own company, and

- The VAT ID number of the client SI55560822 for IJS

- The VAT ID number of the co-financer SI55332862 for UL FMF

- The VAT ID number of the co-financer SI45597162 for IMFM.

The Supplier's invoice must refer to the contract number on the basis of which it is issued. This contract number must also appear on all accompanying documents submitted with the invoice.

The parties responsible for payment will settle the invoice within 30 days of receiving a correctly issued Invoice, paid directly to the Supplier.

The payment period begins on the day following receipt of the document that serves as the basis for payment. If the final day of the payment period falls on a public holiday as defined by labor law, the deadline is extended to the next working day.

The contractual obligations will be settled by payment to the Supplier's account:

IBAN: ___________________________________

Bank: ___________________________________

In the event of delayed payment, the Supplier is entitled to charge statutory default interest.

ACCEPTANCE

Article 6

The acceptance shall be performed on the basis of the Final Acceptance Certificate (APPENDIX 1), signed by the authorised representatives of both contracting parties, upon the proper delivery of the equipment of adequate quality and quantity.

The Supplier will submit to the Client the following documentation:

- Invoice

- Delivery note (with an exact list of delivered parts, data about the gross and net weight)

- Technical documentation and instructions for use

- Handover note – final acceptance certificate

- Warranty certificate

The documentation must be in Slovene or English.

SUPPLIER GUARENTEES

Article 7

The responsibility for keeping the equipment undamaged is transferred from the Supplier to the Client once the equipment is delivered to the Client.

The Supplier guarantees to the Client:

- that the equipment is new, not used and free of any actual and legal defects;

- that the supplied equipment functions perfectly and that it complies fully with all technical descriptions, characteristics and specifications, provided by the Client;

- that additional activities related to the supplied equipment (installation and training) are carried out professionally,

- that the Client shall acquire all the rights related to the equipment, while the Supplier shall duly perform all the obligations related to the equipment.

The Supplier provides the Client with a warranty to cover the quality of the equipment and its smooth operation at least to an extent equal to the extent of the warranty given by the producer of the equipment to the Supplier, i.e., ……. years following the date of the receipt of the equipment.

During the warranty period the Supplier shall provide for emergency services responding to the Client by E-mail or telephone within two working days after the receipt of a notification. If a repair service is needed, the Supplier shall provide it typically within 30 working days after the problem has been identified.

In the case of a defect identified in the minutes upon delivery, the warranty period begins on the day the defect is rectified. If the fault cannot be rectified, the warranty period begins on the day of receipt of the new goods.

ELIMINATION OF ERRORS AND SUPPLY OF SPARE PARTS

Article 8

The Supplier undertakes to eliminate the errors of the delivered goods within the warranty period providing uninterrupted service to its own costs, including transport costs to a specific location. The period of notice is considered to be the time when the message reaches to the supplier at the phone number or e-mail stated in this agreement, provided that it has been awarded by the Client and shall contain at least the necessary data for identification of the goods.

The supplier is obliged to eliminate defects that occur during the warranty period within 30 working days after the problem has been identified.

The Supplier commits to provide the spare parts and to provide service support against payment for another 5 (five) years from the end of the warranty period.

In the case of failure to fulfil obligations, the Supplier must reimburse any additional costs and damage to the Client, which would consequently suffer.

MAINTENANCE AND SERVICE OF THE QUIPMENT

Article 9

The maintenance of the equipment is carried out regularly, according to predetermined schedule in line with the manufacturer's instructions and recommendations to ensure safe operation, to reduce the possibility of failures and to ensure the continuous optimal operation of the equipment. It is performed during the warranty period and after the expiration of the warranty. After expiration of the warranty the bidder provides payable after-warranty support. The price includes material, work, daily allowances and travelling costs.

Service of the equipment is a repair service that is performed in case of equipment failure on the basis of the customer's order after the expiration of the warranty period. Service services are intended to establish such condition of the equipment in which the equipment performs its intended function. Service maintenance costs are calculated according to the actual time and material used. The Supplier informs the Client about at least approximate costs before service is performed.

The responsibility for keeping the equipment undamaged is transferred from the Supplier to the Client once the equipment is delivered to the Client.

The Supplier shall provide for emergency services responding to the Client within 7 (seven) working days after the receipt of a notification.

EARLY TERMINATION OF THE CONTRACT

Article 10

The Client may withdraw from the contract without obligation to the Supplier if the Supplier:

- does not comply with all the requirements of the client and does not meet them despite the warning of non-compliance

- increases prices during the validity of the contract,

- does not perform the subject of the contract in the agreed quality or within the agreed deadlines,

- does not fulfil all its obligations under the contract,

- in other cases and to the extent specified in this contract.

The Client will notify the Supplier of the withdrawal with a withdrawal notice sent by registered mail. Withdrawal from the contract is valid from the date of receipt of written notice of withdrawal if the client does not write in the notice a later date of termination of the contract. If the Supplier cannot be served with the withdrawal notice for any reason, the contract shall expire on the day the withdrawal notice was submitted to the post. In this case, the Supplier is not entitled to any compensation or any other claim.

The Supplier has the right to withdraw from this contract in the event of violation of the terms of the contract by the Client. In this case, the contract will expire when the Client receives a written notice of withdrawal, stating the reason for the assignment by the registered mail.

In the event of withdrawal, the parties shall be obliged to fulfil the obligation so far made as agreed prior to the termination.

TERMS FOR TERMINATION OF THE CONTRACT

Article 11

This contract is concluded under the termination terms that will be met if one of the following circumstances is fulfilled:

- if the Contracting Authority is informed that a competent court has established with a final decision that the Suppler or its Subcontractor has infringed applicable obligations in the area of mandatory environmental, social and labour law, or

- if the Contracting Authority is informed that the competent state authority established that the Supplier or its Subcontractor has at the time of the contract infringed at least two infringement in relation to:

- payment for work,

- working hours,

- on rest,

- the performance of work on the basis of civil law contracts, despite the existence of elements of an employment relationship or in connection with the employment of illegal workers and for which an offence has been imposed by a final decision or several final decisions of the competent authority.

If the Client becomes aware of a violation, he must notify the Supplier within ten days. The Supplier may, within the time limit set by the client, which may not be longer than 15 days, submit evidence that he has taken sufficient measures to prove his reliability despite the existence of violations. If there is a violation by the subcontractor, the Supplier can submit evidence within the same period that the subcontractor has taken sufficient measures to prove its reliability despite the existence of violations. If the Supplier has not submitted evidence for the subcontractor, or if there is evidence, but the Client considers that these measures are insufficient, the Supplier can replace the subcontractor within a period determined by the Client, which may not be longer than 15 days in accordance with Article 94 of this law. or takes over the part he has subcontracted to this subcontractor himself, if this replacement or takeover does not constitute a substantial change to the contract. If the Supplier has not submitted evidence for himself or the subcontractor, or if there is any, the Client assesses that these measures are insufficient, or if the Supplier does not take over the work himself or proposes a new subcontractor, or if the Client, in accordance with Article 94 of this law, timely proposes a new subcontractor refuses, the termination condition is implemented on the condition that there is at least six months between the client's knowledge of the breach and the expiry of the contract.

If the termination condition is met, the contract is considered to be terminated on the date of conclusion of the new public procurement contract, and the client must start a new public procurement procedure immediately, but no later than within 60 days of becoming aware of the violation.  If the Client does not start a new public procurement procedure within this period, the contract is considered to be terminated on the sixtieth day after becoming aware of the violation.

ANTI-CORRUPTION CLAUSE

Article 12

If it is found that, during the execution of the public tendering, on the basis of which this contract has been made, or during the execution of this contract, anyone acting in the name of – or on the account of – the Supplier, has promised, offered or given any undue advantage to the Client (the Contracting Authority), or those employees that had a decisive influence on the choice of the Supplier, in order:

- to win this public contract,

- to conclude this contract under more favourable conditions,

- to omit due diligence in the supervision of the contract implementation,

- to carry out any other act or omission causing damage to the Client or allowing the Client or its employees to obtain an undue advantage,

this contract will be void.

In the case of an alleged existence of the conditions from the first paragraph of this article or from a notification of the Commission for the Prevention of Corruption, or any other responsible body, relating to their alleged existence, the Client shall initiate the proceedings for invalidity of the contract from the first paragraph of this article, or any other measures, in line with the regulations of the Republic of Slovenia.

RESOLUTION OF CONFLICTS

Article 13

The contracting parties shall resolve any disagreement relating to the realisation of this contract by common consent in accordance with the contract. In the case that the contracting parties cannot settle a dispute by common consent, they shall abide by the decision of the competent court in Ljubljana.

ADMINISTRATORS OF THE CONTRACT

Article 14

The administrator of the contract on the part of the Client is: ______________________________________

The administrator of the contract on the part of the Supplier is: ___________________________________

It is not necessary to conclude an annex to the contract to change the administrators referred to in this article. Both parties shall be notified of the change in writing.

FINAL PROVISIONS

Article 15

The contracting parties shall determine all the amendments to this contract, after a previously reached agreement, with the annexes to this contract.

This contract comes into force on the day when it is signed by both contracting parties.

The contract is drawn up in two (2) identical copies so that each contracting party receives one copy.

| Client: |  | Supplier: |
| --- | --- | --- |
| #### Jožef Stefan Institute |  |  |
| Director: |  | Director: |
| Date: |  | Date: |

We herewith declare that the format of this sample contract is acceptable for the Bidder.

|  |  | The Bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

A STATEMENT WITH DETAILS ABOUT THE NATURAL AND LEGAL ENTITIES OWNED BY THE BIDDER including the partners and business entities that are considered to be associated with the Bidder in line with the provisions of the law governing companies

Appendix 1

Details about the Contracting Authority/the tender

| The Contracting Authority | Jožef Stefan INSTITUTE Jamova cesta 39, 1000 Ljubljana, Slovenia |
| --- | --- |
| The code of the tender | JN52/2025 |
| Subject of the tender | SUPPLY AND INSTALLATION OF AN INTEGRATED LOW-TEMPERATURE HIGH-MAGNETIC-FIELD MEASUREMENT SYSTEM FOR PHYSICAL PROPERTIES |

Details about the Bidder

| Name of the Bidder: |  |
| --- | --- |
| VAT identification number: |  |
| Legal representative of the Bidder: |  |

For the purpose of the 6th paragraph of Article 14 of the Integrity and Prevention of Corruption Act (the Official Gazette of the Republic of Slovenia, No.  69/11 – official consolidated text and 158/20), to ensure the transparency of the business and to mitigate corruption risks when concluding legal transactions in the capacity of the Bidder's representative in the process of awarding public contract, I hereby confirm

1.1. that are in the bidder's ownership participated following natural persons:

| No. | Name and surname | Permanent address | Share of ownership in % |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| … |  |  |  |

1.2. that are in the bidder's ownership participated following legal persons:

| No. | Name and registered office of the legal person | VAT identification number | Share of ownership in % |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| … |  |  |  |

the legal person being owned by the following natural persons:

| No. | Name and surname | Permanent address | Share of ownership in % |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| … |  |  |  |

1.3. that are in according of the Companies Act associated companies with the Bidder following economic operators:

| No. | Name and registered office of the economic operator | VAT identification number |
| --- | --- | --- |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |
| … |  |  |

is in mutual relationship with the following legal person:

| No. | Name and registered office of the legal person | Vat identification number | Associated as follows |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| … |  |  |  |

I hereby declare that, in the capacity of a natural person - participant in the Bidder's ownership, I stated:

- any natural person who owns through direct or indirect ownership at least 5% of shares or participates in the management rights, management or capital of the legal person with more than a 5% share or has the controlling position in the management of the legal person's funds;

- any natural person who indirectly provides or is providing funds to a legal person and is on such grounds given the possibility of exercising control, guiding or otherwise substantially influencing the decisions of the management or other administrative body of the legal person concerning financing and business operations.

By signing this statement, I hereby guarantee that there are no other natural or legal persons, dormant partners and companies, considered to be related companies pursuant to the provisions of the law regulating companies, in the entire ownership structure.

My signature on this statement warrants that the data provided are true and accurate and I am aware that the contract/the tender/the framework agreement shall be considered null and void, should my statement be false or data untrue. I undertake to inform the Contracting Authority of any changes to the data submitted.

|  |  | Name and surname of the legal representative: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the legal representative) |

The Contracting Authority’s note: The statement must be completed and signed by the Bidder, as well as by all individual members of a group of Bidders (partners) in the case of a joint bid, all subcontractors (if the bidder is executing the public contract with subcontractors), and all economic operators whose capacities the bidder is relying on.

final acceptance certificate

Appendix 2

| Name of the bidder: |  |
| --- | --- |
| Address: |  |
| VAT identification number: |  |
| Subject: | SUPPLY AND INSTALLATION OF AN INTEGRATED LOW-TEMPERATURE HIGH-MAGNETIC-FIELD MEASUREMENT SYSTEM FOR PHYSICAL PROPERTIES |
| Contract No.: |  |

| Following Delivery note No. |  | of |  | herewith we issue: |
| --- | --- | --- | --- | --- |

| FINAL ACCEPTANCE CERTIFICATE |  | No. |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | place: |  |  |  |
|  | date: |  |  |  |

| Type of purchasing: |  |  |  | Service |  |  |  | Material | X | Equipment |  |  |  |  | Construction |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Delivery method: | x |  | Supplier |  |  |  | JSI arrangement: |  |  |  |  | Other: |  |  |  |  |

| No. | Findings |  | YES |  | NO |  |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | All the ordered parts or services have been delivered. |  |  |  |  |  |
| 2 | The quality of the supplied parts or services complies with all the terms of the purchase agreement. |  |  |  |  |  |
| 3 | The required quantity of individual components or services has been supplied. |  |  |  |  |  |
| 4 | All the required or necessary documentation is included. |  |  |  |  |  |
| 5 | All the manuals including the instruction for use are included. |  |  |  |  |  |
| 6 | All the other required acceptance tests, as specified in Technical specifications have been carried out and are approved (QA) |  |  |  |  |  |
| According to the above findings, we confirm acceptance |  |  |  |  |  |  |

Additional comments:

Supplier shall receive a copy of the signed final acceptance certificate.

| The Supplier |  | Institut “Jožef Stefan” |
| --- | --- | --- |
| ((name and surname of the supplier’s representative) |  | (name and surname of the responsible for the equipment) |