Project title: Package 23 »Small-angle X-ray scattering (SAXS) beamline at Elettra-Sincrotrone Trieste«

Declaration of Compliance with the ContRacting Authority’s Minimum Technical Requirements and Specifications of the Offered Equipment

Name of the tenderer:

Address of the tenderer:

Subject of the public procurement: Procurement and delivery of an in-vacuum undulator system for integration into a synchrotron beamline

Designation of the public procurement: 302/6- VU-ŽL/25

| Is the tenderer also the manufacturer of the offered equipment? | YES ☐ | NO ☐* |
| --- | --- | --- |
| *If the answer is NO, enter the equipment manufacturer's name and address: |  |  |

INSTRUCTIONS:

- This Attachment No. 2 has to be completed and signed. The tenderer must mark whether the offered equipment is meeting the required technical specifications of the offered equipment according to all Contracting Authority’s requirements.

- The tenderer uploads this Attachment No. 2 in PDF format and additionally in Word format, if the form contains active web links to supporting documents, to the e-JN information system, section “Other attachments”.

We declare that the offered equipment fully meets the minimum technical requirements of the Contracting Authority for the subject matter of the public procurement entitled " Procurement and delivery of an in-vacuum undulator system for integration into a synchrotron beamline ", and that we are prepared to perform the contract according to the following specification:

|  | EQUIPMENT |  |  |  |
| --- | --- | --- | --- | --- |
| No. | REQUIREMENTS | OFFERED |  | DATASHEET |
| 1.1 1.2 1.3 1.4 1.5 1.6 1.7 1.8 1.9 1.10 1.11 1.12 1.13 1.14 1.15 1.16 1.17 1.18 1.19 1.20 1.21 | Main parameters, magnetic field characteristics and tolerances |  |  |  |
|  | The main parameters of the undulator are summarized in Table 1. Table 1: Main parameters of the undulator Name device; Beam Line; l0 (mm); KMAX; gmin (mm); gmax (mm); NPER; LMAG (mm); LTOT (mm). IVU17; mXRD; 17.2; 1.55; 5.2; > 30; 116; ~ 2000; 2800.00. The working curves of the undulator, based on the parameters indicated in Table 1, are shown in Figure 1. Figure 1: Angular flux density for IVU17 An essential condition is that the energy of the third harmonic photons, at the minimum gap and for an electron energy of 2.4 GeV, is equal to or less than 4400 eV (dotted line). | - YES | - NO |  |
|  | Magnetic structure |  |  |  |
|  | The system should include a complete magnetic structure composed of permanent magnets manufactured from sintered NdFeB powders with a remanent magnetic field sufficient to achieve the specified deflection parameters. The magnets must have an intrinsic coercivity (Hcj) not lower than 2400 kA/m (30 kOe). Enhancement treatments to increase coercivity, such as grain boundary diffusion, may be applied. All magnets should be coated with a protective layer compatible with ultra-high vacuum conditions and resistant to high radiation environments; titanium nitride (TiN) is required as the coating material. The structure must incorporate ferromagnetic poles and holders, ensuring high mechanical stability and performance. The permanent magnets shall be thermally stabilized (annealed) at a temperature of 110 °C, or at the maximum temperature required by the vacuum system bake-out procedure, to prevent irreversible demagnetization. The final heat treatment must be defined during the design review based on the validated vacuum conditioning plan. | - YES | - NO |  |
|  | Metallic foil for eddy current shielding |  |  |  |
|  | The magnetic structure shall be entirely covered with copper-nickel (Cu-Ni) foils, designed to provide an effective low electrical impedance path for image currents induced by the electron beam. These foils must be stretched over the magnetic assemblies to form a smooth surface facing the beam, free from scratches, indentations, or other surface irregularities. The continuous contact between the rear side of the foil and the underlying magnetic components to maintain optimal electrical continuity must be ensured. The shielding must support the operational requirements of the undulator in ultra-high vacuum and high-radiation environments. | - YES | - NO |  |
|  | Flexible transitions and connection flanges to the external vacuum chamber |  |  |  |
|  | The system shall include flexible transitions installed at both the entrance and exit of the in-vacuum device. These transitions must ensure continuous electrical conductivity for image currents generated by the circulating electron beam. They must also minimize impedance variations at the interface between the undulator and the adjacent vacuum chambers. Each transition must be equipped with connection flanges compatible with the external vacuum system. The design must accommodate mechanical tolerances while maintaining vacuum integrity and preserving beamline alignment. All components shall be suitable for ultra-high vacuum operation. | - YES | - NO |  |
|  | Photon absorber |  |  |  |
|  | The system shall include a photon absorber specifically designed to protect the exit flange of the in-vacuum device and any downstream components. The absorber must be installed downstream of the flexible transition and directly connected to the fixed part of the exit flange. The supplier is required to design a suitable cooling circuit for the absorber, considering the following beam and magnetic parameters: Electron beam current: 400 mA Electron beam energy: 2.4 GeV Bending magnet field: 0.7856 T The absorber must effectively dissipate the synchrotron radiation generated under these operating conditions and must be compatible with ultra-high vacuum requirements. The absorber design must prevent thermal damage or misalignment of adjacent structures during continuous operation. | - YES | - NO |  |
|  | Vacuum system, pumps, valves, and pressure sensors |  |  |  |
|  | The vacuum system shall ensure ultra-high vacuum conditions suitable for the operation of an in-vacuum undulator. Under standard operating conditions, the pressure inside the undulator vacuum chamber must not exceed 2·10–10 mbar, and the helium leak rate must be equal to or less than 1·10–10 mbar·L/s. The system must include: - All necessary vacuum pumps, including those integrated into the undulator housing, - Valves required for vacuum control and safety, - Pressure sensors for continuous monitoring of vacuum conditions. All components must be fully compatible with the required vacuum levels and must be integrated to ensure stable and reliable vacuum operation during the device's operational lifetime. The system must include all in-vacuum cables, controllers, and gauges necessary for complete functionality and control. Materials and surface finishes of internal vacuum components must be suitable for bake-out and free of contaminants. The vacuum system must include the following specific components and meet the associated performance standards: - Ion pumps in triode configuration with negative polarity, rated for ≥ 1·109 rads radiation and ≥ 200 °C temperature; equipped with high-voltage vacuum-compatible connectors such as Fischer-type or equivalent - NEG pumps with sintered elements, mounted on CF flanges (e.g., SAES CapaciTorr Z series or equivalent), to operate in synergy with ion pumps - Two Penning-type pressure gauges (e.g., Pfeiffer IKR 070 or equivalent) with suitable controllers (e.g., TPG500), resistant to radiation (≥1·109 rads) and heat (≥ 200 °C), with remote electronics and 30-meter radiation-hard cables - High-voltage cables, 30 meters in length, with Fischer-type connectors at both ends, capable of withstanding ≥ 1·109 rads and ≥ 200 °C One DN63CF flange fitted with an all-metal angle vacuum valve (e.g., VAT Series 57 or equivalent), capable of bake-out at ≥ 200 °C and ≥ 10,000 cycles | - YES | - NO |  |
|  | Alignment structure |  |  |  |
|  | The system shall include a support and alignment structure designed to securely hold the in-vacuum device during installation, operation, and maintenance. The structure must: - Be equipped with removable or liftable wheels to facilitate transportation and exact placement, - Provide stable mechanical support with structural integrity during vacuum operation and magnetic load conditions, - Allow for alignment of the undulator to the nominal beam height, which is specified as 1.3 meters above the floor level, - Be compatible with the defined space envelope and load limitations: - Maximum length (Z): 2800 mm - Maximum width (X): 1500 mm - Maximum height (Y): 2800 mm - Maximum weight: 7500 kg The structure must be manufactured to high mechanical exactness and allow for accurate and repeatable positioning within the accelerator or beamline environment. It shall also be compatible with lifting and handling procedures defined for heavy and sensitive scientific instrumentation. The total height of the fully assembled device, including any lifting slings or lifting bar, must not exceed 2.8 meters. If slings are used, the angle between them must not exceed 90°. If a sling bar is used, it must ensure a minimum vertical clearance of 0.5 meters from hook to lifting eye. If these constraints cannot be met, the supplier may propose a modular or partially dismountable design, under the condition that each module: - Meets individual height and weight limits (≤ 2800 mm height, ≤ 7500 kg total weight) - Is self-standing and safe to park without additional support - Can be safely lifted, handled, and assembled on site The lifting eyes must be positioned to maintain proper load balance during hoisting. If special lifting tools (e.g., sling bars) are required, they must be delivered with the equipment. | - YES | - NO |  |
|  | Lifting spreader |  |  |  |
|  | A dedicated lifting spreader shall be provided if required for the safe handling and installation of the in-vacuum device. The lifting device must be dimensioned and engineered specifically for the weight and geometry of the undulator assembly, considering a total weight of up to 7500 kg and maximum dimensions of 2800 mm × 1500 mm × 2800 mm (L × W × H). The spreader must ensure safe and balanced lifting without causing mechanical deformation or misalignment of sensitive components. It shall comply with applicable safety regulations for lifting equipment and be compatible with standard rigging and crane systems available at the installation site. | - YES | - NO |  |
|  | Electromechanical system for gap, taper, and offset adjustment |  |  |  |
|  | The device must be equipped with a fully integrated electromechanical positioning system capable of reproducible adjustments of the magnetic gap, taper, and offset. The required specifications are: Gap adjustment: Reproducibility: better than 1 μm Adjustment step: < 0.5 μm Adjustable speed range: 1 μm/s to 1 mm/s Mechanical end stops must be present at both extremes of the working range Taper adjustment: Adjustment range: ±50 μrad Step size: 1 μrad Reproducibility: better than 2 μrad Offset adjustment: Adjustment range: ±1 mm Step size: 1 μm Reproducibility: better than 5 μm The system must include absolute encoders. Encoders shall be placed in order to avoid radiation damage and beam-induced interference. Inclinometers must be installed on both magnetic beams to detect and protect against parallelism loss. The entire positioning system shall be operated via a real-time industrial automation platform capable of sub-micron motion control, using EtherCAT bus or an equivalent real-time fieldbus protocol. | - YES | - NO |  |
|  | Control cabinet and connection cables |  |  |  |
|  | The system shall include a complete control cabinet housing all required electronics for the actuation and monitoring of the in-vacuum device’s mechanical and vacuum systems. The cabinet must include: - Controllers and power supplies for all motion axes (gap, taper, offset) - Interfaces for all sensors and actuators, including temperature monitoring and alignment feedback - Integration of the motion control system based on TwinCAT 3 and EtherCAT bus, or equivalent high-performance industrial automation platform - All necessary wiring, cabling, and connector boxes, delivered as a complete and functional unit The cabinet must be pre-assembled, factory-tested, and labeled for integration at the site. All cables must be compatible with in-vacuum or high-radiation environments where required and include sufficient length and mechanical protection for installation. Electrical and wiring diagrams must be supplied as part of the documentation. | - YES | - NO |  |
|  | Field integral correction coils |  |  |  |
|  | The system must include field integral correction coils designed to allow fine-tuning and local correction of the magnetic field integrals along the beam path. These coils must be compatible with the control system and fully integrated into the device’s magnetic assembly. Their function is to compensate for magnetic field imperfections caused by manufacturing tolerances, mechanical deformations, or material inhomogeneities. The correction system shall allow precise adjustment to meet the required field quality for beam dynamics and spectral performance. All hardware and electrical interfaces necessary for operation and adjustment must be included, as well as provisions for integration into the control cabinet and software system. | - YES | - NO |  |
|  | Temperature monitoring sensors for the magnets and flexible transitions |  |  |  |
|  | The system shall be equipped with temperature monitoring sensors specifically installed on: - The magnetic structures (permanent magnets) - The flexible transitions These sensors must provide real-time temperature data to ensure stable thermal conditions during operation. The system must maintain the magnets at an operational temperature of 23 ± 1 °C, with a maximum heat dissipation of 10 W/m on both the upper and lower foil surfaces due to image currents. The sensors shall be fully integrated into the control system for continuous monitoring, with alarms or interlocks if deviations exceed acceptable thresholds. The temperature data shall be accessible locally and remotely via the user interface. Cooling circuit parameters required for proper thermal management: - Water temperature: 23 °C ± 1 °C - Conductivity: 5 μS/cm ± 3 μS/cm - Supply pressure: 3.5 bar - Max pressure drop: 2.5 bar All piping for the cooling system must comply with metric standards, and no pipe joints may be located inside the ultra-high vacuum (UHV) region. Materials used in contact with the cooling water must be compatible with the demineralized water properties specified in this section. Connectors used for cooling and vacuum assemblies shall be of Swagelok compression type or technically equivalent. The use of PTFE (Teflon) tape or any similar sealing material is strictly prohibited in all vacuum-related assemblies. | - YES | - NO |  |
|  | Local and remote control software and user interface |  |  |  |
|  | The system shall include fully developed control software capable of operating all mechanical, vacuum, and diagnostic functions of the in-vacuum device. The software must: - Be accessible via a local user interface and support remote operation and diagnostics - Allow for control and monitoring of gap, taper, and offset positioning, with integration of the following hardware performance parameters: - Gap reproducibility: better than 1 μm - Taper reproducibility: better than 2 μrad - Offset reproducibility: better than 5 μm - Adjustment step sizes: < 0.5 μm (gap), 1 μrad (taper), 1 μm (offset) - Use of absolute encoders for high-precision position feedback and inclinometers on both magnetic beams to detect loss of parallelism - Adjustable mechanical end stops must be present at both limits of the motion range - Interface with temperature sensors, magnetic field correction coils, vacuum monitoring, and motion encoders - Be compatible with the installed control hardware (e.g., Beckhoff TwinCAT 3 or equivalent), based on EtherCAT bus or equivalent real-time control protocols - Provide a graphical user interface (GUI) that enables intuitive operation, system status feedback, and parameter logging The control software must be delivered with all necessary licenses and documentation, and must support data acquisition, logging, and alarm functions. Remote access must allow full diagnostic capabilities during the warranty period, including performance monitoring and configuration verification. The motion control system shall include the following safety and communication features: - A PLC digital output to indicate “gap-is-open” status for gaps exceeding a programmable threshold - A PLC digital input to trigger emergency gap opening to a safe value - A PLC digital output to signal temperature excursions beyond preset limits - Remote control via Ethernet with DHCP and TCP/IP protocol - Compatibility with the TANGO control system architecture used in Elettra 2.0 - A high-speed data interface (e.g., UDP/IP or equivalent) for real-time synchronization with an external monochromator. Configurable switching between the standard and fast synchronization modes | - YES | - NO |  |
|  | Magnetic measurement data |  |  |  |
|  | The supplier shall provide complete magnetic measurement data for the in-vacuum device, including a summary table detailing all relevant magnetic parameters. The data must verify compliance with the specified magnetic performance, including: - The undulator shall be designed such that at minimum gap and 2.4 GeV electron beam energy, the third harmonic photon energy does not exceed 4400 eV - Magnetic field errors due to imperfections (e.g. material inhomogeneities, tolerances, deformations) must not exceed specified limits within the operational gap range Magnetic field errors caused by imperfections of various kinds (non-homogeneity of materials, machining tolerances, deformations, etc.) must not be exceeded in the indicated gap range given in Table 2: Table 2: Tolerances on magnetic errors Physical quantity; Units of measure; Limit value; Gap range where this value must not be exceeded. First Integrals (IX and IY); G×m; 0.1; gmin ÷ gmax. Integral seconds (I2X and I2Y); G×m2; 0.2; gmin ÷ gmax. Integrated quadrupole (normal and "skew"); G; 50; gmin ÷ gmax. Integrated sextuple (normal and "skew"); G/cm; 50; gmin ÷ gmax. Integrated octupole (normal and "skew"); G/cm2; 100; gmin ÷ gmax. Phase Error (rms); °; 2; gmin ÷ 15 mm. Appropriate correction coils will be part of the supply, integrated into the structure, capable of reducing the first and second integrals within the previous specifications if it is not possible to achieve this result with purely passive methods. All measurements must be performed using calibrated instrumentation, and results must be clearly documented in digital format. The summary shall include magnetic field maps, integral values, and any field correction data applied. The data must be delivered prior to shipment and included as part of the final documentation package. | - YES | - NO |  |
|  | Supply of manufacturing drawings |  |  |  |
|  | The supplier shall deliver complete manufacturing drawings of the in-vacuum device and all associated subsystems in CAD format. These drawings must include: - Full mechanical layouts of the magnetic structure, vacuum chamber, and support frame - Assembly drawings of the electromechanical positioning system, photon absorber, and vacuum interfaces - Schematics for cable routing, cooling lines, and sensor placement - Dimensional tolerances, materials used, and part references All drawings must be supplied in standard CAD formats (e.g., .dwg, .dxf, or .step) and be suitable for engineering review, maintenance planning, and future modification if required. The drawings must reflect the as-built state of the delivered equipment. All drawings must be included in the final documentation package when the undulator is delivered. | - YES | - NO |  |
|  | Operation and maintenance manuals, electrical diagrams, list of commercial components, CE certification for all the above systems |  |  |  |
|  | The supplier shall provide a complete documentation package, which must include: - Operation manual detailing start-up, standard use, safety instructions, and shutdown procedures - Maintenance manual specifying regular service intervals, component lifespan, preventive actions, and troubleshooting guidelines - Electrical diagrams showing full wiring schematics for all subsystems, including power distribution, control circuits, sensor integration, and safety interlocks - A list of commercial components, including model numbers, manufacturer references, and sourcing information for all non-custom parts - A valid and complete CE Declaration of Conformity, certifying that all components and subsystems of the device comply with the applicable EU safety, health, and environmental protection legislation All documentation must be delivered in English, in both printed and digital formats. The documentation must correspond to the final configuration of the delivered system and reflect the as-built and as-tested state. The documentation must be included in the final documentation package when the undulator is delivered. | - YES | - NO |  |
|  | Device bake-out procedure |  |  |  |
|  | The supplier shall provide a detailed bake-out procedure for the in-vacuum device to ensure proper degassing and preparation of internal surfaces prior to operation under ultra-high vacuum conditions. The procedure must include: - Step-by-step instructions for initiating and performing the bake-out process - Maximum allowable temperatures for each component (e.g., vacuum chamber walls, foils, sensors) - Recommended ramp rates for heating and cooling - Hold times at peak temperature to achieve full degassing - Monitoring guidelines to ensure system integrity during bake-out (e.g., pressure monitoring, thermal stability) The procedure must also include safety precautions, compatible heating methods, and environmental conditions required during the process. All system components must be designed and selected to withstand the specified bake-out conditions without degradation or vacuum performance loss. | - YES | - NO |  |
|  | Spare permanent magnets |  |  |  |
|  | The supplier shall deliver a set of spare permanent magnets identical in material, dimensions, magnetic grade, and coating to those used in the magnetic structure of the in-vacuum device. The magnets must be manufactured from sintered NdFeB with an intrinsic coercivity not lower than 2400 kA/m (30 kOe) and coated with titanium nitride (TiN) to ensure compatibility with ultra-high vacuum conditions and resistance to radiation. The quantity and configuration of the spare magnets must be sufficient to replace individual damaged or degraded magnets within the undulator structure. The supplier must include handling and storage instructions to preserve magnetic performance and prevent contamination or corrosion. Certification of magnetic properties for the spare units shall be included in the delivery documentation. | - YES | - NO |  |
|  | Warranty: 2 years or more must cover: |  |  |  |
|  | The replacement costs of the defective parts, - the corresponding labor expenses together with travel and living expenses of the provider’s personnel. All parts and services necessary to maintain optimum condition and a high availability of the instrument should be provided. Everything that is necessary for successful operation should be covered at no additional cost throughout the term of the warranty. This includes: - all spare parts needed, - all consumables needed including its exchange on-site performed by the provider’s expert, - unlimited service visits, - free software updates. | - YES | - NO |  |
|  | Delivery on the location defined by the Contracting Authority (INCOTERMS 2020 DDP) |  |  |  |
|  | Condition for the successful execution of the item and signing of the acceptance report: - the supply of the equipment by the provider on the location defined by the Contracting Authority, - introduction to equipment referred to items from 1.1 to 1.17, - submission of operating instructions and other technical documentation - all performance specifications (items from 1.1 to 1.17) must be demonstrated on-site defined by the client before signing technical acceptance documents. - CE certification is a requirement for acceptance | - YES | - NO |  |
|  | Project timeline and review milestones |  |  |  |
|  | The following project milestones must be respected: - Kick-off meeting within 1 month after contract signature - Preliminary Design Review (PDR) within 3 months - Final Design Review (FDR) within 6 months - Written progress reports every 2 months during contract execution At PDR, the supplier shall present design proposals and quality control procedures. At FDR, the supplier shall deliver validated simulations, 3D models, and detailed plans covering mechanical, thermal, magnetic, and vacuum aspects. | - YES | - NO |  |

| Place and date: | Stamp: | Authorized signature: |
| --- | --- | --- |

Project No P23-79 titled “Small-angle X-ray scattering (SAXS) beamline at Elettra-Sincrotrone Trieste ”, acquired under the Public call for (co-)funding of research equipment (Package 23, No 6316-9/2024-4), is co-funded by the Slovenian Research and Innovation Agency.

The tenderer must mark whether the offered equipment is meeting the required technical specifications of the offered equipment according to all Contracting Authority’s requirements.

The tenderer has to indicate the name of the attached supporting document. The tenderer may also provide a link to the manufacturer's website where the relevant technical specification is stated.