The contracting authority intends to procure laboratory instruments and associated accessories for research and soil analysis and environmental monitoring purposes. The required devices must meet the following minimum technical and functional requirements. All equivalent systems that fully comply with the outlined specifications and intended functions are acceptable.

LOT 1

1. Laboratory Equipment and Environmental Monitoring Sensors and Systems Specifications

1.1 Water Potential Meter

Functional Description

Instrument for laboratory-based measurement of water potential using the chilled mirror dew point method, suitable for plant and agricultural research.

Minimum Technical Requirements

- Measurement method: Chilled mirror dew point technique

- Measurement range: 0 to –300 MPa

- Measurement accuracy:

  - ±0.05 MPa for range 0 to –5 MPa

  - ±1% of reading for range –5 to –300 MPa

- Operating temperature range: 15–40 °C

- Temperature measurement:

  - Accuracy: ±0.2 °C

  - Resolution: up to 0.1 °C

- Sample reading time:

  - Soil samples: 5 – 15 minutes

  - Plant samples: up to 20 minutes

- Sensor types: Chilled mirror dew point sensor and infrared temperature sensor

- Sample cup capacity: up to 15 mL

- Data communication: RS-232A serial interface with USB converter included

- Environmental operating conditions: 5 – 40 °C

Delivery Requirements

- Quantity: 2 units, supplied with necessary cups, standards, cables, software, and user documentation

1.2 Soil Water Retention Curve and Unsaturated Hydraulic Conductivity Measurement System (Evaporation Method)

Functional Description

System designed for determining soil water retention curves and unsaturated hydraulic conductivity using the evaporation method, based on DIN 11274 and Wind/Schindler theory.

Minimum Technical Requirements

- Measurement of soil water tension at two depths using tensiometers

- Continuous sample weight measurement

- Pressure transducer range: +0.3 to –100 kPa (optionally extendable to –400 kPa)

- Pressure accuracy: ±0.1 kPa (auto-zero function)

- Pressure resolution: up to 0.001 kPa

- Temperature sensor range: –20 to +60 °C

- Temperature measurement accuracy: ±0.2 K; resolution: up to 0.01 °C

- Supported sample volumes: up to 100 mL and up to 250 mL

- Default data logging interval: up to 10 minutes

- Sensor unit: Made of POM, splash waterproof (IP65), resistant to pH 3–10 chemicals

- Tensiometer construction: Ceramic tip, acrylic glass shaft, length options up to 24 mm and up to 49 mm

- Multi-sensor support: Operation of up to 20 sensor units per USB interface

- Balance specification: Capacity up to 2200 g, readout up to 0.01 g, internally calibrated

Delivery Requirements

- 8 complete systems, each including:

  - Sensor unit

  - Balance with required specifications

Additional Accessories/Parts

- Refilling equipment:

  - 2 refilling kits, each including a vacuum pump (achievable vacuum up to 8 mbar, up to 20 L/min), stainless steel stand with approx.1000 mL vacuum bottle, vacuum gauge, and adapters for simultaneous refilling of up to 4 tensiometers and 1 sensor unit

  - 4 beaker-mounts stainless steel stands with refilling adapters for up to 16 tensiometer shafts and up to 4 sensor units

- Soil sampling tools:

  - 3 hammering holders for 250 mL rings (approx. 300 mm length, approx. 0.9 kg, with hand protection)

  - 2 shock-free hammers (approx. 2 kg, up to Ø 70 mm)

  - 1 hammering holder for 100 mL rings (approx. 300 mm length, with hand protection)

  - 1 universal adapter set for compatibility with 250 mL ring holders

- Sampling rings:

  - 71 stainless steel rings (250 mL) with low-friction polished surface, engraved ID and tare weight, plus 2 caps per ring, Approx. dimensions: outer Ø: 84 mm; Inner Ø: 80 mm; Height: 50 mm

  - 50 stainless steel rings (100 mL) with low-friction polished surface, engraved ID and tare weight, plus 2 caps per ring, Approx. dimensions: outer Ø: 53 mm; Inner Ø: 50 mm; Height: 51 mm

1.3 Automated System for Saturated Hydraulic Conductivity Measurement

Functional Description

Automated laboratory system for measuring saturated hydraulic conductivity (Ksat) with precise control of water flow and high-resolution pressure sensing.

Minimum Technical Requirements

- Ksat measurement range: 0.01 – 5000 cm/day

- Porous plate conductivity: ≥ 14,000 cm/day

- Pressure sensor accuracy: ±1 Pa

- Temperature sensor accuracy: ±0.2 °C

- Sampling ring dimensions: Volume 250 mL, approx. height 50 mm, internal diameter approx. 80 mm (optional compatibility with 100 mL rings)

- Measurement accuracy: Up to 2% under constant conditions

- USB data communication, compatible with Windows 10 or newer

- Power supply: 110–240 V AC, 50/60 Hz

Delivery Requirements

- 2 complete systems with all required accessories and software

1.4 Laboratory Thermal Conductivity Measurement Device

Functional Description

Device based on transient line heat source method (ASTM D5334) for automated soil thermal conductivity measurements.

Minimum Technical Requirements

- Measurement range: 1.3 – 10.3 W/m·K

- Measurement accuracy: ±5%

- Operating temperature range: –40 °C to +85 °C

- Temperature measurement:

  - Accuracy: ±0.1 °C

  - Resolution: up to 0.01 °C

- Heater current range: 20 – 200 mA with 10 µA resolution

- Heater power output: 0.5 – 20 W/m with 20 mW/m resolution

- Power supply: 12 V / 1 A

- USB data interface; compatible with Windows 10 or newer

Delivery Requirements

- 1 system with thermal conductivity sensor needle, pre-drilled sampling rings, and required accessories

1.5 Automated Soil Particle Size Distribution Analyzer

Functional Description

Automated analyzer for soil particle size distribution based on sedimentation principles, compliant with DIN 18123 and ISO 11277.

Minimum Technical Requirements

- Particle size detection range: 2 – 63 µm

- Measurement error: ±1% of mass fraction

- Pressure measurement:

  - Accuracy: ±1 Pa

  - Resolution: ±0.1 Pa

- Data logging interval: up to 10 seconds

- Typical sample mass: 25–50 g per 1-liter suspension

- USB powered; compatible with Windows 10 or newer

- Measurement accuracy not affected by temperature fluctuations within ±1.5 °C

Delivery Requirements

- 4 systems with sedimentation cylinders, temperature equilibration cylinders, and accessories

1.6 Portable Dual-Head Infiltrometer

Functional Description

Portable system for unattended field measurement of saturated hydraulic conductivity (Kfs) using a dual-head infiltration method.

Minimum Technical Requirements

- Pressure head control range: 0 – 40 cm

- Maintained water level: up to 5 cm above soil surface during measurement

- Infiltration rate range: 0.0038 – 115 cm/h

- Measurement accuracy: ±5% of reading

- Operating temperature range: 0 – 50 °C

- Rechargeable power supply with USB data handling and charging capability

Delivery Requirements

- 1 complete system with insertion rings, collapsible water tanks, tools, and accessories

#### 1.7 Volumetric Water Content Sensor

##### Functional Description

Sensor for precise measurement of volumetric water content (VWC) in soil and soilless substrates, with optional compatibility for low-disturbance borehole installation.

##### Minimum Technical Requirements

- VWC range:

  - Mineral soils: 0.00–0.64 m³/m³

  - Soilless media: 0.00–0.70 m³/m³

- Dielectric permittivity range: 1–80

- Resolution: up to 0.0010 m³/m³

- Accuracy:

  - Mineral soils: ±0.03 m³/m³ (solution EC < 8 dS/m)

  - Soilless media: ±0.05 m³/m³ (solution EC < 8 dS/m)

  - Medium-specific calibration: ±0.01–0.02 m³/m³

- Dielectric measurement frequency: approx. 70 MHz

- Measurement volume: approx. 430 mL

- Operating temperature: –40 to +60 °C

- Compatible with borehole installation tool up to 2 m depth

Delivery Requirements

- 34 units

#### 1.8 Volumetric Water Content, Temperature, and Electrical Conductivity Sensor

##### Functional Description

Combined sensor for simultaneous measurement of volumetric water content, temperature, and electrical conductivity, with optional borehole installation capability.

##### Minimum Technical Requirements

- VWC range:

  - Mineral soils: 0.00–0.70 m³/m³

  - Soilless media: 0.00–1.00 m³/m³

- Dielectric permittivity range: 1–80

- Resolution: up to 0.0010 m³/m³

- Accuracy:

  - Generic calibration: ±0.03 m³/m³ (solution EC < 8,000 μS/cm)

  - Medium-specific calibration: ±0.01–0.02 m³/m³

- Temperature measurement:

  - Range: –40 to +60 °C

  - Resolution: approx. 0.1 °C

  - Accuracy: ±0.5 °C (–40 to 0 °C), ±0.3 °C (0 to +60 °C)

- Dielectric measurement frequency: approx. 70 MHz

- Bulk EC range: 0–20,000 μS/cm; Resolution: up to 1 μS/cm; Accuracy: ±(5% + 10 μS/cm) to ±8%

- Measurement volume: approx. 1010 mL

- Compatible with borehole installation tool up to 2 m depth

Delivery Requirements

- 16 units

#### 1.9 Matric Potential and Temperature Sensor

##### Functional Description

Sensor for measuring soil matric potential and temperature, utilizing a porous substrate and dielectric moisture sensing with integrated data communication capabilities.

Minimum Technical Requirements

- Matric potential range: 0 to –100,000 kPa

- Resolution: up to 0.1 kPa

- Accuracy: ±(10% of reading + 2 kPa) from –100 to –5 kPa

- Temperature measurement:

  - Range: –40 to +60 °C

  - Resolution: approx. 0.1 °C

  - Accuracy: ±1.0 °C

- Dielectric measurement frequency: approx. 70 MHz

- Output protocols: DDI serial, SDI-12

- Data logger compatibility: 3.6–15 VDC powered systems with serial or SDI-12 communication

- Operating temperature: –40 to +60 °C

- Compatible with borehole installation tool up to 2 m depth

Delivery Requirements

- 16 units

#### 1.10 Borehole Installation Tool

##### Functional Description

Tool for low-disturbance installation of soil sensors up to 2 m depth using a hand-augered borehole, suitable for field deployment.

##### Minimum Technical Requirements

- Compatible with volumetric water content, temperature, EC, and matric potential sensors

- Installation depth: up to 2 m (extendable with optional 1 m extensions)

- Includes:

  - Approx.. 10 cm auger

  - Sensor carriers

  - Soil surface flange

  - Depth selection clamp

Delivery Requirements

- 1 unit

#### 1.11 Field Tensiometer with Integrated Temperature Sensor

##### Functional Description

Field tensiometer for direct measurement of soil water potential and temperature, designed for plug-and-play operation with field data loggers, refillable without extraction.

##### Minimum Technical Requirements

- Water potential range: –85 to +50 kPa

- Resolution: up to 0.0012 kPa

- Accuracy: ±0.15 kPa

- Temperature measurement:

  - Range: –30 to +60 °C

  - Resolution: approx. 0.01 °C

  - Accuracy: ±0.3 °C (–0 to +40 °C)

- Communication protocols: DDI serial, SDI-12, RS-485 (4-wire)

- Dimensions:

  - Diameter: approx. 2.5 cm

  - Length options: 40, 80 cm

- Installation angle range: 10°–80° (downward or upward)

- Materials: Al2O3 ceramic, PMMA shaft, POM GF corpus, stainless steel refilling tubes

- Operating temperature: –30 to +50 °C

- Includes:

  - Approx. 60 mL syringe with refilling faucet

  - Installation auger with hammering head and depth gauge (diameter approx. 25 mm, length approx. 125 cm, blade approx. 50 cm)

Delivery Requirements

- 7 units (3 units 80 cm length, 4 units 40 cm length)

#### 1.12 Soil Water Content and Temperature Profile Probe

##### Functional Description

Multi-depth soil moisture and temperature probe with integrated sensors at multiple depths, installed with dedicated tools for accurate placement.

##### Minimum Technical Requirements

- VWC range:

  - Mineral soils: 0.00–0.70 m³/m³

  - Soilless media: 0.00–1.00 m³/m³

- Resolution: up to 0.001 m³/m³

- Accuracy:

  - Generic calibration: ±0.05 m³/m³ (solution EC < 8 dS/m)

  - Medium-specific calibration: ±0.02–0.03 m³/m³

- Temperature measurement:

  - Range: –20 to +60 °C

  - Resolution: approx. 0.03 °C

  - Accuracy: ±0.5 °C (–20 to 0 °C), ±0.25 °C (0 to +60 °C)

- Dielectric measurement frequency: approx. 70 MHz

- Measurement depths: approx. 15, 30, 45, and 60 cm

- Compatible with installation and removal tools

Delivery Requirements

- 7 units

#### 1.13 Installation and Removal Tool for Soil Profile Probes

##### Functional Description

Tool set for installing and extracting soil moisture profile probes with minimal disturbance.

##### Minimum Technical Requirements

- Includes:

  - Slide hammer

  - Auger

  - Eyebolt

  - Carabiner

  - Tensioning strap

  - Lever tool (pile puller)

  - Base

Delivery Requirements

- 1 unit

#### 1.14 Compact All-in-One Weather Station

##### Functional Description

Compact weather station measuring 12 meteorological variables with integrated leveling, plug-and-play operation, and MODBUS communication option.

##### Minimum Technical Requirements

- Measures: Solar radiation, precipitation, relative humidity, air temperature, humidity sensor temperature, vapor pressure, barometric pressure, wind speed, wind gust, wind direction, tilt, lightning strike count and average distance

- Integrated solar radiation sensor: 0–1750 W/m², ±5% accuracy

- Precipitation range: 0–1,500 mm/h, ±5% accuracy

- Wind speed range: 0–60 m/s, accuracy ±0.3 m/s or up to 6%

- Temperature range: –50 to +60 °C, accuracy ±0.2 °C

- With all necessary installation equipment, including bird spikes and masts

- Communication: MODBUS option, integrated leveling bracket

Delivery Requirements

- 10 units

#### 1.15 Bluetooth Sensor Read-Out Device

##### Functional Description

Handheld Bluetooth interface for spot-checking and configuring field sensors.

##### Minimum Technical Requirements

- Compatible with analog, pulse, and digital sensors

- Supports Bluetooth pairing with mobile devices (iOS/Android)

- Displays sensor readings and metadata

- Firmware update support

- Does not log or store data

Delivery Requirements

- 2 units

#### 1.16 Cloud-Enabled Data Logger with Remote Management

##### Functional Description

Field data logger with GPS, Bluetooth setup, solar power, and cloud-based remote management with integrated metadata features.

##### Minimum Technical Requirements

- Logging interval: 5 min to 12 h

- Reporting interval: 1 h to 24 h

- Connectivity: USB, Bluetooth 5.2, 3G/4G cellular, GPS

- Data storage: approx. 8 MB flash memory (40,000+ records)

- Battery options: 6 AA NiMH or alkaline, solar recharging

- Operating temperature: –40 to +60 °C

- Enclosure: min. IP56 rated, impact and weather-resistant

- Sensor ports: approx. 6 universal input ports (analog, digital, pulse)

- Cloud features:

  - Remote device management

  - Data delivery in near-real time

  - Firmware over-the-air updates

  - Model applications (e.g., GDD, ET, chill hours…)

  - Soil-specific calibration management

  - Notifications and alerts

  - Unlimited user sharing

  - API access for data integration

Delivery Requirements

- 16 complete units

- 120 months cloud subscription per 1 unit

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided for all relevant equipment, where applicable

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 2

2. Soil Physical Properties, Sampling and Sensor systems Equipment

The contracting authority intends to procure laboratory instruments and associated accessories for research and soil analysis and environmental monitoring purposes. The required devices must meet the following minimum technical and functional requirements. All equivalent systems that fully comply with the outlined specifications and intended functions are acceptable.

2.1 Pipette Apparatus (Table Model)

Functional Description: Laboratory device for determining the fine fraction content in soil samples by sedimentation in accordance with NEN 5753 and ISO/DIS 11277. Allows simultaneous analysis of at least 6 samples and includes a heated glass tank with stirrer.

Minimum Technical Requirements:

- Table-mounted unit with frame and graduated runner

- Heating element with thermostat and stirrer

- Glass tank, 100–110 L capacity

- Glass sample cylinders (1000 mL), with stoppers and weighting rings

- Maximum immersion depth of pipette: up to 350 mm

- Includes pipette upper/lower sections, pipette balloon, sodium hexametaphosphate

Quantity Required

- 2 complete sets with all accessories

2.2 Granular Composition Test Set

Functional Description

Electromagnetic sieve shaker with stainless steel mesh sieves for particle size distribution of soil samples, suitable for dry and wet sieving.

Minimum Technical Requirements

- Electromagnetic shaker for up to 8 sieves (approx Ø 200 mm, approx 50 mm height)

- Timer: 1–99 min; adjustable vibration: 0,20–3 mm

- Includes up to 8 sieves with apertures: 0.063–8.00 mm

- Includes lids and receivers for dry/wet sieving

Quantity Required

- 1 complete set with all accessories

2.3 Wet Sieving Apparatus

Functional Description

Equipment for aggregate stability analysis using wet sieving method.

Minimum Technical Requirements

- Shaking machine suitable for up to 8 sieves

- Power: 100–240 V AC adapter

- Stainless steel cans (up to Ø 64 × 45 mm): up to 16 pcs

- Sieve cans (up to Ø 39 × 39 mm, approx. 0.250 mm mesh): up to 8 pcs

Quantity Required

- 1 complete set with all accessories

2.4 Tension Infiltrometer Set

Functional Description

Device for measuring unsaturated hydraulic properties in field soil using a tension-controlled infiltration plate.

Minimum Technical Requirements:

- Infiltration disc diameter: up to 20 cm

- Adjustable tension settings; membrane pressure: approx. 30 cm H₂O

- Disc mountable under or separate from water tower

- Reservoir diameter ≥ 5 cm; length ≥ 80 cm; bubbling tower ≥ 2.5 cm

- Accuracy: up to 1 mm

Quantity Required

- 1 complete set with all accessories

2.5 Edelman Auger

Functional Description

One-piece auger for sandy/clayey soils, forged high-tensile steel

Quantity Required

- 1 piece

2.6 Handle for Auger

Functional Description

T-handle, up to 60 cm length, with bayonet connection and coupling sleeve

Quantity Required

- 1 piece

2.7 Bayonet Connection Bottom Part for Edelman Auger

Functional Descriptio

Forged steel bottom part, approx. Ø 5 cm, with rapid-connect coupling

Quantity Required

- 1 piece

2.8 Digital Soil Colour Identification Tool

Functional Description

A portable device designed for accurate and quick soil color assessment in the field. It uses the Munsell Color System, which has been established and widely adopted for over 55 years to support natural resource management. The device includes all colors from the Munsell Soil Color Book and allows voice note recording and automatic library updates via application. The device is intended for soil scientists, geologists, and environmental professionals for accurate soil classification. It is suitable for field use under various weather conditions.

Minimum Technical Requirements

- Light source: approx. 25 LEDs (8 visible wavelengths, 1 UV)

- LED lifespan: >10 million flashes

- Display: Color TFT, approx. 45 mm

- Measurement time: up to 1.8 seconds

- Calibration: Integrated white calibration slider

- Operating temperature: 0 °C to 40 °C

- Storage temperature: -20 °C to 60 °C

- Battery: minimum 3.7 V lithium-ion rechargeable or better

- Battery life: minim 5 hours continuous use, 100 hours standby, or better

Additional Features

- Voice note recording capability

- USB connectivity for software updates

- Compact and rugged design for field use

- Palette software for library synchronization

Quantity Required:

- 1 unit with all necessary accessories

2.9 Sensor

Functional Description

Rugged SDI-12 sensor for fixed installation and long-term measurement of soil moisture and temperature in all soil types.

Minimum Technical Requirements

- Measurement principle: Coaxial impedance dielectric reflectometry

- Moisture accuracy: ±0.01 WFV (0–100% range)

- Temperature accuracy: ±0.3 °C (–10 to +60 °C)

- Pin length: approx. 57 mm

- Communication: SDI-12 v1.2; 3-wire, 1200 baud

- Power supply: 9–20 VDC; current: <1 mA idle / 10 mA active

Quantity Required

- 4 sensors with approx. 7.5 m cable

- 4 sensors with approx. 15 m cable

2.10 Modem

Functional Description

Wireless field modem with cloud connection, supporting SDI-12 sensors and equipped with onboard barometer, USB data plug, and 6 sensor ports.

Minimum Technical Requirements

  - Data transmission: via web portal/email/SMS

  - Up to 4 sensor ports for SDI-12 sensors

  - Battery-powered with external DC option (5.5–18 V)

  - Alarm types: high-low, fluctuation, low battery

  - Up to IP67 environmental protection

  - Cloud features: Real-time data access, modeling tools, API, multi-user sharing

  - At least 5 years of cloud data visualisation subscription

  - SIM card included

  - Accessories: Above-ground housing, solar power kit, mounting set, spare battery,

  - Training and installation support

Quantity Required

  - 2 complete systems with all accessories

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 3

3. Gas Exchange and Environmental Monitoring Systems

3.1 Integrated Soil CO₂ Flux Chamber

Functional Description

Soil respiration chamber designed for integration with the gas exchange and environmental control system for accurate CO₂ flux measurement from soil surfaces in field and laboratory environments.

Minimum Technical Requirements

- Full integration with the Portable Photosynthesis and Fluorescence System sensor head

- Designed for direct measurement of CO₂ flux with system’s analyzers and environmental controls

- Includes sealing hardware and airflow management for precise soil flux measurement

- Air Temperature Thermistor: Operating Range: -20 to 45 °C; Accuracy: ±0.5 °C over 0 °C to 70 °C

- 20 cm Soil CO2 Flux Chamber

- Data output can be formatted for SoilFluxPro™ Software

Quantity: 1 unit

3.2 Field Evapotranspiration and Energy Balance Monitoring Node

Functional Description

Compact integrated system for direct field measurement of evapotranspiration (ET), latent heat (LE), and sensible heat (H) using eddy covariance, with embedded data analytics, power system, and cellular data connectivity.

Minimum Technical Requirements

- Measures ET, LE, H, VPD, Ta, RH, Pa, AH, SVP, Td with flux data output approx. every 30 minutes

- Uses high-frequency wind/water vapor eddy covariance method at approx. 10 Hz

- Power consumption ≤ 1.5 W (≤ 26 W at startup), solar and battery compatible

- SDI-12 communication with LTE CAT1 cellular connectivity; EU region compatible

- Operating range: –20 °C to +50 °C

- Compact design of ET sensor: approx. 58 × 17.5 × 7.6 cm, ~1.4 kg, up to IP54 rating

- Adjustable telescoping mast up to 5 m height

- Cloud-based platform supporting data cleaning, gap filling, analytics

- With integration with Internet of Environment system, including solar panel, battery, cellular modem, GPS, SDI-12 ports

- Subscription to cloud data service for advanced analytics and data storage for years 2–5 (4 years).

Quantity: 1 complete unit

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 4

4. Soil Bulk Density Measurement Equipment

4.1 Soil Bulk Density Meter

Functional Description

Field instrument for in-situ measurement of soil bulk density using gamma-ray backscatter technology with integrated GPS, moisture sensor, and onboard data logging. Designed for direct field use without requiring soil pits or laboratory analysis.

Minimum Technical Requirements:

- Measurement depth: Up to 80 cm, using a borehole diameter of 30 mm

- Measurement accuracy:  up to 0.01 g/cm³

- Operating temperature: –20 °C to +85 °C

- Environmental rating: up to IP68 for dust and water protection

- Integrated sensors: Soil density, moisture, GPS

- Internal storage: approx. 16 GB flash memory

- Battery: Rechargeable Li-ion, minimum 8 hours operating time per charge

- Interfaces: RS-232, Ethernet (up to 100 Mbps), Wi-Fi (2.4/5 GHz), Bluetooth

- Data formats: JSON, NMEA, CSV

- Included software: Dedicated software for data processing and export

- Construction: Housing made of titanium, aluminium, and high-strength plastic

- Accessories: Gauge auger, hammer, cabling, mobile device for onboard logging

Additional Requirements:

- Measurement technique must allow operation without special permits (low-activity gamma source exempt from regulation or equivalent method)

- Integrated moisture sensor to convert bulk density to dry density

- Integrated GPS to geolocate measurements

Quantity Required

- 1 unit with all accessories

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 5

5. Essential Laboratory Devices for Precision Weighing and Mixing

  - Precision Laboratory Balances

Functional Description

High-precision laboratory balances designed for analytical and routine laboratory tasks. Suitable for soil physics laboratories and related research facilities, as well as laboratories requiring high-precision analytical equipment. Equipped with advanced weighing technology that combines the precision of monolithic weighing cells with the modularity and repairability of traditional sensor systems, providing high accuracy while allowing individual components to be replaced or serviced if damaged.

Minimum Technical Requirements

- Capacity: At least 2 kg and 6 kg models

- Readability: up to 0.01 g for both models

- Repeatability:

  - 2200 g model: ≤0.01 g

  - 6200 g model: ≤0.01 g

- Minimum weight:

  - 2200 g model: up to 2 g

  - 6200 g model: up to 2 g

- Linearity:

  - 2200 g model: ±10 mg

  - 6200 g model: ±15 mg

- Pan size: approx. 186 × 186 mm

- Response time: up to 1.5 s

- Automatic self-calibration with option for external calibration using standard weights

- Interfaces: RS232 and USB Host for PC/printer connection

- HID function for direct PC transfer

- Integrated clock for date/time stamping of measurements and calibration

- Built-in statistical, percentage, counting, density applications

- Checkweighing (upper/lower limit comparisons)

- Stability indicator and automatic repeatability test

- Multi-language support with LCD backlit display

- Theft protection: mechanical and PIN-code security

- User-defined unit conversion factors

- Compliance-ready with IQ/OQ/PQ qualification support

Quantity Required

- 1 unit with 2200 g capacity

- 1 unit with 6200 g capacity

  - Compact Laboratory Shaker

Functional Description

Programmable orbital shaker designed for reproducible shaking tasks in laboratories, suitable for use with flasks, funnels, and various vessels. Features a robust design with adjustable speed and stroke, and optional rack systems for flexible applications.

Minimum Technical Requirements

- Shaking motion: Orbital

- Platform size: approx. 400 × 300 mm

- Maximum load capacity: approx. 15 kg

- Speed range: 30–420 rpm, adjustable

- Stroke: approx. 17 mm

- Runtime: Programmable or continuous operation

- Touch display with programmable control

- Vessel capacity:

  - 100 mL flasks: 20 pcs

  - 250 mL flasks: 12 pcs

  - 500 mL flasks: 6 pcs

  - 1000 mL flasks: 4 pcs

  - Separating funnels (250 mL): 1 pc

- Ambient temperature: 5 °C to 50 °C

- Relative humidity: up to 85%

- Protection rating: approx. IP21

- Heat emission: approx. 7 W

- Electrical supply: 230 V or 115 V, 50/60 Hz

Additional Accessories:

- Rack System: Platform approx. 400 × 275 mm with at least 3 clamping strips and rubber mat, suitable for upright vessels.

Quantity Required

- 1 complete system, including shaker and accessories

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 6

6. Computing Equipment

6.1 Portable Computers (Laptops)

Functional Description

High-performance business laptops designed for professional laboratory and office use, featuring AI-enhanced performance, robust security, and collaboration features. Optimized for compatibility with IT environments and remote management.

Minimum Technical Requirements

- Processor: Equivalent to Intel Core Ultra 5-125H (0.7–4.5 GHz, 14 cores/18 threads) or better

- Operating System: Windows 11 Pro (64-bit, EN/SLO) or better

- Memory: At least 16 GB DDR5 5600 MHz or better

- Storage: At least 1 TB NVMe SSD (Gen4) or better

- Graphics: Equivalent to NVIDIA GeForce RTX 2050 or better

- Display: Equivalent to 16" WUXGA (1920 × 1200), non-glare, 300 nits, 45% NTSC or better

- Connectivity: Wi-Fi, Bluetooth, Ethernet

- Ports: 2 × USB-C, 2 × USB-A, HDMI, RJ-45, audio combo

- Smart card reader

- Security:Equivalent to TPM 2.0, fingerprint reader

- Camera: Equivalent to 5MP IR camera

- Audio: Equivalent to Poly Studio dual stereo speakers, dual-array microphones or better

- Keyboard: Spill-resistant, backlit with numeric keypad (Slovene layout)

- Battery: Equivalent to 3-cell 56 Wh Li-ion or better

Additional Requirements

- ESET Security Antivirus (4-year license)

- Microsoft 365 Business Basic (4-year license)

Quantity Required

- 5 units

6.2 Monitors

Functional Description

High-resolution professional monitors suitable for laboratory and office work, ensuring sharp visuals, ergonomic adjustments, and extensive connectivity.

Minimum Technical Requirements

- Screen size: At least 27" (68.58 cm) or more

- Resolution: Equivalent to QHD 2560 × 1440 (16:9) or better

- Panel: Equivalent to IPS, anti-glare or better

- Refresh rate: At least 100 Hz or better

- Brightness: approx. 350 cd/m²

- Contrast ratio: approx. 1500:1

- Color: approx. 16.7 million colors, 99% sRGB

- Response time: 2 - 5 ms

- Viewing angle: approx. 178° (H/V)

- Connectivity: USB-C, HDMI, USB-B, USB-A, DisplayPort

- Ergonomics: Height adjustable, tilt, swivel, rotation

- VESA mount: approx. 100 × 100 mm

- Energy efficiency: Class A – E

Quantity Required

- 3 units

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided

- Availability of spare parts and software updates for a minimum of 3 years

LOT 7

7. Laboratory Equipment Specifications for Public Tender

The contracting authority intends to procure laboratory instruments and associated accessories for research and laboratory use. The required devices must meet the following minimum technical and functional requirements. All equivalent systems that fully comply with the outlined specifications and intended functions are acceptable.

7.1 Steam Sterilizer (Autoclave)

Functional Description

Vertical steam sterilizer designed for sterilizing laboratory instruments, glassware, porous materials, infectious waste, and liquid media. The unit must include advanced control features, user programming capability, integrated safety systems, and comply with relevant EU standards.

Minimum Technical Requirements

- Supporting frame made of stainless steel

- Pressure chamber constructed of stainless steel

- Vertical configuration

- Chamber working volume: approx. 95 L

- Maximum power consumption: up to 9 kW; 400 VAC/16 A

- Integrated steam generator with automatic demineralized water filling

- Capability to create and store at least 50 user-defined programs

- Adjustable sterilization temperature up to 140 ºC

- Adjustable sterilization time

- Adjustable drying time

- Multi-level parameter programming secured by password protection

- Automatic door opening option at the end of the cycle

- Automatic door locking and sealing without manual screwing

- Safety mechanism preventing door opening in case of malfunction

- Safety mechanism preventing door opening if chamber temperature exceeds 80 ºC

- Overpressure safety system to prevent door opening during pressurized conditions

- Movable temperature probe in chamber

- Rapid water cooling of the chamber for liquid sterilization programs

- Chamber heating suitable for sterilization of glassware, instruments, and porous materials

- Radial fan in chamber ensuring uniform temperature distribution during sterilization and faster cooling during media cooling phase

- Multiple pre-vacuum and vacuum drying with additional chamber heating

- Vacuum air removal system using motor pump with water ring seal

- User-friendly operation via colour touchscreen (HMI), minimum size 5.7”

- Mobile app (iOS and/or Android) providing real-time monitoring of autoclave status and program sequence

- Built-in infectious filter for sterilization of infectious waste with automatic in-cycle sterilization controlled by temperature probe

- Condensate sterilization during infectious waste autoclaving cycles

- Internal data storage for at least 10 years

- USB data transfer for PDF and/or CSV reports

- Ethernet connectivity

- Integrated printer for cycle parameters and program phase reporting

- Option for scheduled program start by date/time

- Integrated condensate cooler (max. 70 ºC)

- Two (2) stainless steel mesh baskets

- One (1) stainless steel solid container with top-edge perforations

- Delivered with compatible silent air compressor for compressed air supply

- CE certification

Delivery Requirements

- Quantity Required: 1 unit, delivered with all specified accessories and installation-ready.

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 8

8. Sampler for Field-Based Biological Monitoring

  - Multi-Vial Cyclone Sampler

Functional Description

Field-deployable air sampling device designed for the simultaneous collection of airborne particles into multiple vials for downstream DNA and ELISA analyses. Suitable for continuous field monitoring and integration with weather stations.

Minimum Technical Requirements:

- Capable of collecting airborne samples directly into approx. 1.5 mL vials suitable for molecular (DNA) and immunological (ELISA) analysis

- Multi-vial collection system supporting a minimum of 8 vials per sampling session

- Adjustable sampling intervals and programmable timer functions

- Air intake rate: Minimum 10 L/min

- Battery-powered for field use with sufficient runtime for extended sampling campaigns

- Weatherproof design for reliable outdoor operation in diverse conditions

- Simple tool-free replacement of vials

- Low-maintenance design with minimal consumable requirements

Delivery Requirements:

- Quantity Required: 1 complete unit with all accessories for immediate field deployment

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 9

9. Soil Analysis and Monitoring Equipment

9.1 Laboratory Lysimeter System

Functional Description

A laboratory lysimeter system designed to investigate nutrient and solute transport processes under controlled conditions simulating various soil moisture and precipitation scenarios. The system allows for the study of sediment impacts (surface-applied or incorporated) on soil physical, chemical, and hydrological properties, with adjustable rainfall regimes and long-term monitoring of water and solute transport.

Minimum Technical Requirements (per Lysimeter):

-	Lysimeter Cylinder:

- Diameter: approx. 80 cm (minimizing edge effects, realistic infiltration patterns)

- Height: approx. 60 cm

- Material: Stainless steel (equivalent to quality 1.4301, cold rolled, or better)

- Bottom plate: Approx. diameter 0.88 m, thickness 6 mm, stainless steel

- Watertight bottom sealing with screw mounting set

- Up to10 sensor ports (approx. Ø 75 mm) with watertight flanges and complete mounting set for probes

-	Weighing System:

- 3 load cells per lysimeter, nominal load up tp 300 kg per cell, accuracy class up to C6

- Accuracy: ≤ 50 g (equivalent to ≤ 0.1 mm precipitation/ET for approx. 800 mm diameter)

- Recording interval: up to 1 minute

- Weighing transducer and complete mounting set included

-	Seepage Water Measurement:

- Closed tipping counter, resolution: up to 100 ml per tip

- Sample bottle: approx. 50 ml capacity

-	Sensors:

- 3 matrix potential + soil temperature sensors:

- Range: –5 to –100,000 kPa

- Accuracy: ±(10% of reading + 2 kPa) from –100 to –5 kPa

- Resolution: op to 0.1 kPa; Frequency: approx. 70 MHz

- Temperature: Range –40 to +60 °C; Accuracy ±1 K; Resolution up to 0.1 K

- 3 water content + bulk EC + soil temperature sensors:

- Water content: Range 0–100%, Accuracy ±3% FS, Resolution up to 0.1%

- Bulk EC: Range 0–3,000 µS/cm; Pore EC: 0–55,000 µS/cm

- Accuracy: ±25 µS/cm (0–1,000 µS/cm); ±2.5% (1,000–2,500 µS/cm); ±5% (2,500–3,000 µS/cm)

- Temperature: Range –30 to +55 °C; Accuracy ±0.25 K (5–35 °C), ±0.5 K (–15 to +55 °C)

Facility Requirements

-	Control cabinet and wiring documentation

-	Data logger: equivalent to Campbell Scientific CR6 (programmed for full system control) or better

-	COM-Server for remote operation and support

-	Power supply: 12/24 V system (220 V input provided)

-	Ethernet connectivity for remote access

-	Installation, wiring, test run, and commissioning included

Soil Core Requirements

-	Monolithic undisturbed soil core extraction (B horizon field monoliths)

-	Upper soil section: 20–30 cm Ap horizon soil, with or without approx. 5 cm sediment layer (surface or mixed)

Installation Requirements

-	Installation of sensors and weighing components

-	Installation of lysimeter system in laboratory

-	Full wiring, integration, and test operation

-	Training for users in operation, maintenance, data retrieval, and troubleshooting

Monitoring & Sampling Capabilities

-	Continuous monitoring of soil moisture, matric potential, EC, and temperature at depths (10, 30, 50 cm)

-	Automated weighing and water balance recording

-	Leachate sampling for chemical/physical analysis (metals, nutrients, EC, pH, suspended particles)

Delivery Requirements

-	Quantity Required: 6 lysimeters with 1 central measuring cabinet

-	Includes all installation, commissioning, test operation, and user training

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 10

10. System for Moisture Sorption Characterization

10.1 Moisture Sorption Analyzer (DVS & DDI Methods)

Functional Description

Automated analyzer for determining moisture sorption characteristics using Dynamic Vapor Sorption (DVS) and Dynamic Dew Point Isotherm (DDI) methods. The system must offer the ability to measure at least:

- Soil Moisture retention and release behavior,

- Sorption hysteresis (quantified through adsorption and desorption curves),

- Soil-specific surface area,

- Static and dynamic isotherm curves.

Minimum Technical Requirements

- Measurement range: –10 MPa to –475 MPa

- Water potential measurement accuracy: ±1 MPa or ±1% of reading

- Sample weight capacity: 500 – 5000 mg

- Weight measurement accuracy: ±0.1 mg

- Temperature control range: 15 – 60 °C

- Operating conditions: 15 – 40 °C

- Sample cup capacity: up to 10 cc

- Water reservoir volume: up to 20 mL

- Universal power supply: 110–220 V AC, 50/60 Hz

- USB interface; compatible with Windows 10 or newer

Delivery Requirements

- 1 system with calibration weights, desiccant tubes, sample cups, filters, cleaning kit, and verification standards

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 12 months

- Technical support available in English or Slovene

- Complete user manuals and installation guides and EU declaration of conformity (CE declaration) provided for all relevant equipment, where applicable

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide technical support for all delivered equipment

- The technical support shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Response time for technical support shall not exceed 48 hours from the time of notification. Technical support may be provided on-site or remotely as required. In addition, the supplier must provide continuous technical support, including phone and email assistance, for a minimum of 5 years after delivery

LOT 11

11.  Analytical Measuring Systems (AMS)

11.1 Analytical Measuring Systems (AMS)

Functional Specification

In the initial phase, the system will be used for conducting research on the deposition of plant protection products (PPPs) in hop fields and apple orchards, and later also in other agricultural crops.

In recent years, increasingly strict environmental regulations and a shortage of available PPPs have led to growing challenges in crop production. Proper application and deposition of PPPs are crucial, as deposition studies allow for even distribution of the spray across the entire plant with minimal negative environmental impact.

The Optomax Image Analyser enables us to assess deposition by analysing droplet patterns – including droplet count, droplet size, and percentage of surface coverage. Based on the evaluation of these parameters, we will be able to:

- provide recommendations for optimal spraying parameters,

- evaluate new nozzle types, and

- develop optimal nozzle configurations to recommend to end users.

This will help reduce water usage for spraying as well as PPP consumption.

Additional potential applications of the system

- Detection and analysis of plant parts damaged by diseases and pests (e.g. feeding damage), including pest and spore counting on leaf surfaces

- In plant breeding: measuring and analysing the assimilative surface area of plants, including leaf morphology, pollen measurement, and stomatal analysis

Minimum Technical Requirements

- New camera (Basler camera featuring GigE interface and similar, fully compatible with Windows 10 (64-bit)

- Supply, configuration, and testing of the new camera (Basler GigE camera and similar)

- Included accessories: Power supply, cabling, Ethernet cable, network interface card

- Compatibility with existing C-mount lens

- Development and testing of a software interface for the new camera compatible with Windows 10

- Test environment setup and validation of system performance and reliability

- Appropriate software for new camera compatibility

- Desktop PC specifications:

  - CPU: Intel i5 or higher

  - RAM: 16 GB minimum

  - Storage: 512 GB SSD or larger

  - Monitor, keyboard, and mouse included

- Included technical documentation

Delivery Requirements:

- Quantity Required: 1 complete unit with all hardware, software, accessories, and documentation included

Example:

Image Analysis System with Basler GigE camera and APA software or equivalent

GENERAL REQUIREMENTS

General Warranty and After-Sales Support/Training Requirements for All Devices:

- Minimum warranty of 6 months (unless otherwise stated)

- Certified service technicians available in Slovenia

- Technical support available in English or Slovene

- Complete user manuals and installation guides provided

- Availability of spare parts and software updates for a minimum of 5 years

- The supplier shall provide user training for all delivered equipment

- The training shall cover: Proper operation and handling of the equipment, Safety procedures and maintenance requirements, Basic troubleshooting and calibration (where applicable), Use of any associated software or data interfaces, Data handling, exporting, and integration (where applicable)

- Training may be provided on-site or remotely. Additionally, the supplier shall provide technical support (e.g., via phone or email) for a minimum of 2 years after delivery.

Note on Reference Devices: The devices must meet all the technical and functional specifications listed in this document. Any system commercially available that meets these specifications will be accepted. Examples of known devices that meet these specifications are provided for reference only and do not imply preference or exclusivity.