PUBLIC-TENDER DOCUMENTATION

OPEN PROCEDURE

40. člen

(type of procedure)

»SUPPLY OF A MULTI-BENCH HARDWARE-SHARING CONFOCAL IMAGING PLATFORM FOR HUNTING OF FUNCTIONAL CELLULAR EVENTS«

(subject of the public procurement)

JN01/2026

(internal reference number)

January 2026

CONTENTS

Tender documentation consists of:

1. INSTRUCTIONS FOR THE BIDDERS

2. TECHNICAL DOCUMENTATION OF THE PUBLIC TENDER

3. TENDER DOCUMENTATION

# 1.  INSTRUCTIONS FOR THE BIDDERS

# 1.1 THE CONTRACTING AUTHORITY

The contracting authority Jožef Stefan Institute, Jamova cesta 39, 1000 Ljubljana, tax number SI55560822, registration number 5051606, bank account: 01100-6030344242 registered with the Public Payment Administration (PPA) Ljubljana, has published on the public procurement portals information about the public tender »SUPPLY OF A MULTI-BENCH HARDWARE-SHARING CONFOCAL IMAGING PLATFORM FOR HUNTING OF FUNCTIONAL CELLULAR EVENTS«.

The Contracting Authority hereby invites all interested tenderers to submit a tender in compliance with the requirements from the tender documentation. The tender has to be prepared in line with the tender documentation, fulfilling all the conditions for the participation in this public tender.

# 1.2 GENERAL INFORMATION

| Subject: | SUPPLY OF A MULTI-BENCH HARDWARE-SHARING CONFOCAL IMAGING PLATFORM FOR HUNTING OF FUNCTIONAL CELLULAR EVENTS |
| --- | --- |
| Short description: | The subject of the public tender is the supply and installation of a multi-bench hardware-sharing confocal imaging platform for hunting of functional cellular events dedicated to the process of knowledge discovery – beyond the classic confocal microscopes or today’s emerging high-throughput imaging instruments that can generate lots of images or at best large quantities of data. Together with the analysis algorithms for quantification of biology-informed cellular events and their connection into the proposed mechanism of action, we aim to offer a unique platform for streamlined research and development, accessible to every lab anytime. With the microscope that is subject of this public tender we aim to increase accessibility of the microscopes by reduced price per measuring station as well as create a unique platform that will enable development of a proof of a more general hardware sharing concept (sharing all lasers and detectors) that has the potential to reduce the price-per-measuring station even further. |
| Internal reference: | ZN25_15402/MT |
| Type of the bidding procedure: | Open procedure in accordance with Article 40 of PPA-3 |
| Type of purchasing: | Equipment |
| Division into lots: | The public tender is not divided into lots. |
| Variants: | Variants are not permited. |
| The language of the bids: | The Bidders must submit their bids in Slovene or English. |
| Requested validity of the tender: | 30 April 2026 |
| Deadline for queries: | 10 February 2026 until 16:00 |
| Deadline for submission of bids: | 17 February 2026 until 10:00 |
| Date and time of the public opening: | 17 February 2026 until 12:00 |

Detailed specification of the tender can be found from the technical specifications - point 2 of these instructions.

# 1.3 implementation of the tender

The Contracting Authority carries out the process of awarding a public contract on the basis of the applicable law and by-laws governing public procurement, in accordance with the applicable legislation governing the field of public finance and the field that is the subject of public procurement.

An entity that is subject to an absolute business ban based on the provision of Article 35 of the current Act on Integrity and Prevention of Corruption may not participate in the execution of a public contract.

If the tenderer submits a false statement or gives untrue information about the stated facts, this results in the inadequacy of the offer or the nullity of the contract.

# 1.4 access to the tender documentation

The forms from the tender documentation (TD) are available at the Public-Procurement Portal (https://www.enarocanje.si/). The tender documentation can be accessed free of charge

# 1.5 preparation, submission and delivery of the bids

Any individual or legal entity registered for the service that is the subject of this tender can submit a bid to this tender.

To prepare their bids, the Bidders must give all the details required with the attached specification. The subject of the bid must meet at least the minimum technical requirements specified in the technical specification, which is an integral part of this tender documentation.

The tenderer submits the tender documentation in such a way that, after registration or logging into the ePonudbe system at the address: http://eponudbe.si, for the relevant public procurement, he selects the option "Login and submit a tender", which opens the page for preparing the tender.

Detailed instructions are available in the General Terms and Conditions of the ePonudbe portal and published at address https://eponudbe.si/en/General-Terms-of-Use   (Article 5)

# 1.6 the bidding price

All prices must include all costs for the realisation of the relevant public contract. The final price must include all its constituent elements. VAT must be stated separately (for the bidders from Slovenia). If the price is not expressed in EUR, the contracting authority will convert the offered price into EUR at the exchange rate of the Bank of Slovenia on the day of bid opening.

# 1.7 a joint bid

A bid may be submitted by a group of contractual partners. This group shall submit a Partnership Agreement stipulating how the public contract awarded within this public-procurement procedure will be executed, specifying, in particular, the leading partner authorized to assume, from the Contracting Authority, the obligations, instructions, and possibly also the payments on behalf of, and for the account of, all the partners, as well as the share and type of services/goods to be provided by each individual partner. The Agreement should clearly stipulate that, in relation to the Contracting Authority, all the partners are collectively and individually liable for the realisation of the entire commitment and each part of it, in accordance with paragraph 3 of Article 81 of the PPA-3.

All the partners in a group have to meet the conditions for the recognition of their competencies with respect to their legal statuses (they must meet all the conditions in section 1.13).

All the bidders in the joint bid must fill out the ESPD individually and provide all the required information.

Form 1 "Proforma invoice" is submitted by all providers participating in the joint offer together (one form signed by at least one of the providers participating in the joint offer).

In the case of a joint offer, legal entities should submit in form "ESPD" and provide all participants in the joint offer (Section A Part II ESPD). Providers participating in the joint offer may also specify one of the legal entities with whom the client will communicate until a decision on the order is made, otherwise the client will address all documents to all providers participating in the joint offer.

# 1.8 a bid involving subcontractors

In a subcontracting relationship the main Contractor transfers the execution of the whole or a part of the awarded public contract to a third party, namely, a Subcontractor. The definition of a Subcontractor is given in point 1 of Article 94 of the PPA-3. The Bidder can perform the award of the contract alone or together with Subcontractor(s).

If the tenderer offers the execution of the contract by Subcontractor(s), it is required to:

- provide the list of all the Subcontractors (each Subcontractor must meet all the requirements under point 1.13) and the type of work that the Subcontractor will be undertaking,

- provide the list of contact information and the legal representatives of the Subcontractors,

- complete Declaration of compliance with the conditions for the Economic Operator for each Subcontractor,

- complete FORM 2.2 – The Subcontractor’s authorisation relating to direct payments made by the Contracting Authority to the Subcontractor(s) and consent, if Subcontractor required.

The main Contractor will be required during the realization of the public tender to inform about any changes to the existing information specified in the preceding paragraph and to send information about new Subcontractors, that are planned to be subsequently integrated into carrying out the work no later than 5 days after the change. In the case of the involvement of new Subcontractors the main Contractor must together with the notice provide the information and documents, specified in the second, third and fourth indents of the preceding paragraph and meet all the requirements under point 1.13

In the event that the Subcontractor in accordance with and in such a way specified in the second and third paragraph of Article 94 of the PPA-3 requires direct payment, a direct payment to the Subcontractor shall be deemed to be obligatory in accordance with this Act and bound by an obligation of the Contracting Authority and the main contractor.

Where the tenderer intends to carry out the contract with a Subcontractor, which requires a direct payment in accordance with this Article:

- the main contractor in the contract shall authorize the Contracting Authority, on the basis of an approved invoice or situation by the main contractor to pay directly to the Subcontractor,

- the Subcontractor submits the consent on the basis of which the Contracting Authority instead of the tenderer settles the Subcontractor's claim against the tenderer,

- the main contractor to its invoice attaches an invoice of a Subcontractor, which was previously approved.

If a direct payment to the Subcontractor is not required, the Contracting Authority of the main contractor will require that not later than 60 days from the payment of the final invoice the main contractor sends a written statement and a written statement from the Subcontractor that the Subcontractor has received payment for the supplied goods directly related to the subject of the contract.

If the main contractor does not comply with Article 94 of PPA-3, the Contracting Authority shall submit a proposal to the National Review Commission about the introduction of a minor offence procedure referred to in point 2 of the first paragraph of Article 112 of the PPA-3.

In the event that the tenderer acts with subcontractors, he indicates in FORM 2 that he is applying with subcontractors and at the same time submits FORM 2.2 - Information and consent of the subcontractor for direct payments.

# 1.9 EXAMINATION AND EVALUATION OF THE BIDS

Based on the conditions and criteria specified in the tender documentation, the Contracting Authority will select the Bidder with whom he will enter into a contract for the supply of goods/performance of the service.

Before awarding the contract, the Contracting Authority will verify the existence and content of data or other information indicated in the tender of the tenderer to which it has decided to award the contract. The Contracting Authority will perform an examination and evaluation in line with Article 89 of the PPA-3.

# 1.10 submission and delivery of the bid

Bidders must submit their bids through the IT ePonudbe.si system at the web address http://eponudbe.si/en

Before submitting their bids, Bidders must register at the web address http://eponudbe.si/ . If the Bidder is already registered in the IT ePonudbe.si, they must sign in at the same address.

A user, that is in the ePonudbe.si information system and authorised to submit tenders for an Economic Operator, submits a tender by clicking the »Submit tender« button. The ePonudbe.si information system records the identity and the time of the submission of tender. The user by submitting the tender expresses and states its will to submit a binding tender in the name of the Economic Operator (Article 18 of the Obligations Code, Official Gazette of the Republic of Slovenia, no. 97/07 – official consolidated version, 64/16 – odl. US and 20/18 – OROZ631). With its submission, the tender is binding for the time stated in the tender except if the user withdraws or changes it before the time for submitting the tenders has expired.

A bid is deemed to be submitted on time if the Contracting Authority receives it via the ePonudbe.si system http://eponudbe.si/en  by 17 February 2026 at 10.00 AM. A bid is deemed submitted if it is marked with the status "SUBMITTED" (“ODDANO”) in the IT ePonudbe.si system.

Bidders may withdraw or change their bids by the deadline for submission of the bids. Withdrawal of the bid can be performed on the ePonudbe.si system in the profile of the Bidder that has submitted the bid to the system, when the bid is submitted, with the optional statement about the reason for the withdrawal, the change of the bid is performed within the same system first by withdrawal of already submitted bid, and afterwards with submission of a new bid before expiry of the deadline.

It is not possible to submit any bid after the expiry of the deadline for the submission of bids.

The page for submitting an e-bid in this e-procurement procedure can be accessed here http://eponudbe.si/en.

Instructions for foreign bidders:

We kindly ask foreign providers to start the process at least three days before the deadline. Please follow the instructions https://eponudbe.si/en/General-Terms-of-Use (Article 5)

# 1.11  THE PUBLIC OPENING OF THE BIDS

Bids shall be opened automatically in the IT ePonudbe.si system on 17 February 2026 and will start at 12.00 PM at the web address http://eponudbe.si/en.

The opening takes place in the following way: at the hour set for the public opening of bids, the IT ePonudbe.si system automatically displays data on the Bidder and variants, if requested or allowed, and enables access to the PDF file that the Bidder uploaded in the ePonudbe.si system under the »Proforma Invoice« (»Predračun«) section. The public display automatically ends after 60 minutes. The Bidders that submitted bids have these data available in the IT ePonudbe.si system under the "Minutes on the bid opening" (»Zapisnik o odpiranju ponudb«) section.

# 1.12 QUERIES RELATING TO THE TENDER DOCUMENTATION

Communication with the Bidders concerning queries about the tender and the preparation of a bid shall be organised through the Public-Procurement Portal of the Official Gazette of the Republic of Slovenia: www.enarocanje.si.

Queries about the tender documentation or any other query about the call for tender placed through the Public-Procurement Portal by 16:00 on 10 February 2026 shall be considered as timely. After this time, the Contracting Authority shall not answer any queries about the tender.

Information provided by the Contracting Authority or through the procurement portal shall be, in accordance with the paragraph 2 of Article 67 of PPA-3 understood as an amendment, supplement or clarification regarding the award of the contract, if the content of this additional information results in changing this documentation, or if the explanation eliminates the ambiguity of the statements in this document or the public announcement.

# 1.13 GROUND FOR EXCLUSION OF THE BID

The economic operator shall verify compliance with the conditions by submitting the completed FORM 3.

| P1: Reasons related to criminal convictions | The proof |
| --- | --- |
| The Economic Operator must be excluded by the Contracting Authority from participation in a public procurement procedure in the case it is found that the Economic Operator itself or any person who is a member of its administrative, management or supervisory body or has powers of representation, decision or control therein has been the subject of a conviction by final judgment which has the elements of the criminal offenses that are defined in the first paragraph of Article 75 of the PPA-3 In the case that the Economic Operator is in the situation referred to in the above paragraph, in accordance with paragraph 9 of Article 75 of the PPA-3, the Contracting Authority may submit evidence that it has taken sufficient measures to prove its reliability despite the existence of reasons for exclusion. Payment or commitment to pay compensation for all damage caused by a criminal act or violation, active cooperation with investigative authorities for the complete clarification of facts and circumstances, and the adoption of concrete technical, organizational, and personnel measures appropriate to prevent further criminal acts or violations are considered sufficient measures. When evaluating the measures taken by the economic entity, the client takes into account the seriousness and special circumstances of the criminal act or violation. If the client assesses that the evidence submitted by the business entity is sufficient, the business entity shall not be excluded from the public procurement procedure, regardless of the first paragraph of Article 75 of the Public Procurement Act-3. If the client assesses that the measures are not sufficient, it sends the justification for such a decision to the economic entity. | FORM 3 – Completed Declaration of compliance with the conditions for the Economic Operator for all Economic Operators in the procedure. |
| P2: Reasons related to the payment of taxes or contributions for social security: The proof |  |
| The Economic Operator must be excluded by the Contracting Authority from participation in a public procurement procedure in the case that it does not fulfil the obligations related to taxes and other monetary non-fiscal obligations pursuant to the law regulating the financial administration, collected by the tax authority in line with the provisions of the country in which it has its our registered office or in the Member State of the Contracting Authority and have no outstanding liabilities as on the day of submitting the bid worth EUR 50 or more. On the day of submitting the bid, the Economic Operator must submit all accounts of tax deductions for the incomes arising from employment relationships for the period of the last 5 years. | FORM 3 – Completed Declaration of compliance with the conditions for the Economic Operator for all Economic Operators in the procedure |
| P3: National provision – Register of economic operators with negative references The proof |  |
| The Economic Operator must be excluded by the Contracting Authority from participation in a public procurement procedure in the case it is on the deadline for the submission of bids excluded from the public award procedures due to being included in the record of entities with negative references, as defined in point a) of the fourth paragraph of Article 75 of the PPA-3 | FORM 3 – Completed Declaration of compliance with the conditions for the Economic Operator for all Economic Operators in the procedure |
| P4: National provision – offence related to payment for work | The proof |
| The Economic Operator must be excluded by the Contracting Authority from participating in a public procurement procedure it has been in the 3 years prior to the expiry of the deadline for submitting the bids fined twice for an offence related to the payment for work, working hours, on rest, on the performance of work on the basis of civil law contracts, despite the existence of elements of an employment relationship or in connection with the employment of illegal workers by way of a final decision of the competent authority of the Republic of Slovenia or another Member State or a third country. An economic entity that is in one of the positions from the above paragraph can submit evidence to the client by the deadline for submission of applications or offers at the latest that it has taken sufficient measures to prove its reliability despite the existence of reasons for exclusion. Payment or commitment to pay compensation for all damage caused by a criminal act or violation, active cooperation with investigative authorities for the complete clarification of facts and circumstances, and the adoption of concrete technical, organizational, and personnel measures appropriate to prevent further criminal acts or violations are considered sufficient measures. When evaluating the measures taken by the economic entity, the client takes into account the seriousness and special circumstances of the criminal act or violation. If the client assesses that the evidence submitted by the economic entity is sufficient, the economic entity shall not be excluded from the public procurement procedure, regardless of point b) of the fourth paragraph of Article 75 of the Public Procurement Act-3. If the client assesses that the measures are not sufficient, it sends the justification for such a decision to the economic entity. | FORM 3 – Completed Declaration of compliance with the conditions for the Economic Operator for all Economic Operators in the procedure |

In all cases 1 copy of a proof is sufficient, even when a particular proof is mentioned several times!

If the country where the bidder is established cannot issue the proofs, or documents, required by the tender documentation, the bidder can, in line with Article 42 of PPA-2, submit an appropriate sworn statement given by the bidder or by a witness in the presence of a judicial or administrative authority, a notary, or an authorised body representing the commercial entities in the bidder’s country.

# 1.14 COMPLETION OF THE BID – ADMISSIBLE TENDER

»Admissible tender« shall mean a tender that is submitted by a tenderer for which there are no grounds for exclusion and which meets the selection criteria, which meets the needs and requirements of the Contracting Authority set out in the technical specifications and the procurement documents, which was received in due time, for which there is no evidence of collusion or corruption, which has not been found by the Contracting Authority to be abnormally low, and the price does not exceed the Contracting Authority’s budget.

Tender documentation must consist of the forms submitted in the following order:

  - FORM 1 – The bid (Proforma invoice)

  - FORM 2 – Details about the Bidder

  - FORM 3 – Completed »ESPD« form

  - FORM 4 – Statement of Technical suitability

  - FORM 5 – The completed, signed and stamped sample contract initialled on each page

  - Appendix 2 – Statement with details about the natural and legal entities owned by the bidder

  - Evidence relating to the fulfillment of the requirements of the technical specifications (i.a. brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included)

Only when acting as consortium and/or with Subcontractors the Bidder must submit the following forms:

  - Partnership Agreement

  - FORM 1 – The bid (Proforma invoice)

  - FORM 2.1 – General data about the Bidder and the consortium

  - FORM 2.2 – Data about the subcontractors and subcontractor's authorisation for direct payments

  - FORM 3 – Completed »ESPD« form

  - FORM 4 – Statement of Technical suitability

  - FORM 5 – The completed, signed and stamped sample contract initialled on each page

  - Appendix 2 – Statement with details about the natural and legal entities owned by the bidder

  - Evidence relating to the fulfillment of the requirements of the technical specifications (i.a. brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included)

# 1.15 FORM »BID (PROFORMA INVOICE)«

The Bidder must offer all listed items from the Form 1 – Bid (Proforma invoice), considering the technical description and the specification of the equipment, which is part of the Tender Documentation.

The Bidder completes all the listed items in the Proforma invoice, to a maximum of two decimal places.

If the Bidder declares a price of zero (0), it is understood this item is offered free of charge.

The Bidder must not change the content of Form 1.

The price must include all discounts and costs.

Where the Contracting Authority, during the examination and evaluation of tenders, finds that a calculation error has occurred, it may correct it in line with Article 89, paragraph 7, of the PPA-3.

Bidder uploads in the IT ePonudbe.si system under the section »Pro-forma Invoice« (»Predračun«) in a pdf file the Proforma Invoice.

# 1.16 FORM »ESPD« FOR ALL ECONOMIC OPERATORS

European Single Procurement Document (ESPD) is a formal statement by the economic operator that the relevant grounds for exclusion do not apply, that the relevant selection criteria are fulfilled and that it will provide the relevant information as required by the contracting authority or contracting entity. The ESPD form shall also include a formal statement that the contracting authority may ask any tenderer at any moment during the procedure to submit all or part of the required certificates and supporting documents where this is necessary to ensure the proper conduct of the procedure.

The entries in the ESPD form and/or the evidence submitted by the economic operator must be valid.

The economic operator imports XML.file prepared by the contracting authority to the ESPD file on the website: https://ejn.gov.si/espd and directly enters the requested data into it.

Where groups of economic operators (including temporary associations, economic operators on whose capacity is referred to by the economic operator and subcontractors) participate together in the procurement procedure, a separate ESPD must be given for each of the participating economic operators.

The bidder who submits the bid in the ePonudbe.si system uploads its ESPD in the section »ESPD-Bidder« section and places the other ESPDs in the »ESPD-Other Participants« section. A bidder who submits an offer in the ePonudbe.si system shall upload the electronically signed ESPD into xml. format or unsigned ESPD in xml. in the latter case, in accordance with the General Terms and Conditions of use of the ePonudbe.si information system, a legally binding document having the same validity as is signed shall be deemed to have been submitted.

For the other participating economic operators from the group, the bidder attaches signed ESPDs in the pdf. form or electronically signed xml. in the »ESPD – Other Participants« section.

# 1.17  CRITERIA FOR ASSESSING AND EVALUATING THE BIDS

The contracting authority will award the contract based on the most economically advantageous tender. Timely submitted bids will be assessed in accordance with the criteria specified below.

| No, |  | Scale |  | points |
| --- | --- | --- | --- | --- |
| 1 |  | Vrednost ponudbe: A = vrednost najcenejše ponudbe, ki ustreza pogojem. Vrednosti so ocenjene na osnovi odstopanja od A | A + 3% | 50 |
|  |  |  | more than 3% till 6% 90 | 40 |
|  |  |  | more than 6% till 9% | 30 |
|  |  |  | more than 9% | 20 |
| 2 |  | Lead time in weeks | less than 15 weeks | 10 |
|  |  |  | 15-20 weeks | 5 |
|  |  |  | more than 20 weeks | rejection |
| 3 | Compatibility of acquisition modules with most commercially available microscope bodies offered by main high-end microscope suppliers (Olympus, Leica, Zeiss, Nikon), not only with a body used in specific installation, where module is installed and proven to work | Yes | 20 |
|  |  | No | 0 |
| 4 | Number of new microscope bodies | 2 or more | 5 |
|  |  | less than 2 | 0 |
| 5 | Number of user control PCs for all stations together | 3 or more | 5 |
|  |  | less than 3 | 0 |

The bidder obtaining the most points out of 90 possible points shall be asked to realise its bid.

EXAMPLE OF BIDS WITH EQUAL TO THE BEST RESULTS

In the event that two or more bidders receive the same number of points after the completion of the evaluation, the bid with the shortest delivery time shall be selected. If the bids are still tied after that, the bidder offering the lowest bid price in EUR excluding VAT shall be selected. If the bid prices are the same the bidder with better conditions under No. 3, 4 and 5 respectively (if bidders have same number of points, next No. is considered).

# 1.18 ANTI-CORRUPTION CLAUSE

In the process of awarding a public contract, the client and bidders may not initiate and perform actions that would predetermine the selection of a specific offer, or that would cause the contract not to enter into force or not be fulfilled. Any kind of lobbying in public procurement procedures is prohibited.

# 1.19 CONTRACT AWARD NOTICE

The Contracting Authority will publish the signed decision on awarding the contract on the public procurement portal. The decision is considered to have been served on the date of its publication on the public procurement portal.

# 1.20 RIGHT OF WITHDRAWAL FROM THE PUBLIC CONTRACT

The Contracting Authority may, according to the eighth paragraph of Article 90 of the ZJN-3, after the decision to award the contract has been made, withdraw from the performance of the public contract until the conclusion of the contract for justified reasons, that they no longer need the object of the public contract or that they do not have guaranteed funds for it, or that a valid reason arises with the client suspicion that the content of the contract was or could be the result of a committed criminal act or that other extraordinary circumstances arose which the client could not influence and foresee and which made the performance of the public contract with the selected provider impossible. In this case, the contracting authority will notify the bidders in writing of its decision and the reasons for withdrawing from the execution of the public contract.

# 1.21 SIGNING OF THE CONTRACT

The Contract shall be prepared for signature by the Jožef Stefan Institute as The Contracting Authority and forwarded to the selected Bidder for execution.

In line with the 6th paragraph of Article 14 of the Integrity and Prevention of Corruption Act (Official Gazette of the Republic of Slovenia, No. 69/11–official consolidated version and 158/20), the selected Bidder is obliged to submit, prior to signing the contract and at the request of the Jožef Stefan Institute, the awarding authority, a statement or details about the natural or legal entities owned by the selected Bidder, including the silent partners and business entities that are considered to be associated with the selected Bidder in line with the provisions of the law governing companies. If the Bidder submits a false statement or gives untrue information about the required details, the contract shall be annulled.

The selected bidder shall submit to the Contracting Authority, within 8 days of the receipt of the request, the details about:

- its founders, partners, shareholders, limited partnerships or other owners, as well as the equity shares of these entities;

- business entities that are considered to be associated with the selected bidder on the basis of the law governing companies.

Prior to signature, the Contract shall be amended and finalised as necessary to reflect the specific circumstances of the Selected Bidderer’s offer, in particular in cases where the Selected Bidder has submitted a joint tender, intends to engage subcontractors, or has otherwise arranged the performance of the Contract in a specific manner.

The selected Bidder shall sign the contract within 8 working days of the receipt of the contract signed by the Contracting Authority.

# 1.22 LEGAL INSTRUCTION

A review claim referring to the contents of the call and/or the tender documentation, may be filed in 10 working days from the date of publication of the call for tender or the notice of supplementary information, information on the pending procedure or correction if this is a notice to amend or supplement requirements or criteria for selecting the most advantageous tenderer from the tender documentation or previously published contract notice; the review claim may refer to the amended, supplemented or clarified content of the publication or the tender documentation or a directly related reference to it in the original publication or tender documentation. A review claim may not be submitted after the deadline for submission of the bid, unless the deadline submission of the bid is less than 10 working days. In this case, a review claim may be filed within 10 working days of the date of publication of the contract notice.

A request for a review shall be submitted via the eRevizija (eReview) portal (https://www.portalerevizija.si).

The applicant shall pay a tax of 4.000,00 euros to the bank account of the Ministry of Finance, No. SI56 0110 0100 0358 802, opened at the Bank of Slovenia, Slovenska 35, 1505 Ljubljana, Slovenia, SWIFT CODE: BS LJ SI 2X; IBAN: SI56011001000358802 – a tax for a revision of public procurement.

# TECHNICAL DOCUMENTATION OF THE PUBLIC TENDER

The subject of the public tender is the supply and installation of a multi-bench hardware-sharing confocal imaging platform for hunting of functional cellular events dedicated to the process of knowledge discovery - beyond the classic confocal microscopes or today’s emerging high-throughput imaging instruments that can generate lots of images or at best large quantities of data. Together with the analysis algorithms for quantification of biology-informed cellular events and their connection into the proposed mechanism of action, we aim to offer a unique platform for streamlined research and development, accessible to every lab anytime.

With the microscope that is subject of this public tender we aim to increase accessibility of the microscopes by reduced price per measuring station as well as create a unique platform that will enable development of a proof of a more general hardware sharing concept (sharing all lasers and detectors) that has the potential to reduce the price-per-measuring station even further.

To achieve such a platform, the following TECHNICAL SPECIFICATIONS need to be met:

- A. Functionality of the microscope

The provided instrument should consist of four independent measuring stations. Three of the stations should have independent control via control software and should share at least one laser. The fourth station needs to be a platform ready for development of synchronization and hardware sharing with one of the three stations.

Being ready for development of synchronization and hardware sharing consists of:

- Laser sources can be moved from one station to a breadboard. The fibres that guide the laser light to the scan heads can be moved to a breadboard.

- Detectors from one station can be moved to a breadboard or optics fibre from the detectors is provided to the breadboard.

- Provided plan for coupling of detection of two stations (station providing the lasers and detectors and fourth station) with appropriate optics cables provided.

- Long enough signal cables to relevant hardware (e. g. to scanner head) to reach both stations from a control box.

- A full microscope body with the same stage, z-piezo, scanner module and autofocus module as the stations with independent control (this body can have non-motorized nosepiece – manual focus and objective quick-swap)

Each of the stations with independent control should have the following functionalities:

  - confocal fluorescence microscopy

  - back-scatter detection

  - 2D STED fluorescence microscopy

  - full software control for each independently controlled station

  - programmatic/scripting (e.g. Python) control over key microscope functionalities needed for multichannel high-throughput imaging (e.g., running the measurement, ROI selection, upload of custom settings of 8 channels or more, gating control, laser power control, number and position selection of z-slices, time-point selection)

  - full compatibility with high-throughput multi-day time-lapse imaging, where the bid must include the incubators with controlled environment (5 % CO2, 100% humidity, 37C) and 96-well compatible holder

  - system must consist of upgrading the commercially available microscopy body of high-end microscope suppliers (for example Olympus IX83 or an equivalent class) by use of free camera port to a STED super-resolution system

  - hardware ready for FLIM (TCSPC), upgrade possible with simple add-on of TCSPC electronics with Phaser plot analysis implemented in the basic measurement software.

- For some of the more robust component we allow the incorporation of refurbished components (f. a. microscope body), but the following sensitive components need to be new:

- detectors

- lasers

- xy stages

- z-piezo stages

- incubators

- objectives

In this respect each independetly controlled station should include a set of following independent components, unless otherwise specified:

- 1. microscope

- inverted microscope body

  - fully motorized

  - control by computer and touchpad

  - mass <50 kg (including objectives, turrets, ocular, stage, etc.)

  - two decks

- motorized stage (XY)

  - stroke of at least 110 mm × 80 mm

  - max speed >25 mm/s

  - step <0.00001 mm

  - possibility of changing holders (for petri dishes, microtiter plates, incubator, etc.)

  - control by computer and rotating knobs/joystick

  - compatibility with incubator that can hold 96-well plate holder (MW-oil holder)

  - max load >0.5 kg

- motorized focus

  - stroke >1 cm

  - precision <15 nm

  - max speed >2.5 mm/s

- motorized switching between objectives

  - at least 6 places for objectives

  - turret sealed against liquids

  - switching time between adjacent objectives < 15 s

- fast piezo for z-scan acquisition

  - precision <15 nm

  - max speed >6 mm/s

  - piezo scan range 200 µm or more

objectives

- magnification 20x

  - air (no liquid immersion)

  - NA >0.7

  - WD >0.5 mm

  - suitable for transmission and epifluorescence microscopy

- magnification 60× - at least one for all stations together

  - water immersion

  - NA >1.2

  - WD >0.25 mm

  - transmittance within wavelength range 400–1300 nm >0.3

  - correction ring for compensation of different coverslip thickness in the range 0.13–0.20 mm

  - suitable for super-resolution microscopy

  - sealed against liquids

  - safety spring

2D beam scanner

- max line frequency > 800 Hz, acquisition time of an image of 512 × 512 pix at 15 nm pixel size <1,1 s

- size of the FOV with 100× objective at least 90 µm × 80 µm

- compensation of image distortions towards the edges of the FOV

- lens-free design for minimal wave-front distortion

- line-interleaved scanning (also between STED and confocal mode)

- rotation of whole scanned area in the whole field of view in both confocal nad STED mode

VIS laser system for (at least) two-colour 1PE fluorescence of red fluorescent dyes

- lasers with wavelengths: 488nm, 561nm and 640nm +- 5 nm of the central wavelength

- pulse power variability below 10%

- pulse duration <200 ps

- max pulse repetition rate >30 MHz

- max average power at max repetition rate >150 µW

- continuously adjustable power of the excitation light through the program

2D STED module for red fluorescent dyes

- STED laser with wavelength within 750–800 nm

- pulse duration within 0.5–5 ns

- max repetition rate >30 MHz

- average output laser power >1.8 W

- continuously adjustable power of the laser light through the control program

- synchronization with excitation lasers, scanner, and detectors

  - adjustable delay with respect to other lasers and detectors with resolution <100 ps

- hands-free optimization of 2D STED beam polarization, e.g., when adding an additional dichroic into the beam path.

- alignment of the STED beam with excitation beams through software or automatic

- final spatial resolution of the system in 2D STED mode (XY) <60 nm, when using Olympus 100x 1.45 oil lens, when using 100 mW fibre output in the main optics path

- aligned-by-design optical system: alignment and maintenance-free optical beam path

- motorized pinhole with at least 12 steps for high freedom of adjustability

  - One pinhole needs to be of 1.1 Airy units’ size

- B. Detection

- confocal pinholes of variable sizes within 10–1000 µm (corresponding to the sizes in the intermediate image at 1x magnification)

- 4 avalanche photodiodes (APD) or hybrid single photon-counting detectors

  - QE >0.6 within wavelenght range 600–700 nm

  - low dark current < 250 Hz

  - temporal resolution of single photon detection <0.5 ns

  - pulse width <15 ns

  - dead time <30 ns

  - saturation >35 Mcounts/s

  - probability of afterpulsing <1%

  - enabling gated detection

- dischroics and filters for four detection channels

  - 505-550nm,

  - 575-625nm,

  - 650-700nm

  - 485-495nm

    - Filtering before the pinhole is adapted to allow roughly 1% transmission of 488nm light.

- C. control, connectivity, synchronization

- sync resolution <100 ps for excitation, STED, and (gated) detection

- at least 8 logical detection channels (freely assignable to physical detectors, gating settings, routing signals, etc.)

- ability to control of at least 5 lasers (including STED)

- specifications of user PC controling the software

  - RAM

    - size at least 64GB

    - type DDR5 (or faster)

    - file transfer >2 GT/s

  - fast multi-core processor

    - at least 20 cores

    - at least 8 threads

    - frequency >3 GHz

    - cache memory at least 28 MB

  - high-performance graphical card

    - CUDA support

    - at least 6000 CUDA cores

    - memory at least 15 GB GDDR6

  - disc

    - SSD 2 TB or larger

      - At least Max Sequential Write CDM 4000 Mb/s

      - At least Max Sequential Read CDM 5500 Mb/s

    - connection SATA3 ready for big storage upgrade

      - including at least 2 SATA disc holders

  - I/O

    - USB 3.2 Gen 1 at least 2×

    - USB 3.2 Gen 2 (Type-C) at least 1×

    - USB in total at least 5×

    - digital video (DVI, HDMI, DP) at least 2×

    - at least 1 free PCIe place

    - audio

    - 10 Gb/s Ethernet Card

  - monitor

    - 30''or larger (or 2 smaller with equivalent surface)

    - refresh frequency >50 Hz

    - gamut coverage at least 100% sRGB

    - colour depth at least 30 bit

    - luminosity >250 cd/m2

  - operating system

    - 64-bit

    - Windows, version 11

  - external devices

    - full-size keyboard

    - cordless mouse

- D. software

- browser based operational on any modern PC, Mac or tablet

- full control of motorized/electronic components

- freely customizable acquisition routines (delays, gating, dimensions, synchronization, number of z-slices, ROI selection, time-lapse with multiple ROI setup)

  - possibility of programatic definition (e. g. Python) of at least 7 + 1 channels (7 fulorescent and 1 scatter channel)

- auto-save function

- overview funtction

  - needs to enable fast overview images of the sample (<2.4s/image)

    - 100x objective at 200 nm pixel size, 10 us dwell time, 90 µm x 80 µm FOV

    - 20x NA 0.8 objective at 1 µm pixel size, 10 us dwell time, 450 µm x 400 µm FOV

  - user driven navigation of the sample during overview

  - definition of regions of interest via overview functionality

- tiling functionality for imaging of bigger FOVs

- handling of multidimensional experiments

  - at least 4D (f.a. xyz + time) together with channels and multi-roi information

- data analysis

  - intensity in ROI

  - intensity along a line

  - deconvolution

- exporting of images and analysis results into a common image or data format (e.g., tif, gif, png, jpg, and txt, dat, csv)

- saving of settings for different experiments

- licence transferable to a new control computer

- free licences for offline analysis on personal computers

- License includes warranty to provide free critical updates/upgrades to ensure compatibility with the currently available operating system and its updates for at least 3 years; after that, the license should enable the machine to work without updates to the operating system and drivers if the compatibility with the operating system and its updates cannot be maintained.

- E. Long-term live-cell imaging compatibility

- Stage-top incubator for controlled sample environment (temperature, humidity, CO2) during acquisition:

  - heating of the sample and objective within 25–60 °C, feedback loop for temperature stability and accuracy <0.5 °C

  - manual mixing of air and CO2 (connection for 100% CO2) with indicator

  - humidifier to reach near saturated humidity

  - pump and tubing for flow of the gas mixture into the sample chamber

  - at least one heating ring with diameter 25 to 32 mm per incubator

  - inserts for samples of various sample geometries

- standard microscopy slides,

- 35-mm petri-dishes,

- 96-multiwell plates

  - compatible with brightfield microscopy

  - compatible with the microscope stage

- Autofocus module to hold focus position over time during both confocal and STED imaging.

- F. Physical properties

- dimensions of the system on the table preferably not more than 0.7 m × 0.8 m (top view)

- combined weight (microscope body, controller, STED module): < 100 kg

- light-tight optical path

- compliance with all laser-safety standards (safety interlocks etc.)

- all power supplies with 240 V AC, 50/60 Hz.

- the system should operate in the electrical grid 220/400 V

- operation in a room with temperature stabilization within 1 K in the range 20–25 °C

- G. Customization

- Combiner module for at least one of the lasers, splitting the power to three different optics paths, fibre coupled into the same optics path as excitation lasers.

  - Splitting ratios should be 40 % to independently controlled station assoicated with development station and 20% to each of the two remaining stations.

- scatter detection with 488 nm laser for each of the independently controlled stations

ACCEPTANCE TESTS (acceptance criteria):

- For each of the measuring stations

- confirm all connected modules (microscope, stage, scanner, lasers, pinhole exchanger, detectors, etc.) are accessible by software

- confirm user computer PC specifications meet minimal requirements specified in the tender

- confirm laser power at the objective entrance pupil to be in accordance with their typical values

- confirm the size of the FOV with 100× objective at least 90 µm × 80 µm, calibrated to real units

- total range of the stage confirmed to be at least 100 mm x 70 mm

- confirm, detectors and pinholes are aligned

- spatial resolution demonstrated with immobilized fluorescent beads of a known size (diameter <45 nm) for every suitable combination of excitation and STED lasers using 100x 1.45 NA oil immersion objective:

  - XY resolution in confocal mode better than 300 nm

  - XY resolution of 60 nm or better in 2D STED

All acceptance resolution tests should be performed as follows:

- intensity profile along a line across a sample object on a non-processed image is taken

- FWHM of the intensity profile is determined by fitting of a Gaussian function

- the process is repeated for at least 10 sample objects from at least 3 independently acquired images

- 50% of the obtained values must be lower than or equal to the required

- max scanning speed demonstrated, e.g., by measuring the time of the image acquisition of an image 512 × 512 pixels at 15 nm pixel size which should be less than 1,3 s

- sync and gating of lasers/detectors demonstrated

- full functionality of the software demonstrated, according to the required specifications above (control, acquisition, analysis, exporting)

- customization demonstrated according to the requirements above:

  - at least 20% (40% on control station that will be connected to the development part) of expected STED power on each station compared to a single station coupled STED power (single station coupled STED should have > 500 mW from the optics fibre output on the main optics path).

  - demonstration of scatter detection on sub-diffraction sized metal beads using 488 nm laser

- demonstration of 8 channel recording with files of a size > 10 MB (preferrable 15 MB) per position with following parameters run with script:

  - < 16s per slice for 750 x 750 pixels image with 0.5 µm pixel size and saved file size > 10 MB

    - one ROI, one tile, one slice is recorded and timed

  - two successfully ran time points 60 regions of interest, 2 tiles, 3 z-slices with time per slice the same as confirmed in a. (no wait time buildup)

    - 10x times in a row, a single ROI recording time is measured and the number divided by tile and z-slice number, average value is calculated

    - total recording time is recorded and divided by number of ROIs, tiles and z-slices

    - both numbers need to be within 15 % of the number determined in a.

All stations together

- demonstration of simultaneous acquisition independently at the same time on all independently controlled stations

- demonstration of STED capability on independently controlled stations using shared STED laser

OTHER REQUIREMENTS:

- Documentation:

- the equipment should be CE compatible

- documentation and manuals provided, including basic optical schemes and description of electronic interface to connect external devices, optical checks and procedures to maintain the machine working at the required characteristics

- the price should include delivery charges and transport insurance charges, installation, testing and on-site crew training at customer’s site in Ljubljana; price should include all travel and lodging costs of the supplier’s crew.

- Delivery:

- delivery time: no longer than 15 weeks ARO (desired delivery date is no later than 30th April 2026)

- installation at the customer's site to full functionality (acceptance tests above)

- at least one-day of hands-on training for researchers on the installed instrument at the customer’s site

- Technical support:

- the bidder is responsible for technical support and service

- online support for hardware and software issues (at least weekdays 9am–3pm)

- all remote support must be provided for free during warranty period

- response time for service at the customer’s site of 10 working days or less after the receipt of a notification if it is identified problem cannot be solved remotely

- warranty period at least 12 months on all parts, 36 months on control (FPG) electronics, excluding consumables

- out-of-warranty support against payment, at least 10 years after the purchase date

- expert suggestions on experimental issues (sample preparation, probes and labelling, etc.) by email or telephone

# 3. TENDER DOCUMENTATION

The Bidder prepares the offer by entering all the required data in the forms that are an integral part of the tender documentation and attaching all other required documentation. The offer must be submitted on the forms from the tender documentation or on the bidder's forms, which must not differ significantly in content from the attached forms.

All required tender documents must be completed, signed by the legal representative and stamped, if the tenderer uses a company stamp in its business operations.

Tender documentation must consist of the forms submitted in the following order:

  - FORM 1 – The bid (Proforma invoice)

  - FORM 2 – Details about the Bidder

  - FORM 3 – Completed »ESPD« form

  - FORM 4 – Statement of Technical suitability

  - FORM 5 – The completed, signed and stamped sample contract initialled on each page

  - Appendix 2 – Statement with details about the natural and legal entities owned by the bidder

  - Evidence relating to the fulfillment of the requirements of the technical specifications (i.a. brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included)

The tenderer submitting the offer accepts criminal and material responsibility that all data and documents included in the tender are true. Otherwise, the tenderer shall be liable to the Contracting Authority for any damage caused to it.

BID (PROFORMA INVOICE)

FORM 1

| Name of the Bidder: |  |
| --- | --- |
| Address of the Bidder: |  |
| Contact person: |  |
| Contact person's E-mail address: |  |

On the basis of your public tender we are pleased to submit the following bid:

| BID No. |  |
| --- | --- |

| Type of purchasing: |  | Service | X | Equipment |  | Const |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

| No. | Description of the goods | Quantity | Price | Discount % | Value |
| --- | --- | --- | --- | --- | --- |
| 1 | SUPPLY OF A MULTI-BENCH HARDWARE-SHARING CONFOCAL IMAGING PLATFORM FOR HUNTING OF FUNCTIONAL CELLULAR EVENTS | 1 Set |  | . |  |

| TOTAL |  |
| --- | --- |
| Discount |  |
| VALUE without VAT |  |

| Payment terms: |  |
| --- | --- |
| Delivery time in weeks: |  |
| Delivery term: | DDU Ljubljana |
| Validity of the bid: | 30 April 2026 |
| Warranty period: |  |

MANDATORY ENCLOSURE:

Brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included (please, enclose it to the second part of the tender documentation)

|  |  | The Bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

DETAILS ABOUT THE BIDDER

FORM 2

| Name of the Bidder: |  |
| --- | --- |
| Address of the Bidder: |  |
| Legal representative(s) of the Bidder: |  |
| VAT identification number: |  |
| Name of the bank, SWIFT, IBAN: |  |
| Contact person: |  |
| Phone: |  |
| E-mail: |  |
| Responsible person for signing the contract: |  |

On the basis of a public tender published on the Public Procurement Portal we are applying to your invitation to tender and enclose our offer documents in accordance with the Instructions for the bidders.

Acting as a Bidder, we are submitting the bid:

|  | Individually, without contract partners or subcontractors |
| --- | --- |
|  | As a joint bid as a consortium, involving contract partners |

|  | As a bid invoilving subcontractors |
| --- | --- |

|  | As a bid using the capabilities of other operators |
| --- | --- |

|  |  | The Bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The Contracting Authority’s note: Form must be completed and attached to the bid, regardless the fact whether you submit the bid individually, as a joint bid as a consortium or with a joint bid with subcontractors.

GENERAL DATA ABOUT THE BIDDER AND THE CONSORTIUM

FORM 2.1

| Name of the Bidder: |  |
| --- | --- |
| Address of the Bidder: |  |
| Legal representative(s) of the Bidder: |  |
| VAT identification number: |  |
| Name of the bank, SWIFT, IBAN: |  |
| Contact person: |  |
| Phone: |  |
| E-mail: |  |
| Responsible person for signing the contract: |  |

On the basis of a public tender published on the Public Procurement Portal we are applying to your invitation to tender and enclose our offer documents in accordance with the Instructions for the bidders.

Contract partners in a joint bid and description of work:

| No. | Name of the partner | Description of the works | Value | Share in % |
| --- | --- | --- | --- | --- |
| The group leader in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |

|  |  | The Bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The Contracting Authority’s note: Form must be completed and attached to the bid, only whether you submit the joint bid as a consortium.

DATA ABOUT THE SUBCONTRACTORS AND SUBCONTRACTOR'S AUTHORISATION (relating to direct payments)

FORM 2.2

| Name of the Subcontractor: |  |
| --- | --- |
| Address of the Subcontractor: |  |
| Legal representative of the Subcontractor: |  |
| VAT identification number: |  |
| Phone: |  |
| E-mail: |  |

| Part of the execution of the contract that will be performed by a Subcontractor |  |
| --- | --- |
| Description of work |  |
| Količina del podizvajalca v % |  |
| Value (without VAT) |  |
| Time of completion |  |
| Place |  |

In accordance with paragraph 5 of Article 94 of PPA-3 we declare (tick):

|  | YES, we request direct payments by the Contracting Authority and give our consent, and on the basis of the fourth indent of the second paragraph of Article 94, we request that the Contracting Authority "JOŽEF STEFAN" INSTITUTE for the public contract , will settle our claims against the Bidder (name of the Bidder) instead of the Bidder directly to us, to our transaction account. |
| --- | --- |
|  | We do NOT require direct payments from the Contracting Authority. In this case, we will receive payment for the work performed by the selected Bidder. |

Subcontractors who authorize the Contracting Authority to make direct payments and above tick YES, by signing this statement give consent to the Contracting Authority to pay the Subcontractors' claims against the Bidder.

If the Bidder will carry out the contract with Subcontractors, it must provide a list of all the Subcontractors and the type of work that the Subcontractors will be undertaking, contact information and legal representatives of the Subcontractors and complete An ESPD of Subcontractors in accordance with 79. Article of PPA-3 and accompanied by a request for direct payment of the Subcontractor if requested by the Subcontractor.

|  |  | The Subcontractor: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The Contracting Authority’s note: Form must be completed and attached to the bid, only whether you submit the joint bid as a consortium

ESPD for all economic operators

FORM 3

European Single Procurement Document (ESPD) is a formal statement by the economic operator that the relevant grounds for exclusion do not apply, that the relevant selection criteria are fulfilled and that it will provide the relevant information as required by the contracting authority or contracting entity. The ESPD form shall also include a formal statement that the contracting authority may ask any tenderer at any moment during the procedure to submit all or part of the required certificates and supporting documents where this is necessary to ensure the proper conduct of the procedure. The entries in the ESPD form and/or the evidence submitted by the economic operator must be valid.

The economic operator imports XML.file prepared by the contracting authority to the ESPD file on the website: https://ejn.gov.si/espd and directly enters the requested data into it.

Where groups of economic operators (including temporary associations, economic operators on whose capacity is referred to by the economic operator and subcontractors) participate together in the procurement procedure, a separate ESPD must be given for each of the participating economic operators.

The bidder who submits the bid in the ePonudbe.si system uploads its ESPD in the section »ESPD-Bidder« section and places the other ESPDs in the »ESPD-Other Participants« section. A bidder who submits an offer in the ePonudbe.si system shall upload the electronically signed ESPD into xml. format or unsigned ESPD in xml. in the latter case, in accordance with the General Terms and Conditions of use of the ePonudbe.si information system, a legally binding document having the same validity as is signed shall be deemed to have been submitted.

For the other participating economic operators from the group, the bidder attaches signed ESPDs in the pdf. form or electronically signed xml. in the »ESPD – Other Participants« section.

STATEMENT OF TECHNICAL SUITABILITY

FORM 4

We hereby declare that we can provide the equipment with the required technical suitability in line with all technical requirements included in the tender documentation.

The supplied a multi-bench hardware-sharing confocal imaging platform will achieve the following TECHNICAL SPECIFICATIONS:

- A. Functionality of the microscope

The provided instrument will consist of four independent measuring stations. Three of the stations will have independent control via control software and will share at least one laser. The fourth station will be a platform ready for development of synchronization and hardware sharing with one of the three stations.

Being ready for development of synchronization and hardware sharing will consist of:

- Laser sources can be moved from one station to a breadboard. The fibres that guide the laser light to the scan heads can be moved to a breadboard.

- Detectors from one station can be moved to a breadboard or optics fibre from the detectors is provided to the breadboard.

- Provided plan for coupling of detection of two stations (station providing the lasers and detectors and fourth station) with appropriate optics cables provided.

- Long enough signal cables to relevant hardware (e. g. to scanner head) to reach both stations from a control box.

- A full microscope body with the same stage, z-piezo, scanner module and autofocus module as the stations with independent control; this body ______________ (has/does not have) non-motorized nosepiece – manual focus and objective quick-swap (bidder needs to specify)

Each of the stations with independent control will have the following functionalities:

  - confocal fluorescence microscopy

  - back-scatter detection

  - 2D STED fluorescence microscopy

  - full software control for each independently controlled station

  - programmatic/scripting (e.g. Python) control over key microscope functionalities needed for multichannel high-throughput imaging (e.g., running the measurement, ROI selection, upload of custom settings of 8 channels or more, gating control, laser power control, number and position selection of z-slices, time-point selection)

  - full compatibility with high-throughput multi-day time-lapse imaging, where the bid includes the incubators with controlled environment (5 % CO2, 100% humidity, 37C) and 96-well compatible holder

  - system consists of upgrading the commercially available microscopy body of high-end microscope suppliers (for example Olympus IX83 or an equivalent class) by use of free camera port to a STED super-resolution system

  - hardware ready for FLIM (TCSPC), upgrade possible with simple add-on of TCSPC electronics with Phaser plot analysis implemented in the basic measurement software.

In this respect each independetly controlled station will include a set of following independent components, unless otherwise specified:

- 1. microscope

- inverted microscope body

  - fully motorized

  - control by computer and touchpad

  - mass <50 kg (including objectives, turrets, ocular, stage, etc.)

  - two decks

- motorized stage (XY)

  - stroke of at least 110 mm × 80 mm

  - max speed >25 mm/s

  - step <0.00001 mm

  - possibility of changing holders (for petri dishes, microtiter plates, incubator, etc.)

  - control by computer and rotating knobs/joystick

  - compatibility with incubator that can hold 96-well plate holder (MW-oil holder)

  - max load >0.5 kg

- motorized focus

  - stroke >1 cm

  - precision <15 nm

  - max speed >2.5 mm/s

- motorized switching between objectives

  - at least 6 places for objectives

  - turret sealed against liquids

  - switching time between adjacent objectives < 15 s

- fast piezo for z-scan acquisition

  - precision <15 nm

  - max speed >6 mm/s

  - piezo scan range 200 µm or more

- 2. objectives

- magnification 20x

  - air (no liquid immersion)

  - NA >0.7

  - WD >0.5 mm

  - suitable for transmission and epifluorescence microscopy

- magnification 60× - at least one for all stations together

  - water immersion

  - NA >1.2

  - WD >0.25 mm

  - transmittance within wavelength range 400–1300 nm >0.3

  - correction ring for compensation of different coverslip thickness in the range 0.13–0.20 mm

  - suitable for super-resolution microscopy

  - sealed against liquids

  - safety spring

2D beam scanner

- max line frequency > 800 Hz, acquisition time of an image of 512 × 512 pix at 15 nm pixel size <1,1 s

- size of the FOV with 100× objective at least 90 µm × 80 µm

- compensation of image distortions towards the edges of the FOV

- lens-free design for minimal wave-front distortion

- line-interleaved scanning (also between STED and confocal mode)

- rotation of whole scanned area in the whole field of view in both confocal nad STED mode

VIS laser system for (at least) two-colour 1PE fluorescence of red fluorescent dyes

- lasers with wavelengths: 488nm, 561nm and 640nm +- 5 nm of the central wavelength

- pulse power variability below 10%

- pulse duration <200 ps

- max pulse repetition rate >30 MHz

- max average power at max repetition rate >150 µW

- continuously adjustable power of the excitation light through the program

2D STED module for red fluorescent dyes

- STED laser with wavelength within 750–800 nm

- pulse duration within 0.5–5 ns

- max repetition rate >30 MHz

- average output laser power >1.8 W

- continuously adjustable power of the laser light through the control program

- synchronization with excitation lasers, scanner, and detectors

  - adjustable delay with respect to other lasers and detectors with resolution <100 ps

- hands-free optimization of 2D STED beam polarization, e.g., when adding an additional dichroic into the beam path.

- alignment of the STED beam with excitation beams through software or automatic

- final spatial resolution of the system in 2D STED mode (XY) <60 nm, when using Olympus 100x 1.45 oil lens, when using 100 mW fibre output in the main optics path

- aligned-by-design optical system: alignment and maintenance-free optical beam path

- motorized pinhole with at least 12 steps for high freedom of adjustability

  - One pinhole needs to be of 1.1 Airy units size

- B. Detection

- confocal pinholes of variable sizes within 10–1000 µm (corresponding to the sizes in the intermediate image at 1x magnification)

- 4 avalanche photodiodes (APD) or hybrid single photon-counting detectors

  - QE >0.6 within wavelenght range 600–700 nm

  - low dark current < 250 Hz

  - temporal resolution of single photon detection <0.5 ns

  - pulse width <15 ns

  - dead time <30 ns

  - saturation >35 Mcounts/s

  - probability of afterpulsing <1%

  - enabling gated detection

- dischroics and filters for four detection channels

  - 505-550nm,

  - 575-625nm,

  - 650-700nm

  - 485-495nm

    - Filtering before the pinhole is adapted to allow roughly 1% transmission of 488nm light.

- C. Control, connectivity, synchronization

- sync resolution <100 ps for excitation, STED, and (gated) detection

- at least 8 logical detection channels (freely assignable to physical detectors, gating settings, routing signals, etc.)

- ability to control of at least 5 lasers (including STED)

- specifications of user PC controling the software

- RAM

- size at least 64GB

- type DDR5 (or faster)

- file transfer >2 GT/s

- fast multi-core processor

- at least 20 cores

- at least 8 threads

- frequency >3 GHz

- cache memory at least 28 MB

- high-performance graphical card

- CUDA support

- at least 6000 CUDA cores

- memory at least 15 GB GDDR6

- disc

- SSD 2 TB or larger

      - At least Max Sequential Write CDM 4000 Mb/s

      - At least Max Sequential Read CDM 5500 Mb/s

- connection SATA3 ready for big storage upgrade

      - including at least 2 SATA disc holders

- I/O

- USB 3.2 Gen 1 at least 2×

- USB 3.2 Gen 2 (Type-C) at least 1×

- USB in total at least 5×

- digital video (DVI, HDMI, DP) at least 2×

- at least 1 free PCIe place

- audio

- 10 Gb/s Ethernet Card

- monitor

- 30''or larger (or 2 smaller with equivalent surface)

- refresh frequency >50 Hz

- gamut coverage at least 100% sRGB

- colour depth at least 30 bit

- luminosity >250 cd/m2

- operating system

- 64-bit

- Windows, version 11

- external devices

- full-size keyboard

- cordless mouse

- D. Software

- browser based operational on any modern PC, Mac or tablet

- full control of motorized/electronic components

- freely customizable acquisition routines (delays, gating, dimensions, synchronization, number of z-slices, ROI selection, time-lapse with multiple ROI setup)

  - possibility of programatic definition (e. g. Python) of at least 7 + 1 channels (7 fulorescent and 1 scatter channel)

- auto-save function

- overview funtction

  - needs to enable fast overview images of the sample (<2.4s/image)

    - 100x objective at 200 nm pixel size, 10 us dwell time, 90 µm x 80 µm FOV

    - 20x NA 0.8 objective at 1 µm pixel size, 10 us dwell time, 450 µm x 400 µm FOV

  - user driven navigation of the sample during overview

  - definition of regions of interest via overview functionality

- tiling functionality for imaging of bigger FOVs

- handling of multidimensional experiments

  - at least 4D (f.a. xyz + time) together with channels and multi-roi information

- data analysis

  - intensity in ROI

  - intensity along a line

  - deconvolution

- exporting of images and analysis results into a common image or data format (e.g., tif, gif, png, jpg, and txt, dat, csv)

- saving of settings for different experiments

- licence transferable to a new control computer

- free licences for offline analysis on personal computers

- License includes warranty to provide free critical updates/upgrades to ensure compatibility with the currently available operating system and its updates for at least 3 years; after that, the license will enable the machine to work without updates to the operating system and drivers if the compatibility with the operating system and its updates cannot be maintained.

- E. Long-term live-cell imaging compatibility

- Stage-top incubator for controlled sample environment (temperature, humidity, CO2) during acquisition:

  - heating of the sample and objective within 25–60 °C, feedback loop for temperature stability and accuracy <0.5 °C

  - manual mixing of air and CO2 (connection for 100% CO2) with indicator

  - humidifier to reach near saturated humidity

  - pump and tubing for flow of the gas mixture into the sample chamber

  - at least one heating ring with diameter 25 to 32 mm per incubator

  - inserts for samples of various sample geometries

- standard microscopy slides,

- 35-mm petri-dishes,

- 96-multiwell plates

  - compatible with brightfield microscopy

  - compatible with the microscope stage

- Autofocus module to hold focus position over time during both confocal and STED imaging.

- F. Physical properties

- dimensions of the system on the table preferably not more than 0.7 m × 0.8 m (top view)

- combined weight (microscope body, controller, STED module): < 100 kg

- light-tight optical path

- compliance with all laser-safety standards (safety interlocks etc.)

- all power supplies with 240 V AC, 50/60 Hz.

- the system will operate in the electrical grid 220/400 V

- operation in a room with temperature stabilization within 1 K in the range 20–25 °C

- G. Customization

- Combiner module for at least one of the lasers, splitting the power to three different optics paths, fibre coupled into the same optics path as excitation lasers.

  - Splitting ratios should be 40 % to independently controlled station assoicated with development station and 20% to each of the two remaining stations.

- scatter detection with 488 nm laser for each of the independently controlled stations

- H. Additional requirements:

- Lead time in weeks is ______________ (planned delivery date is ____________). (bidder needs to specify)

- Acquisition modules _______________ (are/are not) compatible with most commercially available microscope bodies offered by main high-end microscope suppliers (Olympus, Leica, Zeiss, Nikon). (bidder needs to specify)

- Number of new microscope bodies is _______________. (bidder needs to specify)

- Number of user control PCs for all stations together ___________________. (bidder needs to specify)

This statement is an integral part of, and an annex to, the application to the public tender for the supply of »SUPPLY OF A MULTI-BENCH HARDWARE-SHARING CONFOCAL IMAGING PLATFORM FOR HUNTING OF FUNCTIONAL CELLULAR EVENTS«.

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

SAMPLE CONTRACT

FORM 5

CONTRACT

concluded between

| The Client: | The Supplier: |
| --- | --- |
| JOŽEF STEFAN INSTITUTE Jamova cesta 39 1000 Ljubljana represented by the director , director | (company name) (address) represented by the director , director |
| VAT identification number: SI55560822 | VAT identification number: |

INTRODUCTORY PROVISIONS

Article 1

The contracting parties have established that, in line with the Public Procurement Act, the Client awarded the contract as the most cost-effective Bidder with regard to the realisation of the supply requirement expressed in the public tender published on the Public Procurement Portal of the Republic of Slovenia under the reference number ……………..

With this contract the Client and the Supplier agree on their mutual rights and obligations, necessary for the smooth operation of purchasing and selling.

SUBJECT-MATTER OF THE CONTRACT

Article 2

Subject-matter of the contract is the

SUPPLY OF A MULTI-BENCH HARDWARE-SHARING CONFOCAL IMAGING PLATFORM

FOR HUNTING OF FUNCTIONAL CELLULAR EVENTS

specified, in detail, in the Quotation/pro-forma invoice with the ref. number ………..... as of ………….... that is an integral part of this contract.

An integral part of this contract the client's tender and the entire tender documentation, including all technical documentation.

VALUE OF THE CONTRACT

Article 3

The value of the above equipment is …….………….. . The price does not include the VAT.

The Supplier shall provide, transportation including transport insurance of the equipment to the Client’s address. The costs for transportation including insurance is included in the price as set out in article 2 of this agreement. The Supplier shall notify the Client of the intended delivery by fax, e-mail, in writing or by telephone, at least 3 working days prior to the delivery.

DELIVERY TERMS

Article 4

The delivery period for the equipment that is the subject of this contract is ………......... after the date of the contract.

The Client shall sign the handover note after the establishing that the equipment has been correctly delivered.

The Supplier will engage a transport company to send the equipment to the Client. If the Client finds that the work is carried out by the Subcontractor not listed in the tender, he is entitled to cancel the contract. The Supplier reserves the right to verify, at any time, the Subcontractor employees performing work. All employees are required to give credible information.

PAYMENT TERMS

Article 5

The Supplier shall issue the invoices to the Client in the following order:

- 30 % as advance payment after receipt of the order and signed contract

- 40 % as progress payment after delivery

- 30 % as final payment after installation and technical acceptance

The Client is obliged to pay an invoice for the equipment, within 30 days of the correct issuing of the invoice.

The contracting parties agree that, in the case of a late payment, the Client has to pay the default interest.

ACCEPTANCE

Article 6

The acceptance shall be performed on the basis of the Final Acceptance Certificate (APPENDIX 1), signed by the authorised representatives of both contracting parties, upon the proper delivery of the equipment of adequate quality and quantity.

The Supplier will submit to the Client the following documentation:

- Invoice

- Delivery note (with an exact list of delivered parts, data about the gross and net weight)

- Technical documentation and instructions for use

- Handover note – final acceptance certificate

- Warranty certificate

The documentation must be in Slovene or English.

The equipment is taken over upon delivery of the equipment at the delivery location with the delivery note. Quantitative and qualitative inspection of the delivered equipment is carried out on record within 8 days after delivery. A representative of the provider may be present during the inspection. The takeover report on the quantity and quality inspection is signed by the customer's representative and the supplier's representative, if the latter is present.

SUPPLIER GUARENTEES

Article 7

The responsibility for keeping the equipment undamaged is transferred from the Supplier to the Client once the equipment is delivered to the Client.

The Supplier guarantees to the Client:

- that the equipment is new, not used and free of any actual and legal defects;

- that the supplied equipment functions perfectly and that it complies fully with all technical descriptions, characteristics and specifications, provided by the Client;

- that additional activities related to the supplied equipment (installation and training) are carried out professionally,

- that the Client shall acquire all the rights related to the equipment, while the Supplier shall duly perform all the obligations related to the equipment.

The Supplier provides the Client with a warrantee to cover the quality of the equipment and its smooth operation at least to an extent equal to the extent of the warrantee given by the producer of the equipment to the Supplier, i.e., ……. years following the date of the receipt of the equipment.

During the warrantee period the Supplier shall provide for emergency services responding to the Client by E-mail or telephone within three working days after the receipt of a notification. If a repair service is needed, the Supplier shall provide it typically within 60 days after the problem has been identified.

In the case of a defect identified in the minutes upon delivery, the warranty period begins on the day the defect is rectified. If the fault cannot be rectified, the warranty period begins on the day of receipt of the new goods.

ELIMINATION OF ERRORS AND SUPPLY OF SPARE PARTS

Article 8

The Supplier undertakes to eliminate the errors of the delivered goods within the warranty period providing uninterrupted service to its own costs, including transport costs to a specific location. The period of notice is considered to be the time when the message reaches to the supplier at the phone number or e-mail stated in this agreement, provided that it has been awarded by the Client and shall contain at least the necessary data for identification of the goods.

The Supplier shall provide for emergency services responding to the Client within three working days after the receipt of a notification.

The supplier is obliged to eliminate defects that occur during the warranty period within 60 days from the date of notification of the defect to the supplier by the client.

The Supplier commits to provide the spare parts and to provide service support against payment for another 10 (five) years from the end of the warranty period.

In the case of failure to fulfil obligations, the Supplier must reimburse any additional costs and damage to the Client, which would consequently suffer.

SERVICE OF THE EQUIPMENT

Article 9

Service of the equipment is a repair service that is performed in case of equipment failure on the basis of the customer's order after the expiration of the warranty period. Service services are intended to establish such condition of the equipment in which the equipment performs its intended function. Service maintenance costs are calculated according to the actual time and material used. The Supplier informs the Client about at least approximate costs before service is performed.

ENGAGEMENT OF SUBCONTRACTORS

Article 10

(applicable where the Supplier engages one or more Subcontractors)

Artcle a.

The Supplier shall perform the Contract in cooperation with the following Subcontractors::

1.

| Name of Subcontractor | Regstraton No.: VAT-ID: Bank Account (IBAN): SI56 … (held at …) |
| --- | --- |
| Direct Payment Requested: | YES/NO |
| Scope of Subcontract Services: |  |
| Quantity of Supplies to be Delivered by Subcontractor | % |
| Contract Value of works | without VAT |
| Place of Performance |  |
| Timeframe for Performance |  |

Supplier performing a public procurement contract with one or more Subcontractors must fully comply with the obligations set out in Article 94 of the Public Procurement Act (PPA-3) and with the requirements of the tender documentation, and must submit, for all listed Subcontractors, the duly completed, signed, and stamped forms required by the tender documentation. If the Supplier does not act in accordance with Article 94 of PPA-3, the contracting authority will submit a proposal to the National Review Commission to initiate misdemeanor proceedings under Point 2 of the first paragraph of Article 112 of PPA-3.

The Subcontractor must meet all conditions and requirements of the contracting authority regarding Subcontractors as specified in the tender documentation and must complete all annexes related to fulfilling Subcontractor requirements.

The Supplier is fully responsible to the contracting authority for the proper performance of contractual obligations, regardless of the number of Subcontractors.

During the execution of the contract, the Supplier must inform the contracting authority of any changes to the information referred to in the second paragraph of Article 94 of PPA-3 and must submit information on any new Subcontractors they intend to subsequently engage, no later than five (5) days after the change. In the case of new Subcontractors, the Supplier must also provide, together with the notification, the data and documents referred to in the second, third, and fourth indents of the second paragraph of Article 94 of PPA-3.

The contracting authority may reject the proposal to replace a Subcontractor or to include a new Subcontractor if such replacement or inclusion could affect the smooth execution of the supply of goods, or if the new Subcontractor does not meet the conditions set by the contracting authority in the tender documentation. The contracting authority must notify the Supplier of any rejection of a new Subcontractor no later than ten (10) days after receiving the proposal.

For a Subcontractor requesting direct payment, the Supplier must attach to each payment claim either:

- the Subcontractor’s invoice for contractual obligations performed, confirmed by the Supplier, on the basis of which the contracting authority will transfer payment directly to the Subcontractor’s account; or

- a signed declaration by the Subcontractor, addressed to the contracting authority, stating that they are aware of the specific invoice issued by the Supplier or that they did not participate as a Subcontractor in the contractual obligations covered by the payment claim, and that they have no claims and will not have any claims against the contracting authority in connection with that payment claim.

If none of these documents is submitted for the registered Subcontractor, the contracting authority will withhold payment of the entire invoice until all documents are provided, without being in default of payment.

By paying an individual amount directly to the Subcontractor, the contracting authority’s obligation to pay the Supplier is extinguished up to the amount paid to the Subcontractor.

Payment deadlines for the Supplier and Subcontractors are identical.

If direct payment to the subcontractor is not mandatory, the Client shall require the main contractor to provide, no later than 60 days after payment of the final invoice or interim payment, its own written statement and the written statement of the subcontractor confirming that the subcontractor has received payment for the executed works or services, or for the goods supplied, directly related to the subject matter of the public procurement contract.

OR

Article b

(applicable in the event that the Supplier does not engage a Subcontractor )

At the time of submitting the tender and at the conclusion of this contract, the Supplier has not declared any Subcontractors for the execution of the subject matter of the contract.

During the execution of the contract, the Supplier must inform the contracting authority of any changes to the information referred to in the second paragraph of Article 94 of the Public Procurement Act (PPA-3) and must submit information on any new Subcontractors it intends to subsequently engage in the performance of such works or services, no later than five (5) days after the change. In the event of including new Subcontractors, the Supplier must also provide, together with the notification, the data and documents referred to in the second, third, and fourth indents of the second paragraph of Article 94 of PPA-3.

The contracting authority may reject a proposal to replace a Subcontractor or to include a new Subcontractor if such replacement or inclusion could affect the smooth execution or completion of the works, or if the new Subcontractor does not meet the conditions set by the contracting authority in the procurement documentation. The contracting authority must notify the Supplier of any rejection of a new Subcontractor no later than ten (10) days from the receipt of the proposal. A nominated Subcontractor may not subcontract the accepted works further.

A subsequently engaged Subcontractor may not commence the performance of works before the contracting authority has approved its engagement.

The Supplier remains fully responsible to the contracting authority for the proper performance of contractual obligations, regardless of the number of Subcontractors.

EARLY TERMINATION OF THE CONTRACT

Article 11

The Client may withdraw from the contract without obligation to the Supplier if the Supplier:

- does not comply with all the requirements of the client and does not meet them despite the warning of non-compliance

- increases prices during the validity of the contract,

- does not perform the subject of the contract in the agreed quality or within the agreed deadlines,

- does not fulfill all its obligations under the contract,

- in other cases, and to the extent specified in this contract.

The Client will notify the Supplier of the withdrawal with a withdrawal notice sent by registered mail. Withdrawal from the contract is valid from the date of receipt of written notice of withdrawal if the client does not write in the notice a later date of termination of the contract. If the Supplier cannot be served with the withdrawal notice for any reason, the contract shall expire on the day the withdrawal notice was submitted to the post. In this case, the Supplier is not entitled to any compensation or any other claim.

The Supplier has the right to withdraw from this contract in the event of violation of the terms of the contract by the Client. In this case, the contract will expire when the Client receives a written notice of withdrawal, stating the reason for the assignment by the registered mail.

In the event of withdrawal, the parties shall be obliged to fulfil the obligation so far made as agreed prior to the termination.

TERMS FOR TERMINATION OF THE CONTRACT

Article 12

This contract is concluded under the termination terms that will be met if one of the following circumstances is fulfilled:

- if the Contracting Authority is informed that a competent court has established with a final decision that the Suppler or its Subcontractor has infringed applicable obligations in the area of mandatory environmental, social and labour law, or

- if the Contracting Authority is informed that the competent state authority established that the Supplier or its Subcontractor has at the time of the contract infringed at least two infringements in relation to:

- payment for work,

- working hours,

- on rest,

- the performance of work on the basis of civil law contracts, despite the existence of elements of an employment relationship or in connection with the employment of illegal workers and for which an offence has been imposed by a final decision or several final decisions of the competent authority.

If the customer becomes aware of a violation, he must notify the contractor within ten days. If the Client becomes aware of a violation, he must notify the Supplier within ten days. The Supplier may, within the time limit set by the client, which may not be longer than 15 days, submit evidence that he has taken sufficient measures to prove his reliability despite the existence of violations.

If there is a violation by the subcontractor, the Supplier can submit evidence within the same period that the subcontractor has taken sufficient measures to prove its reliability despite the existence of violations.

If the Supplier has not submitted evidence for the subcontractor, or if there is evidence, but the Client considers that these measures are insufficient, the Supplier can replace the subcontractor within a period determined by the Client, which may not be longer than 15 days in accordance with Article 94 of this law. or takes over the part he has subcontracted to this subcontractor himself, if this replacement or takeover does not constitute a substantial change to the contract.

If the Supplier has not submitted evidence for himself or the subcontractor, or if there is any, the Client assesses that these measures are insufficient, or if the Supplier does not take over the work himself or proposes a new subcontractor, or if the Client, in accordance with Article 94 of this law, timely proposes a new subcontractor refuses, the termination condition is implemented on the condition that there is at least six months between the client's knowledge of the breach and the expiry of the contract.

If the termination condition is met, the contract is considered to be terminated on the date of conclusion of the new public procurement contract, and the client must start a new public procurement procedure immediately, but no later than within 60 days of becoming aware of the violation.  If the Client does not start a new public procurement procedure within this period, the contract is considered to be terminated on the sixtieth day after becoming aware of the violation.

ANTI-CORRUPTION CLAUSE

Article 13

If it is found that, during the execution of the public tendering, on the basis of which this contract has been made, or during the execution of this contract, anyone acting in the name of – or on the account of – the Supplier, has promised, offered or given any undue advantage to the Client (the Contracting Authority), or those employees that had a decisive influence on the choice of the Supplier, in order:

- to win this public contract,

- to conclude this contract under more favourable conditions,

- to omit due diligence in the supervision of the contract implementation,

- to carry out any other act or omission causing damage to the Client or allowing the Client or its employees to obtain an undue advantage,

this contract will be void.

In the case of an alleged existence of the conditions from the first paragraph of this article or from a notification of the Commission for the Prevention of Corruption, or any other responsible body, relating to their alleged existence, the Client shall initiate the proceedings for invalidity of the contract from the first paragraph of this article, or any other measures, in line with the regulations of the Republic of Slovenia.

RESOLUTION OF CONFLICTS

Article 14

The contracting parties shall resolve any disagreement relating to the realisation of this contract by common consent.

In the case that the contracting parties cannot settle a dispute by common consent, they shall abide by the decision of the competent court in Ljubljana.

ADMINISTRATORS OF THE CONTRACT

Article 15

The representative of the Client who will be responsible for the technical part of the implementation of the subject of the contract is: ______________________________________

The representative of the Supplier who will be responsible for the technical part of the implementation of the subject of the contract is: ______________________________________

It is not necessary to conclude an annex to the contract to change the administrators referred to in this article. Both parties shall be notified of the change in writing.

FINAL PROVISIONS

Article 16

The contracting parties shall determine all the amendments to this contract, after a previously reached agreement, with the annexes to this contract.

This contract comes into force on the day when it is signed by both contracting parties.

The contract is drawn up in two (2) identical copies so that each contracting party receives one copy.

| The Client: |  | The Supplier: |
| --- | --- | --- |
| #### Jožef Stefan Institute |  |  |
| Director: |  | Director: |
| Date: |  | Datum: |

We herewith declare that the format of this sample contract is acceptable for the Bidder.

|  |  | The Bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

final acceptance certificate

Appendix 1

| Name of the bidder: |  |
| --- | --- |
| Address: |  |
| VAT identification number: |  |
| Subject: | SUPPLY OF A MULTI-BENCH HARDWARE-SHARING CONFOCAL IMAGING PLATFORM FOR HUNTING OF FUNCTIONAL CELLULAR EVENTS |
| Contract No.: |  |

| Following Delivery note No. |  | of |  | herewith we issue : |
| --- | --- | --- | --- | --- |

| FINAL ACCEPTANCE CERTIFICATE |  | No. |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | place: |  |  |  |
|  | date: |  |  |  |

| Type of purchasing: |  |  |  | Service |  |  |  | Material | X | Equipment |  |  |  |  | Construction |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Delivery method: | x |  | Supplier |  |  |  | JSI arrangement: |  |  |  |  | Other: |  |  |  |  |

| No. | Findings |  | YES |  | NO |
| --- | --- | --- | --- | --- | --- |
| 1 | All the ordered parts have been delivered. |  |  |  |  |
| 2 | The quality of the supplied parts complies with all the terms of the purchase agreement. |  |  |  |  |
| 3 | The required quantity of individual components or services has been supplied. |  |  |  |  |
| 4 | All the required or necessary documentation is included. |  |  |  |  |
| 5 | Allthe manuals including the instruction for use are included. |  |  |  |  |
| 6 | All the other required acceptance tests, as specified in Technical specifications have been carried out and are approved (QA) |  |  |  |  |
| According to the above findings, we confirm acceptance |  |  |  |  |  |

Additional comments:

* Additional final acceptance criteria enclosed

Supplier shall receive a copy of the signed final acceptance certificate.

| The Supplier |  | Institut “Jožef Stefan” |
| --- | --- | --- |
| ((name and surname of the supplie’s representativer) |  | (name and surname of the responsible for the equipment) |

ADDITIONAL FINAL ACCEPTANCE CRITERIA – ACCEPTANCE TESTS (acceptance criteria):

- 1. For each of the measuring stations

- confirm all connected modules (microscope, stage, scanner, lasers, pinhole exchanger, detectors, etc.) are accessible by software

- confirm user computer PC specifications meet minimal requirements specified in the tender

- confirm laser power at the objective entrance pupil to be in accordance with their typical values

- confirm the size of the FOV with 100× objective at least 90 µm × 80 µm, calibrated to real units

- total range of the stage confirmed to be at least 100 mm x 70 mm

- confirm, detectors and pinholes are aligned

- spatial resolution demonstrated with immobilized fluorescent beads of a known size (diameter <45 nm) for every suitable combination of excitation and STED lasers using 100x 1.45 NA oil immersion objective:

  - XY resolution in confocal mode better than 300 nm

  - XY resolution of 60 nm or better in 2D STED

All acceptance resolution tests should be performed as follows:

- intensity profile along a line across a sample object on a non-processed image is taken

- FWHM of the intensity profile is determined by fitting of a Gaussian function

- the process is repeated for at least 10 sample objects from at least 3 independently acquired images

- 50% of the obtained values must be lower than or equal to the required

- max scanning speed demonstrated, e.g., by measuring the time of the image acquisition of an image 512 × 512 pixels at 15 nm pixel size which should be less than 1,3 s

- sync and gating of lasers/detectors demonstrated

- full functionality of the software demonstrated, according to the required specifications above (control, acquisition, analysis, exporting)

- customization demonstrated according to the requirements above:

  - at least 20% (40% on control station that will be connected to the development part) of expected STED power on each station compared to a single station coupled STED power (single station coupled STED should have > 500 mW from the optics fibre output on the main optics path).

  - demonstration of scatter detection on sub-diffraction sized metal beads using 488 nm laser

- demonstration of 8 channel recording with files of a size > 10 MB (preferrable 15 MB) per position with following parameters run with script:

  - < 16s per slice for 750 x 750 pixels image with 0.5 µm pixel size and saved file size > 10 MB

    - one ROI, one tile, one slice is recorded and timed

  - two successfully ran time points 60 regions of interest, 2 tiles, 3 z-slices with time per slice the same as confirmed in a. (no wait time buildup)

    - 10x times in a row, a single ROI recording time is measured and the number divided by tile and z-slice number, average value is calculated

    - total recording time is recorded and divided by number of ROIs, tiles and z-slices

    - both numbers need to be within 15 % of the number determined in a.

- 2. All stations together

- demonstration of simultaneous acquisition independently at the same time on all independently controlled stations

- demonstration of STED capability on independently controlled stations using shared STED laser

A STATEMENT WITH DETAILS ABOUT THE NATURAL AND LEGAL ENTITIES OWNED BY THE BIDDER including the partners and business entities that are considered to be associated with the Bidder in line with the provisions of the law governing companies

Appendix 2

Details obout the Contracting Authority/the tender

| The Contracting Authority | Jožef Stefan INSTITUTE Jamova cesta 39, 1000 Ljubljana, Slovenia |
| --- | --- |
| The code of the tender | JN01/2026 |
| Subject of the tender | SUPPLY OF A MULTI-BENCH HARDWARE-SHARING CONFOCAL IMAGING PLATFORM FOR HUNTING OF FUNCTIONAL CELLULAR EVENTS |

Details obout the Bidder

| Name of the Bidder: |  |
| --- | --- |
| VAT identification number: |  |
| Legal representative of the Bidder: |  |

For the purpose of the 6th paragraph of Article 14 of the Integrity and Prevention of Corruption Act (the Official Gazette of the Repuclic of Slovenia, No.  69/11 – official consolidated text and 158/20), to ensure the transparency of the business and to mitigate corruption risks when concluding legal transactions in the capacity of the Bidder's representative in the process of awarding public contract, I hereby confirm

1.1. that are in the bidder's ownership participated following natural persons:

| No. | Name and surname | Permanent address | Share of ownership in % |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| … |  |  |  |

1.2. that are in the bidder's ownership participated following legal persons:

| No. | Name and registered office of the legal person | VAT identification number | Share of ownership in % |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| … |  |  |  |

the legal person being owned by the following natural persons:

| No. | Name and surname | Permanent address | Share of ownership in % |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| … |  |  |  |

1.3. that are in according of the Companies Act associated companies with the Bidder following economic operators:

| No. | Name and registered office of the economic operator | VAT identification number |
| --- | --- | --- |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |
| … |  |  |

is in mutual relationsjip with the following legal person:

| No. | Name and registered office of the legal person | Vat identification number | Associated as follows |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| … |  |  |  |

I hereby declare that, in the capacity of a natural person - participant in the Bidder's ownership, I stated:

- any natural person who owns through direct or indirect ownership at least 5% of shares or participates in the management rights, management or capital of the legal person with more than a 5% share or has the controlling position in the management of the legal person's funds;

- any natural person who indirectly provides or is providing funds to a legal person and is on such grounds given the possibility of exercising control, guiding or otherwise substantially influencing the decisions of the management or other administrative body of the legal person concerning financing and business operations.

By signing this statement, I hereby guarantee that there are no other natural or legal persons, dormant partners and companies, considered to be related companies pursuant to the provisions of the law regulating companies, in the entire ownership structure.

My signature on this statement warrants that the data provided are true and accurate and I am aware that the contract/the tender/the framework agreement shall be considered null and void, should my statement be false or data untrue. I undertake to inform the Contracting Authority of any changes to the data submitted.

|  |  | Name and surname of the legal representative: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the legal representative) |