Project name: GIGA Nuclear Magnetic Resonance – GIGA NMR

Number of the public contract: 302/7-O-CATS/26

Technical specifications for the “Purchase and delivery of the Equipment for the Center for Solid-State Analysis”

- SUBJECT OF PUBLIC TENDER

The subject of the public procurement is the purchase and delivery of equipment for the Center for Solid-State Analytics, constituting an integrated system for surface analysis, spatially resolved chemical characterization, and the determination of the chemical composition and chemical states of surfaces in demanding research and development environments. The equipment will enable the correlative investigation of heterogeneous systems, thin films, functional surfaces, and trace species, and will provide key infrastructure for advanced research and support for development in biotechnology, pharmaceuticals, and advanced materials.

The public procurement is divided into five lots:

Lot 1: High-resolution secondary ion mass spectrometry system (Orbitrap SIMS)

Lot 2: X-ray photoelectron spectroscopy equipment for operation at near-ambient pressure (NAP-XPS)

Lot 3: Fully automated multi-analysis UHV system for the analysis of surfaces and interfacial layers

Lot 4: Equipment for three-dimensional surface characterization, optical imaging, and elemental analysis of materials

Lot 5: System for nano- and micro-infrared microscopy

These specifications relate to Lot 1: High-resolution secondary ion mass spectrometry system (Orbitrap SIMS).

Orbitrap SIMS: Orbitrap SIMS is the subject of this public tender for the supply, delivery, installation, commissioning, and acceptance of an advanced secondary ion mass spectrometry system intended for biotechnology studies. The instrument shall enable high-confidence chemical identification and spatially resolved molecular characterization of biofunctional materials, including the analysis of surface and near-surface composition, molecular distributions, and depth-dependent chemical changes.

The scope of procurement includes the complete Orbitrap SIMS configuration, including the main instrument, ultra-high-vacuum infrastructure, sample introduction and handling components, primary ion sources and related options, detectors and control electronics, measurement and data processing software with the required number of licenses, delivery under the specified delivery terms, on-site installation, performance verification and acceptance, user training, remote service capability, and the full warranty and support packages defined in the Technical requirements.

- TECHNICAL REQUIREMENTS

The minimum technical requirements for Lot 1 are set out below.

The tender must meet or exceed all minimum technical requirements.

Delivery deadline: The tenderer (bidder) must deliver the equipment that is the subject of this public procurement and carry out all activities for the successful implementation of item 1.28 of the technical specifications at the location University of Maribor, Faculty of Chemistry and Chemical Engineering, Smetanova ulica 17, 2000 Maribor, no later than 660 days from the entry into force of the contract.

Warranty period: For the delivered equipment, the bidder provides at least 60 months of warranty for a flawless technical operation.

LOT 1: High-resolution secondary ion mass spectrometry system (Orbitrap SIMS)

| Item number | Number of pieces | Required |
| --- | --- | --- |
| 1.1 | 1 | The system must include an ion source producing Bin primary ions with the following specifications: - the source must provide large Bin clusters up to n=7 and a high current of Bi1, - total DC current should exceed 40 nA, in pulsed mode (including bunching), a target current of >20 pA must be obtained, - lateral resolution: less than 50 nm, - mass resolution: more than 17 000 at 29 atomic mass units, and at higher mass (> 200 atomic mass units), a resolution of 26 000 or more must be demonstrated (with Bi1+ and Bi3+), - operation in different modes: high-mass resolution and small spot size at the same time, i.e. a mass resolution of 9 000 at atomic mass units >50 atomic mass units and a spot size of <90 nm must be demonstrated simultaneously, - fast imaging of the features of size 90 nm or less, An internal Faraday cup for measuring the current Bin+ primary ions, and O2+ and Cs+ sputter ions must be included in the system. The time-of-flight (TOF) analyzer with the following characteristics must be included: - high secondary ion transmission and high mass resolution simultaneously (transmission higher than 8.0 x 108 Al+/nC, mass resolution higher than 7 000), - parallel detection of all masses of one secondary ion polarity, - mass range: > 12 000 atomic mass units, - in-column measurement of mass separated, pulsed primary ion currents down to 50 fA - a pumping system for an ultra-high vacuum with a demonstrated base pressure of less than 5 x 10–10 torr, The system must include an Orbitrap mass spectrometer in item 1.8 that can be operated in combination with the TOF analyzer. The scope must include the Orbitrap unit, the ion transfer system from the TOF analyzer into the Orbitrap, a TOF analyzer version required for this integration, and full software integration. The system must include a software extension for the instrument control and data acquisition software that enables single-beam focused ion beam (FIB) applications using the system’s standard liquid metal ion gun primary ion source. The functionality must enable the execution of FIB applications without installing an additional dedicated FIB source on the instrument. All ion sources included in items 1.1, 1.2, and 1.3 must be installed on the main analytical chamber together with the TOF analyzer. Any sample treatment by the ion sources must be accomplished without movement of the sample under analysis. |
| 1.2 | 1 | The system must include a titanium sublimation pumping subsystem for efficient pumping of reactive gases, comprising: - a pumping vessel, - an evaporation device for three titanium filaments, - a smart power supply controlled by the vacuum control unit, - support for fully automated outgassing, - automated switching between titanium filaments. |
| 1.3 | 1 | Internal bakeout of the main chamber, including differential pumping of the TOF analyzer. |
| 1.4 | 1 | Internal bakeout of the sample introduction chamber. |
| 1.5 | 1 | Motorized five-axis sample stage with travel in x, y, and z directions, tilting, and rotation. |
| 1.6 | 1 | Fully automated gas flooding system, including a regulative automated leak valve and nozzle, which allows the exposure of the sample to inert and reactive gases. The fully automated switching of gases by the vacuum control unit must be possible and included. |
| 1.7 | 1 | A self-adjusting charge compensation system with the following specifications: - the analysis of insulating samples must be possible without special pre-treatment of the sample and without special sample mounting, - a low-energy electron flood gun with adjustable energy (1 to 20 eV) shall be used for charge compensation, - the charge compensation must work with a maximum repetition rate of up to 50 kHz, independent of sample size and geometry, - the sample must be permanently held at ground potential throughout the complete measurement. - the system must include a high-energy charge compensation option providing an extended electron beam energy range of 1–300 eV, adjustable by computer control. The option must allow analysis of highly insulating samples in interlaced mode with repetition rates up to 50 kHz instead of non-interlaced mode. |
| 1.8 | 1 | The system must include a high-performance, high-resolution (at masses 200 amu and above, a mass resolution of 240 000 or higher) accurate-mass (lower than 1 ppm at mass 50 amu and above) mass spectrometer based on an orbital electrostatic trapping mass analyser (Orbitrap-type) that can be operated in combination with the time-of-flight analyser. The combination must enable the use of fast imaging capabilities of the time-of-flight analyser together with high-performance, accurate-mass analysis for peak identification, including MS/MS features of the analyser. The scope must include the accurate-mass analyser unit, a dedicated version of the time-of-flight analyser equipped with a secondary ion beam switch, an ion transfer system from the time-of-flight analyser into the accurate-mass analyser, and full software integration. |
| 1.9 | 1 | The system must include an extended dynamic range option for the time-of-flight analyser based on attenuation of intense secondary ion signals. The option must include hardware and software for high-speed and mass-selected secondary ion package rerouting via two different attenuation pathways. This must allow processing of up to one hundred times higher count rates with linearity and high reproducibility. Two different attenuation factors must be available; 1/10 and 1/100 attenuation is needed. |
| 1.10 | 1 | A system for (inorganic) sample analysis must include: - dual-source column with an electron impact gas ion source (for operation with oxygen and Ar, Ne, and Xe gases) and a thermal ionization cesium ion source for dual beam depth profiling in combination with the Bi ion cluster gun, with minimum acceleration energy of 100 eV or lower, - a current density of >48 mA/cm2 and >12 mA/cm2 for the oxygen and cesium sputter beams, respectively, - minimum sputter beam diameters: 10 μm for oxygen and 5 μm for cesium, - hardware and software for high-speed and mass-selected secondary ion imaging, - special sample holder and power supply for in-situ, high-speed sample rotation while the sample is under analysis. |
| 1.11 | 1 | A system for organic samples analysis must include: - argon gas cluster ion source, - oxygen operation option for gas cluster source, - internal Faraday cup, - analysis option for the gas cluster source, including additional pulsing systems and software, - operation in the dual-beam mode of depth profiling in combination with the liquid metal cluster ion gun as well as operation as a short pulse length analysis gun (possible operation also in the single-beam mode of depth profiling), - spot size: less than 50 μm (with an Ar cluster current of 10 nA or more), - minimum beam diameter: 5 μm or less. - minimum pulse length (using analysis option): < 20 ns |
| 1.12 | 1 | Computer-controlled heating and cooling of the sample in the sample introduction chamber and in the main analytical chamber while the sample is under analysis, with the following specifications: - possibility to cool the sample to at least −150 °C and the possibility to heat the sample to at least 600 °C, where the full temperature range (from at least −150 °C to at least 600 °C) must be accessible on the same sample holder, - cooling from room temperature to −130 °C must be possible within less than 15 minutes, - an additional dedicated sample holder must be included, with the possibility of cooling and heating, and full movement of the sample in X, Y, and Z directions, rotation, and tilting of the sample. - vessels for liquid nitrogen for the sample introduction chamber and the main chamber need to be included |
| 1.13 | 1 | High-resolution photo navigation system (reflex camera with 24.0 megapixels or higher, integrated sample illumination, USB interface). |
| 1.14 | 1 | A secondary electron (SE) detector. The detection method should enable the acquisition of SE images using different ion guns, both with and without analyzer extraction voltage. It should enable the user to set up operation parameters such as focus and field of view for both secondary ion polarities using SE images. |
| 1.15 | 1 | The system must be supplied with a set of sample holders, compatible with the sample introduction chamber, comprising at minimum: -3 sample holder for rear mounting of 10 samples or more with dimensions of 10 mm (or larger) x 10 mm (or larger) x 8 mm (or larger), including mounting accessories and a box for storage -3 sample holder for top mounting of multiple samples on an area of at least 90 mm x 70 mm with a thickness of at least 20 mm or higher, including mounting accessories and a box for storage |
| 1.16 | 1 | The system must include a mass spectrometry reference library for MS and MS/MS data analysis with the following minimum requirements: -the library must include spectra of more than 340,000 reference compounds with MS data, -the library must include more than 50,000 reference compounds with MS/MS data, -library searching must be initiated directly from the instrument’s data analysis software as part of the standard data analysis workflow, -acquired spectral data must be exportable in a NIST-compatible format to enable creation of local custom libraries usable in database search queries. |
| 1.17 | 1 | The system must include an adaptation for an external vacuum cryogenic transfer vessel interface at the sample introduction chamber. The adaptation must enable transfer of samples from a vacuum cryo transfer vessel onto a specialised heating and cooling sample holder inside a sample introduction chamber. The scope must include a dedicated adapter for the specialised heating and cooling sample holder. |
| 1.18 | 1 | The system must include two separate sample-holder parking positions inside the sample introduction system. The parking positions must allow storage of up to two sample holders inside the sample introduction chamber. |
| 1.19 | 1 | The system must include a transfer vessel for transporting a complete sample holder under an inert-gas atmosphere or under medium vacuum. The transfer vessel must enable the complete sample holder without exposing the samples to ambient air. The scope must include the transportation vessel and a dedicated version of the sample introduction chamber required for compatibility with the transfer vessel. |
| 1.20 | 1 | The system must include an additional sample introduction system as a second sample introduction chamber for the instrument. The option must include an internal bake-out system for the second sample introduction chamber. |
| 1.21 | 1 | The system must include a glovebox module for the sample introduction system that allows the introduction of samples and sample holders under an inert gas atmosphere. The glovebox must meet the following minimum requirements: -transparent, crack- and bubble-free enclosure, -designed connections for external inert gas supply, -one inlet sluice for transferring items into the glovebox, -two gloves for handling samples and holders inside the glovebox, -bellows connection to the sample introduction system to enable transfer of samples and sample holders under an inert atmosphere, -attainable cleanliness levels of oxygen below 200 ppm and water below 200 ppm -at least one safety valve, -oxygen sensor for the glovebox and for the inlet sluice, -water (humidity) sensor for the glovebox and for the inlet sluice, -automated inertisation functionality, -compatibility with the sample introduction chamber system of the SIMS instrument -additional sample introduction system must be combined with the glove box |
| 1.22 | 1 | Licensed software for measuring and processing data for at least 10 users. |
| 1.23 | 1 | Dedicated process computer with the following specifications:* - MS Windows 11 Professional, 64-bit operating system, - latest-generation processor with at least 16 physical cores and a maximum boost frequency of at least 5.7 GHz, or better, - 64 GB RAM or more, 8 TB SSD or more, - 2 × 27'' (or larger) LED displays with Full HD resolution or higher than Full HD resolution. * Special requirements for item 1.23 in accordance with the Decree on Green Public Procurement are set out at the end of this document. |
| 1.24 | 1 | Warranty: 60 months or more must cover: - the replacement costs of the defective parts, - the corresponding labor expenses, together with travel and living expenses of the bidder’s personnel. All parts and services necessary to maintain the instrument in optimum condition and at high availability shall be provided. Everything necessary for the successful performance of surface analyses must be covered at no additional cost during the warranty period. This includes: - all spare parts needed, - all consumables needed for 1200 hours of operation per year or more, including their on-site exchange performed by the bidder’s expert, - one preventative maintenance visit per year (two working days on-site, 8 h per day), - unlimited service visits, - free software updates. |
| 1.25 | 1 | Basic operation training Acquaintance with the equipment and basic training, provided by the bidder’s engineer, for up to four users, lasting at least 5 working days, 8 hours per working day (this also includes the handover of all necessary data and passwords, as well as the travel and accommodation costs of the bidder’s engineer; the training shall be conducted at the FKKT UM location designated by the contracting authority), in Slovene or English. This item shall be deemed completed after the delivery, installation, and commissioning of the equipment. |
| 1.26 | 1 | An advanced user training package consisting of two separate on-site advanced training sessions to be delivered after acceptance, comprising (training is carried out at the location of FKKT UM, determined by the client): -one advanced training session scheduled no later than 31 days after acceptance, duration of 5 working days (8 hours per working day) -one additional advanced training session scheduled no later than 31 days after acceptance, duration of 5 working days (8 hours per working day) Each advanced training session must be delivered for up to four persons and must include the travel, living, and accommodation costs of the training bidder’s personnel, and it must be conducted in Slovene or English. |
| 1.27 | 1 | Provided remote access to the device for the needs of calibration and control of the device, at least for the duration of the warranty period (60 months). The instrument design must facilitate remote operation and fault diagnostics. |
| 1.28 | 1 | Delivery and installation on the client’s premises; 3rd floor at FKKT UM. Condition for the successful execution of the item and signing of the acceptance report: - the supply, installation, and commissioning of the equipment on the client’s premises by the bidder, - introduction to equipment and training, - delivery of the operating instructions and other technical documentation in digital form and in English, - all performance specifications (items from 1.1 to 1.25) must be demonstrated on-site (FKKT UM) before signing acceptance documents. |

SPECIAL REQUIREMENTS FOR ITEM 1.23

The equipment under item 1.23, Dedicated Process Computer, must fully comply with the technical specifications, ensuring the fulfillment of the objectives in accordance with the Decree on Green Public Procurement (Official Gazette of the Republic of Slovenia, Nos. 51/17, 64/19, 121/21, 132/23, and 43/25), namely:

DEDICATED PROCESS COMPUTER

Power Consumption MANDATORY: The personal computer must belong to the highest energy efficiency class available on the market. Method of proof: The bidder must attach the following to the tender: - a certificate showing that the product has a Type I environmental label or a label from another certification scheme (e.g. Energy Star V9.0, EPEAT, TCO) that meets the specified requirements. Appropriate supporting documents may also include alternative test results obtained by testing bodies accredited in accordance with ISO 17025, in compliance with IEC 62623.

ELECTRONIC DISPLAYS

Electrical energy consumption MANDATORY – Every electronic display supplied under the contract must be classified in the highest energy efficiency class available on the market, an energy class of at least B,* as defined in Regulation (EU) 2017/1369 establishing a framework for energy labelling and in Commission Delegated Regulation (EU) 2019/2013 with regard to the energy labelling of electronic displays, as amended by Commission Delegated Regulation (EU) 2021/340. Method of proof: For each model offered, the bidder must submit: - the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling, or - a declaration stating that all delivered products will bear the valid required energy label issued in accordance with Regulation (EU) 2017/1369 establishing a framework for energy labeling. * Note: The contracting authority refers to the best EU energy efficiency class available at the time of the public procurement, including at least 25 registered models of electronic displays in the European Product Registry for Energy Labeling (EPREL).

The operation is co-financed by the European Union from the European Regional Development Fund under the European Cohesion Policy Programme 2021–2027, Policy Objective 1: A more competitive and smarter Europe by promoting innovative and smart economic transformation and regional ICT connectivity, under Priority 11: Development of Strategic Technologies for Europe – STEP, and under Specific Objective RSO1.6: Support for investments contributing to the objectives of the Strategic Technologies for Europe Platform (STEP).